## Sheet1
Recall Number | Recall Event ID | Device Type | Product Code | Regulation Description | Medical Specialty | Review Panel | Submission Type | Trade Name/Product | Main Name | Recall Class | Date Posted | Year Posted | Recalling Manufacturer | Reason for Recall | FDA Determined Cause | Action | Consumer Instructions | Clean Quantity | Quantity in Commerce | Distribution | Recall Status | Termination Date | Time to Terminate | Merged Quantities | Fault_Class
Z-1788-2014 | 68426 | No | EJJ | Dental amalgam, mercury, and amalgam alloy. | Dental | Dental | 510(k) | Tytin Regular Set, Double Spill, 50 Capsules, Part Number 29948, Lot Number 3-1294. |  |  | May 29, 2014 | 2014 | Kerr Corporation | Kerr Corporation is recalling one (1) lot of Tytin Regular Set ( Part Number 29948, Lot Number 3-1294), because it may set faster than specified in the Directions for Use. | Unknown/Undetermined by firm | Kerr Corporation initiated this recall by sending the recall notification letter accompanied by a return form, to customers via USPS 1st class mail. The consignees are instructed to complete the return form and return any affected product to Kerr Corporation.The recall notification letter, dated May 29, 2014, titled "URGENT: MEDICAL DEVICE RECALL", informed customers of the recall by providing the following information: product description with part number and lot number, distribution date range, reason for recall, risk statement, instructions to customers and/or Kerr distributor on what to do with the recalled product, FDA MEDWATCH reporting system info, and company contact information, and instructed customers to complete the return form titled "Acknowledgement and Recall Return Form". |  | 2030 | 2,030 units | Worldwide Distribution-USA (nationwide) and the countries of Canada, Mexico, Australia, Belgium, Czech Republic, Germany, Spain, France, Great Britain, French Guiana, Italy, Portugal, and Saudi Arabia. | Terminated | October 27, 2014 | 151 | 2030 | Undetermined
Z-1927-2014 | 68537 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Siemens Healthcare ADVIA Acid/Base Reagent 1 and 2 Kit containing ADVIA Centaur Acid Reagent bottles, lot number: 454778 |  |  | June 02, 2014 | 2014 | Siemens Healthcare Diagnostics, Inc | Swollen acid bottles due to a trace metal contamination in the Acid Reagent bottles | Material/Component Contamination | Siemens Healthcare Diagnostics issued Urgent Medical Device Recall (UMDR) to customers in the United States on June 02, 2014 and a Urgent Field Safety Notice (UFSN) issued outside the United States on June 02, 2014.These notices inform the customer of the potential for swollen acid bottles due to confirmed metal contamination. Customers with alternative lots in inventory were instructed to transition to the alternative lot upon receipt of the communication. Customers without an alternative lot of acid in inventory were informed that Siemens proactively shipped an alternative lot to their laboratory and they should transition to that lot upon its receipt. Siemens also informed customers without an alternative lot that they may continue to run and report patient results using the affected lot of acid.Questions contact your Siemens Customer Care Center or your local Siemens technical support representative. |  | 3839 | 3,839 kits | Worldwide Distribution - US Nationwide and the countries of Canada, Argentina, China, Colombia, Korea, Malyasia, New Zealand, Mexico, Guatemala, and Singapore. | Terminated | February 05, 2015 | 248 | 3839 | Undetermined
Z-1846-2014 | 68384 | No | LHN | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | ProBeam Proton Therapy System (HGA and HFB) |  |  | May 09, 2014 | 2014 | VARIAN MEDICAL SYSTEMS PARTICLE THERAPY GMBH | Anomaly with the ProBeam System where under certain conditions, the Treatment Control and Monitoring application could fail to send treatment history records to the ARIA database. | Under Investigation by firm | Varian sen an Urgent Device Medical Correction/Urgennt Field Safety Notice letter dated May 9, 2014 to affected customers. The letter identified the affected product, problem and actions to be taken. For questions cotact Varian Service Support. at debg-PT-service-support@varian.com. |  | 1 | 1 | US Distribution in tthe state of CA | Terminated | November 12, 2014 | 187 | 1 | Undetermined
Z-1745-2014 | 67879 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Brilliance CT Big Bore Oncology Whole Body Computed Tomography X-Ray System &amp; Brilliance CT Big Bore Radiology Whole Body Computed Tomography X-Ray System. |  |  | December 19, 2013 | 2013 | Philips Medical Systems (Cleveland) Inc | During internal testing, the Brilliance CT Big Bore was found to be out of tolerance for radio frequency emissions. Specifically, at the 48MHZ frequency, the testing indicated the Brilliance CT Big Bore was 3.5dB uV/meter higher than the applicable IEC 60601-1-2 standard specification. | Nonconforming Material/Component | Philips Medical System sent an Urgent Medical Device Corrections letters dated December 10, 2013 to affected customers. The letter identified the affected product, problem and actions to be taken. The letter informed customers that a field service Engineer will visit the facility to perform an Electro-Magnetic Interference (EMI) noise reduction up0date on each system. For information or support concerning contact your local Philips representative:For North America and Canada, contact the Customer Care Solutions Center (1-800-722-9377, OptionS: Enter site ID or follow prompts). |  | 454 | 454 Units | Worldwide Distribution - US (nationwide) in the states of AK, AL, AR, AZ, CA, CO, CT, FL, GA, HI, IA, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, NC, ND, NH, NJ, NM, NV, NY, OH, OK, OR, PA, PR, RI, SC, SD, TN, TX, UT, VA, WA, WI &amp; WV.and countries of Algeria, Australia, Austria, Bangladesh, Belgium, Brazil, Canada, China, Czech Republic, Denmark, Ecuador, France, French Polynesia, Gabon, Germany, India, Indonesia, Israel, Italy, Jordan, Libyan Arab Jamahiriya, Lithuania, Luxembourg, Malaysia, Martinique, Mexico, Morocco, Netherlands, New Zealand, Oman, Palestine, Philippines, Portugal, Russia, Saudi Arabia, Singapore, Slovakia, Slovenia, South Africa, South Korea, Spain, Sweden, Switzerland, Syrian Arab Republic, Thailand, Ukraine &amp; United Kingdom. | Terminated | August 25, 2015 | 614 | 454 | Undetermined
Z-1852-2014 | 68491 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Spacelabs Healthcare Qube Compact Monitor, Model 91390. |  |  | June 06, 2014 | 2014 | Spacelabs Healthcare Inc | Spacelabs Healthcare is recalling Qube Compact Monitor, Model 91390, due to a potential chance of losing network connection after switching from hardwired Ethernet connection to using the Qube Docking Station without powering down the device. The bedside monitor and all of its alarms will continue to function normally. No one has been injured as a result of this issue. | Component design/selection | Spacelabs Healthcare Snoqualmie, WA sent an Urgent Medical Device Correction letter, dated June 6, 2014, to all consignees. The letter identified the product, the problem, and the actions to be taken by the customer. The letter instructed customers using docking station ready qube monitors with qube docking stations immediately advise their staff of this situation; and, to follow the steps provided, if they continue to use a docking station ready qube monitor with a docking station.Spacelabs Healthcare will contact customers at the earliest possible date to schedule a convenient time for Spacelabs to correct customer's affected qube monitors at no cost. Questions should be directed to Spacelabs Healthcare Technical Support at 1-800-522-7025. |  | 4842 | 3383 in the USA; 1459 outside the USA | Worldwide Distribution -- USA, including the states of AL, AZ, CA, CO, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, ME, MI, MN, MO, MS, MT, NC, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, VA, WA, WI, and WY; and, the countries of AFG,ARE, ARG, AUS, BHR, BOL, CAN, CHE, CHL, CHN, COL, CRI, DEU, DOM, ECU, FRA, GBR, IND, ITA, JOR, KWT, LBY, MYS, NLD, OMN, PAN, POL, PRT, ROU, SAU, SGP, THA, TTO, TUN, TUR, TWN, and VNM. | Terminated | February 17, 2015 | 256 | 4842 | Undetermined
Z-1853-2014 | 68485 | refractor, manual, non-powered, including phoropter | HKN | Manual refractor. | Ophthalmic | Ophthalmic | 510(K) Exempt | RT-5100 Refractor and RT-3100 Refractor; |  |  | June 09, 2014 | 2014 | Nidek Inc | Near Point Chart Arm of RT 5100 and RT-3100 Refractors may lower spontaneously with potential for patient injury. | Use error | Recall/Field Correction letters were sent to doctors on June 9, 2014 with a preventive maintenance instruction to be performed on the device by user. Also included are the necessary tools to perform actions. Customers are to complete and return the Recall Response Card. Questions are to be directed to Customer Service.Contact Information: Customer Service: 1-800-891-1772 E-mail: RT-5100 update@noritsu.com |  | 167 | ~4,285 units RT-5100; 167 units of RT-3100 | US distribution only. | Terminated | October 30, 2015 | 508 | 167 | Undetermined
Z-1864-2014 | 68588 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | Ingenia 3.0T, magnetic resonance imaging system. |  |  | June 13, 2014 | 2014 | Philips Medical Systems, Inc. | For some identified Ingenia systems shipped between December 2013 and March 2014, cover set rear magnetic M10 nuts were erroneously supplied instead of A4 stainless steel nuts. | Nonconforming Material/Component | On June 14, 2014, Field Safety Notice (FSN 78100432), was be sent to all US and Canadian consignees via certified mail. Philips Key Market representatives will distribute the letter internationally.The FSN instructs consignees to not let hospital technical staff or any person other than Philips service engineers dismount the rear covers before the system is inspected and/or corrected by Philips. Philips has alerted installation and service engineers to this hazard. A Field Change Order with reference FCO781 00432 is being released directing Philips field service engineers to inspect the Ingenia rear cover set nuts and replace them if needed with non-magnetic stainless steel nuts. For further information or support concerning this issue, consignees should contact their local Philips representative at 1-800-722-9377. |  | 18 | 18 | Worldwide Distribution-USA (nationwide) including the states of AL, AZ, CA, CO, GA, ME, MI, NY, OH, PA, TX, and WA, and the countries of Mexico, Argentina, Australia, Belarus, Belgium, Brazil, Chile, China, France, Germany, Hong Kong, India, Indonesia, Ireland, Italy, Japan, South Korea, Laos, New Zealand, Norway, Palau, Poland, Portugal, Qatar, South Africa, Sweden, Switzerland, Taiwan, Tunisia, Turkey, Turkey, United Kingdom, Venezuela, and Vietnam. | Terminated | January 27, 2016 | 593 | 96 | Undetermined
Z-1854-2014 | 68354 | Yes | FKR | Hemodialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | NxStage Dialysate Sack (SAK) with Concentrate |  |  | May 14, 2014 | 2014 | NxStage Medical, Inc. | SAK Dialysate Concentrate contain aluminum levels which exceed internal product specification | Material/Component Contamination | NxStage issued Initial recall letter dated 5/14/14 to dialysis centers. An error was noted and an updated letter dated 5/15/14 was resent to dialysis centers. The updated 5/15/14 letter was emailed to all current active PureFlow SL patients who had provided an email address followed by UPS next business day letter sent to all current active PureFlow SL patients. Accounts are requested to examine inventory and remove product. Contact NxStage Customer Service to arrange for return of all affected product and forreplacement product to be sent. If you have any questions or comments, please feel free to contact NxStage Customer Service at 1-866-NXSTAGE (1-866-697-8243). |  | 13474 | 13474 cases | Worldwide Distribution - USA (nationwide) and Internationally to France, Netherlands, Spain, Sweden, and the UK. | Terminated | March 13, 2017 | 1034 | 69334 | Undetermined
Z-1873-2014 | 68343 | No | CGA | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Glucose Hexokinase Liquid Reagent For the in vitro quantitative determination of Glucose in serum. |  |  | April 07, 2014 | 2014 | JAS Diagnostics Inc. | JAS Diagnostics, Inc. of Miami Lakes, FL is recalling their Glucose Hexokinase reagent due to certain lots not performing according to their specifications. | Mixed-up of materials/components | JAS Diagnostics, Inc. sent an Important Product Recall Information letter dated April 7, 2014 to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. The customers were instructed to: 1) Screen their inventory for any of the lots listed in the letter and discontinue use immediately. 2) Contact the Customer Service Department at 305-748-2762 or 305-748-2763 to obtain replacement product. |  | 407 | 407 | Worldwide Distribution - US including Puerto Rico and the states of CA, FL, TX, PA, GA, KY, MI, NY, OK, MN, TN, VA, NC, NV, NJ and WV, and the countries of Barbados, Ghana, Trinidad, Tobago, Canada and Haiti. | Terminated | August 21, 2014 | 136 | 407 | Undetermined
Z-1872-2014 | 68405 | No | HRX | Arthroscope. | Orthopedic | Orthopedic | 510(k) | Shaver Handpiece Sterilization Tray. Model number 272-700-000. For sterilization of Stryker Endoscopy arthroscopic shaver. |  |  | May 19, 2014 | 2014 | Stryker Endoscopy | STERRAD 100S parameters provided in the Shaver Handpiece reprocessing guide (P10299 Revision A) did not fully consider all worst-case scenarios during validation testing. | Nonconforming Material/Component | Device Correction letters were sent on May 23, 2014 to all users by certified mail. Internation sites will receive notification and be responsible for contacting customers. |  | 1992 | 1992 trays | Worldwide Distribution. | Terminated | May 07, 2015 | 353 | 1992 | Undetermined
Z-1876-2014 | 64247 | No | MBI | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | ANSPACH***REF 98-0020***Modified Arthroscopic Knotless Suture Anchor Deployment Tool with Suture Puller***Rx Only*** |  |  | November 28, 2012 | 2012 | The Anspach Effort, Inc. | The Anspach Effort, Inc. in Palm Beach Gardens, FL is recalling their Anspach Custom Devices. The units were not manufactured consistently with Quality System requirements. | Process design | The Anspach Effort, Inc. sent an Urgent Medical Device Recall letter dated December 21, 2012 to all affected customers. The letter identified the affected products, problems and actions to be taken. Customers were instructed to remove and return the affected products and complete the attached reply form and return to The Anspach Effort Inc. For questions contact Anspach Product Support at (800) 327-6887.For questions regarding this recall call 561-494-3706. |  | 1 | 1 | Worldwide Distribution - USA including AL, CA, FL, IA, IL, KY, MN, MS, NC, OH, OK, PA, TN, TX.and Internationally to Canada, Japan, Hungary, Korea and Sweden.. | Terminated | August 07, 2014 | 617 | 788 | Undetermined
Z-1874-2014 | 68447 | No | IPF | Powered muscle stimulator. | Physical Medicine | Physical Medicine | 510(k) | IF 8000; 1/F Interferential Electrical Stimulation Device; NMES (neuromuscular electrical stimulation) mode. |  |  | January 25, 2009 | 2009 | Zynex Medical, Inc. | Zynex notified customers that a malfunction in the IF 8000 can result in the device ceasing to operate, and in some instances,overheating and becoming hot to the touch. | Component design/selection | Consignees were notified via letter on 6/25/2009. The letter identified the reason for the recall, and stated that customers could contact customer service at 1-866-940-7030 if they encountered any difficulties with the affected product. Otherwise, they could continue using the product as they have in the past. |  | 1461 | 1,461 | Nationwide Distribution | Terminated | July 02, 2015 | 2349 | 1461 | Undetermined
Z-1840-2014 | 68379 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Siemens LANTIS System ; LANTIS Commander, 10 Users ; LANTIS 6.1 Commander ; LANTIS 8.3 Commander |  |  | May 06, 2014 | 2014 | Siemens Medical Solutions USA, Inc | There is a potential risk when using LANTIS OIS System client software with operating systems for which it has not been validated and released, resulting in incorrectly stored data, such as treatment records, and LANTIS database corruption. | Use error | Siemens sent an Urgent Medical Device Correction Field Safety Notice dated was sent to end users on/about May 6, 2014, to notify them of the safety issue and provide instructions to avoid potential risks. Customers were instructed to make sure all computers with the LANTIS client software installed are running on an operating system that was validated by Siemens for opration with LANTIS, e.g. Microsoft Windows XP Service Pack 3 (SP3). Customers were asked to include the Field Safety Notice in their System Owner Manual chapter "Safety Advisory Letters" where it should remain and that all affected personnel affected be informed.For questions regarding this recall call 1-610-219-6300. |  | 108 | 108 | Nationwide Distribution | Terminated | December 15, 2014 | 223 | 108 | Undetermined
Z-1926-2014 | 66364 | No | JIT | Calibrator. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Siemens Dimension LOCI TNI Troponin I Calibrator RC 621. |  |  | September 17, 2013 | 2013 | Siemens Healthcare Diagnostics, Inc. | Siemens Dimension LOCI TNI Troponin I calibrator lot 3DD007 may demonstrate an average upward shift of 24% in QC and patient results following calibration. | Process control | Siemens sent an Urgent Medical Device Recall letter dated September 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Customers were advised immediate discontinuation of use and discard of the affected device.Please complete the attached Effectiveness Check form and fax it to the Siemens Customer Care Center at (302) 631-8467 to indicate that you have received this information. Also, please indicate any replacement product required.Please contact Siemens Customer Care Center- Technical Solutions at (800) 441 9250 if you are aware of any adverse event that occurred as a consequence of this issue, or for immediate product needs.Please retain this letter with your laboratory records, and forward this Jetter to those who may have received this product.We apologize for the inconvenience this situation has caused. If you have any questions, please contact your Siemens Customer Care Center or your local Siemens technical support representative. |  | 1625 | 1625 | Distribution: US Nationwide including Puerto Rico | Terminated | August 27, 2014 | 344 | 1625 | Undetermined
Z-1925-2014 | 68418 | No | JDC | Elbow joint metal/polymer constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | DISCOVERY ELBOW prosthesis, 4 MM X 100 MM RIGHT HUMERAL FLANGED I BOND COAT Tl 6AL 4V ALLOY. |  |  | May 15, 2014 | 2014 | Biomet, Inc. | Surface finish is different than specified. The implant has 30 grit blast on the exterior of the flange when the print calls for the glass bead blast. The patient may experience pain requiring surgical intervention. | Nonconforming Material/Component | Biomet sent an Urgent Medical Device Recall Notice letter dated May 28, 2014 to their customer. The letter identified the affected product, problem and actions to be taken. The letter instructed physician to monitor the patient for any pain that may require surgical intervention. For questions call (574) 372-1570. |  | 1 | 1 | US Nationwide Distribution in the state of California | Terminated | March 16, 2015 | 305 | 1 | Undetermined
Z-1941-2014 | 68396 | No | OMN | Microorganism differentiation and identification device. | Microbiology | Microbiology | 510(K) Exempt | illumipro-10 Incubator/Reader, Meridian Bioscience, Inc. |  |  | April 22, 2014 | 2014 | Meridian Bioscience Inc | Illumipro-10 instruments may have Block B chambers operating at an incorrect temperature as a result of incorrect temperature calibration. | Employee error | Meridian sent an Urgent Medical Device Recall letter dated April 25, 2014, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to immediately return their unit to Meridian Bioscience. Meridian's Technical Support Staff will call them directly to provide them with a return authorization number and instructions for decontaminating, packaging and shipping the unit. For additional information customers were instructed to call 800-343-3858 (North America or 513-271-3700 (Worldwide). |  | 35 | 35 instruments | Worldwide Distribution - USA including AL, AZ, CA, CO, FL, IA, IN, LA, MA, ME, MI, MN, NJ, NY, OH, PA, and WV. Internationally to Australia. | Terminated | June 30, 2014 | 69 | 35 | Undetermined
Z-1947-2014 | 68335 | No | LRO | Surgical drape and drape accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | Enforcement Discretion | Alcon CUSTOM-PAK containing Anterior Chamber Cannula 27G Component Part Number 585006 contained in Alcon Assembly Numbers 9795-37, 11233-14, and 13334-04. |  |  | May 16, 2014 | 2014 | Alcon Research, Ltd. | Alcon CUSTOM-PAK which is supposed to contain a 27G Anterior Chamber Cannula (27G x 7/8 in) actually contains a 27G Sharp Tip Needle | Employee error | Alcon sent an Urgert Market Withdrawal letter dated May 14, 2014, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to review their inventory to determine if they have any affected product, place the affected product into quarantine to ensure it is not used, return the attached Response Form via fax to 817-302-4337 or email market.action@Alcon.com.Customers with questions were instructed to call 1-800-862-5266. |  | 684 | 684 total (525 units in US) | Nationwide Distribution including North Carolina and Maine | Terminated | December 09, 2014 | 207 | 684 | Undetermined
Z-1835-2014 | 67713 | No | IXY | Wall-mounted radiographic cassette holder. | Radiology | Radiology | 510(K) Exempt | The Claymount Harmony is a grid cabinet designed to accept a digital x-ray imaging receptor and stationary grid. It contains signaling electronics to indicate the presence of the grid and detector as well as detector orientation. It may be mounted to a wall stand or incorporated in a radiographic table. |  |  | March 15, 2014 | 2014 | Claymount Assembli | Harmony1 models of the Harmony Bucky is incorrectly labeled in that the label does not list the manufacturer address, the date of manufacture, or a statement of compliance with 21 CFR Subchapter J. | Radiation Control for Health and Safety Act | CDRH approves the CAP subject to the following conditions for Claymount Assemblies BV:1. Notification of all dealers and purchasers is to be made within 15 working days of receipt of this letter in the manner specified in 21 CFR 1003.21 and 1003.22. This office and the Food and Drug Administration (FDA) district office coordinator noted below are to be included in the notification.2. All forward production of Harmony devices must be properly tested, labeled, and reported when imported into the United States. For further questions please call (630) 271-9729. |  | 125 | 125 units shipped to United States | Worldwide Distribution - US including the states of CA, CO, FL, IL, NC, NJ, OH and PA and the countries of Canada, France, Italy, Netherlands, South Africa, Switzerland, Turkey and United Kingdom. | Open, Classified |  |  | 125 | Undetermined
Z-1836-2014 | 67715 | No | IXY | Wall-mounted radiographic cassette holder. | Radiology | Radiology | 510(K) Exempt | The Claymount Harmony is a grid cabinet designed to accept a digital x-ray imaging receptor and stationary grid. It contains signaling electronics to indicate the presence of the grid and detector as well as detector orientation. It may be mounted to a wall stand or incorporated in a radiographic table. It is intended to be used in general radiographic applications as a component of a radiographic system. |  |  | March 15, 2014 | 2014 | Claymount Assembli | Claymount Assemblies BV is recalling the Harmony2 model of Harmony Bucky because 1) the devices are incorrectly labeled in that the label does not list the manufacturer address, the date of manufacture, or a statement of compliance with 21 CFR Subchapter J, and 2) the device's insulation values do not comply with IEC 60601-1 in all cases. | Radiation Control for Health and Safety Act | MEDICAL DEVICE RECALL letters dated March 15, 2014 were sent to all direct accounts. The letters includes instructions to: 1) complete and return the Acknowledgement and Receipt Form accompanying the letter; 2) replace the Harmony2 Bucky(s) with new Harmony3 Bucky(s) that the direct account will be receiving in the near future; 3) if the direct account is a distributor and/or installer and it elects to have their customer(s) perform the replacement, the customer(s) should be provided with a copy of the letter, the replacement Harmony3 Bucky(s), and the replacement notification card, or the direct account can provide the names and addresses of their customers in which case Claymount will contact the customers directly; and, 4) once each unit is replaced, the replacement completion notification card should be mailed or faxed to the Claymount USA office and the Harmony2 Bucky(s) should be returned to "Claymount Americas Corp., ATTN: Alice Packard, 2545 Curtiss St., Downers Grove, IL 60515-4059" or faxed to 630-271-9995.For questions customers were instructed to call the following numbers:USA (630) 271-9729EUR +31 315 659 150 |  | 12 | 12 devices | Worldwide Distribution - USA including NC, OH, CO, CA, NJ, IL, OH, PA, FL and Internationally to Canada, China, France, Italy, Netherlands, South Africa, Switzerland, Turkey, and United Kingdom. | Open, Classified |  |  | 12 | Undetermined
Z-1961-2014 | 68498 | No | NKB | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | MIS Rod 2, a component of the NewPort System. |  |  | June 09, 2014 | 2014 | Integra LifeSciences Corp. | Integra Lifesciences is recalling NewPort System MIS Rod, 2 because it may be up to 30 degrees offset from specification requirements causing misalignment with the inserter. This could potentially result in the lordosis of the rod not aligning with the handle when the inserter is used in accordance with the current NewPort MIS System Surgical Technique. | Under Investigation by firm | A customer notification letter was sent on 6/9/14 to their customers and distributors who purchased the Integra NewPort System MIS Rod, 2 to inform them there may be a potential that the hex on some NewPort System MIS Rod, 2 are up to 30 degrees offset from specification requirements causing misalignment with the inserter. Customers are instructed to complete and return the customer response form and return by email or fax as indicated on the form. Customers with questions are instructed to contact Customer Service at (866) 942-8698. |  | 94 | 94 units | Worldwide Distribution. US including states of CA, FL, IL, MD, MO, TN, TX, and countries of Australia, Czech Republic and Spain. | Terminated | November 06, 2014 | 150 | 94 | Undetermined
Z-1959-2014 | 68398 | Yes | CAC | Autotransfusion apparatus. | Anesthesiology | Anesthesiology | 510(k) | Atrium Express Dry Seal Chest Drain ATS Blood Recovery, Catalog Number: 4050-100N. |  |  | April 28, 2014 | 2014 | Atrium Medical Corporation | Product that was previously recalled because the Chest Drain tubing of the ATS Blood recovery may leak or disconnect, was released for distribution in error. | Nonconforming Material/Component | Smith &amp; Nephew notified Distributors on April 28, 2014 by telephone. A follow up e-mail and a recall notification letter were issued by UPS on 6/27/14. Distributors were requested to notify customers who received the lot. A reply acknowledgement form was to be completed and returned. Return Goods Authorizations for the return of the goods were issued to customers. |  | 195 | 195 Cases; Total 1170 units | Distributed in the states of AK, CA, MD, MN, MO, NC, NJ, OH, and UT. | Terminated | March 17, 2016 | 689 | 195 | Undetermined
Z-1940-2014 | 68333 | freezer | JRM | General purpose laboratory equipment labeled or promoted for a specific medical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Stirling Ultracold Portable Shuttle, ULT-25. |  |  | March 20, 2014 | 2014 | Global Cooling Inc | The external power supply input is overheating at the connection to the freezer. | Use error | Consignees received an Urgent Medical Device Recall letter, dated March 20, 2014, a Recall Response letter, a User Alert, and a label to attach to the freezer power supply input. The letter identified the reason for the recall and how to recognize that the device may fail. Customers are to discontinue use of the device if there are any signs of a loose connection, corrosion, discoloration, or melting; and take the necessary actions identified, including contacting Global Cooling Customer Service at service@stirlingultracold.com. The Recall Return Response was to be completed and returned. |  | 413 | 413 units | Worldwide Distribution -- USA, including the states of AK, AR, AZ, CA, CO, CT, FL, GA, GU, HI, IA, ID, IL, IN, KS, LA, MA, MD, MI, MN, MO, MT, NC, NE, NJ, NM, NY, OH, OK, OR, PA, SC, SD, TN, TX, UT, VA, VT, and WA; and, the countries of Australia, Brazil, Canada, China, India, Indonesia, Japan, Korea, Kuwait, Mexico, Netherlands, Pakistan, Singapore, Switzerland, Taiwan, Thailand, United Arab Emirates, and United Kingdom. | Terminated | July 01, 2015 | 468 | 413 | Undetermined
Z-1841-2014 | 68362 | No | NKB | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | Integra UCR 3mm Hex, Large Axial Driver. |  |  | May 21, 2014 | 2014 | SeaSpine Inc | Integra has identified through an investigation of complaints that there may be the potential for the tip of asingle lot number of the UCR 3mm Hex, Large Axial Driver, 91-1123 to twist or become rounded duringuse. | Under Investigation by firm | Integra sent an Urgent Voluntary Medical Device Recall letter dated May 21, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Customers were instructed:1. If you have distributed the affected product, identify those customers and forward the attached Integra Medical Device Recall Notice to them.2. Inspect your inventory of UCR 3mm Hex, Large Axial Drivers (Catalogue Number: 91- 1123). If you have product with lot number MO16090B, STOP USING IMMEDIATELY. If you have product of any other lot number, you can continue to use it. The event has been isolated to this single lot.3. We also request that regardless if you have a UCR Driver with the affected lot, you complete the attached form and return it to us as soon as possible as indicated. We recommend you keep a copy of the form for your records.4. When your form is received, Customer Service will contact you and provide an RMA number, directions to return the product, and input an order to replace the quantity you indicate on the form.Should you have any questions regarding these instructions, please contact Customer Service at 1-866- 942-8698.We apologize for any inconvenience this may cause and thank you for your cooperation in this effort.:. |  | 7 | 7 | US Distribution including the states of : AL, CA, OH, NM and TN. | Terminated | December 15, 2014 | 208 | 7 | Undetermined
Z-1842-2014 | 68339 | No | LPH | Hip joint metal/polymer/metal semi-constrained porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | VerSysÂż Hip System, Beaded Fullcoat Stems, Zimmer, Inc., Warsaw, IN 46581 |  |  | May 19, 2014 | 2014 | Zimmer, Inc. | During routine inspection, beaded hip stem forgings received from supplier failed to meet ZES 2A-102. Processing of Beaded CoCrMo Alloy Hip Stem, material requirements per ATS #14-04818 and ATS #14-05549, respectively. | Nonconforming Material/Component | On 6/2/2014 the firm sent "Urgent Medical Device Recall" letters to their customers. On 8/13/14, Zimmer contacted a distributor via phone pertaining to one additional lot (00-7843-011-06; lot 62547211) which had been included in the recall. Formal recall notices were issued to the distributor via email on 8/15/14. |  | 147 | 147 units | Worldwide Distribution-USA (nationwide) including the states of AL, AZ, CA, CO, CT, FL, GA, IL, IN, MA, MI, MN, NC, ND, NJ, NY, OH, OK, PA, TN, TX, VA &amp; WI, and the countries of Canada, Colombia, Dominican Republic, Japan, Korea (S. Korea), Peru, Taiwan &amp; Thailand. | Terminated | March 24, 2015 | 309 | 147 | Undetermined
Z-1732-2014 | 68350 | No | KOD | Urological catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | RUSCH Easy Cath Coude Kit, 10 FR, Rx Only, Sterile, Distributed by: Teleflex Medical, Research Triangle Park, NC 27708. |  |  | May 21, 2014 | 2014 | Teleflex Medical | Product was packaged with a straight catheter rather than the correct coude (curved) catheter. | Mixed-up of materials/components | Teleflex Medical sent an Urgent Medical Device Recall Notification dated May 21, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Customers were instructed to immediately discontinue use and quarantine any of the affected product in stock. They were to complete and return the enclosed Recall Acknowledgement Form via fax to Teleflex Medical. Upon receipt, they will be contacted by a Customer Service Representative with instructions for the return of the product to Teleflex Medical.Teleflex Medical is committed to providing high quality, safe and effective products. We sincerely apologize for any inconvenience this action may cause your operations. If you have any other questions, feel free to contact your local sales representative or Customer Service at 1-866-246-6990. |  | 1000 | 1000 eaches | US Distribution including the states of CA, FL, OH and OK. | Terminated | November 23, 2015 | 551 | 1000 | Undetermined
Z-0619-2014 | 67069 | No | JSS | Microorganism differentiation and identification device. | Microbiology | Microbiology | 510(K) Exempt | AdvanDx Gram-Negative QuickFISH BC. |  |  | December 11, 2013 | 2013 | AdvanDx, Inc. | Kits incompatible with certain fluorescence microscope light sources | Process design | AdvanDx issued an email notification along with a PDF file of the new package insert to each customer followed by written notification dated December 6, 2013 and a paper copy of the new package insert via Fed Ex courier. All customers received an acknowledgement form to fill out and return to AdvanDx so that the effectiveness of the replacement can be evaluated. All existing inventory of Gram-Negative QuickFISH BC will be repackaged with the new package inert Customers with questions were instructed to contact AdvanDx Technical Support at 1-866-376-0009 or techwupport@advandx.com.For questions regarding this recall call 781-376-0009. |  | 66 | 49 kits US; 17 kits OUS | Worldwide Distribution - USA including CA, IL, IN, OH, NJ, MI , MO, NY, WI. Internationally to Denmark. | Terminated | January 13, 2014 | 33 | 66 | Undetermined
Z-1740-2014 | 67998 | No | EBZ | Ultraviolet activator for polymerization. | Dental | Dental | 510(k) | bluephase style Light probe Pin-point 6&gt;2 mm black REF # 636 241. The accessory is a pin point light probe, an accessory to a dental curing light |  |  | April 14, 2014 | 2014 | Ivoclar Vivadent, Inc. | A stock check revealed broken glass fibers between the connector and the fiber bundle in a large number of these light probes. | Other | lvoclar Vivadent, Inc. sells the affected medical device through Dental Dealer partners (initial consignees) who in turn sell directly to the dentist's office. In this case, lvoclar Vivadent, Inc. handled direct drop shipment of the device to end user dental offices in 80% of the cases. Where lvoclar Vivadent hasdirect shipment information available, dentist offices will be notified directly.The Dental Dealer partners will be notified of this recall and asked to cooperate and provide name and address listings for all dental office customers where they have shipped the Pin Point probe. The dealer will also be given a customer communication to send directly to the dentists purchasing the device. |  | 241 | 241 units | Worldwide Distribution: US Nationwide and the country of Canada. | Terminated | November 05, 2015 | 570 | 241 | Undetermined
Z-1742-2014 | 68313 | No | NEC | Liquid bandage. | General Hospital | General &amp; Plastic Surgery | 510(K) Exempt | BioDerm Liberty 3.0 External Male Catheter Specialty Type with integral collection chamber (Component: SKIN-PREP Protective wipes). |  |  | May 31, 2011 | 2011 | BioDerm, Inc. | On 06/08/2011, BioDerm, Inc. issued a voluntary recall of a component skin wipes for the BioDerm Liberty 3.0 External male Catheter, BioDerm Liberty Pouch Clean &amp; Dry External Male Catheter and BioDerm Safe n' Dry. This component (wipes) were initially manufacturer by Smith &amp; Nephew and recalled due to a potential bacterial contamination. | Material/Component Contamination | Notice was sent to customers by letter dated June 8, 2011 and by phone. |  | 26040 | 26,040 units. | Worldwide Distribution-USA (nationwide) including the states of AZ, CA, FL, NH, NJ, NY, PA, TX, and WI, and the countries of Canada and Netherlands. | Terminated | June 20, 2014 | 1116 | 26040 | Undetermined
Z-1679-2014 | 68356 | No | DWJ | Thermal regulating system. | Cardiovascular | Cardiovascular | 510(k) | 115 V Blanketrol II, Model 222S Hyper-Hypothermia System . Used to lower or to raise a patients temperature and /or maintain a desired patient temperature through conductive heat transfer. Part number 86165. |  |  | January 08, 2014 | 2014 | Cincinnati Sub-Zero Products Inc | The incorrect voltage component was placed into the device during manufacturing. This may make the device unable to maintain patient temperature during cooling mode. The device will continue to indicate that the unit is cooling because the compressor will continue to run, however water in the reservoir will begin to return to ambient temperature. The actual water temperature will be displayed but | Mixed-up of materials/components | The firm initiated a telephone notification on 01/08/2014, followed by a written notification which was sent to the customers via letter or email on 01/31/2014.URGENT Medical Device Field Action (January 31, 2014)REF: 115V Blanketrol II Model 22S Hyper-Hypothermia System(1) Discontinue Use and contact CSZ(2)Upon receipt of new 115V (p/n 91113) solenoid coil(s), immediately remove and destroy/discard all the affected 220V (p/n 3600) solenoid coil(s)(3)After the affected 220V (p/n 36007) solenoid coil(s) have been removed, discard and replaced, please complete and return the enclosed response form as soon as possible to acknowledge receipt of this notification and to imform CSZ that you have performed and completed the requested actions. Return the form by fax |  | 7 | 7 units | The product was distributed to one customer in MI, and internationally (Taiwan) | Terminated | June 06, 2014 | 149 | 7 | Undetermined
Z-1747-2014 | 68279 | No | OCY | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | Boston Scientific, NovaGold High Performance Guidewire, Rx only, Sterilized using ethylene oxide, UPN M00552010 and M00552000, Manufactured by NeoMetrics Inc, 2605 Fernbrook Lane Suite J, Plymouth, MN 55447. |  |  | March 31, 2014 | 2014 | NEO METRICS, INC. | Neo Metrics, Inc. is recalling NovaGold Guidewire because the tip may uncoil and / detach more frequently than anticipated. There have been no reports of illness or injury. | Nonconforming Material/Component | NeoMetrics sent a letter dated March 31, 2014, to Boston Scientific. On March 24, 2014, Boston Scientific made NeoMetrics aware that Boston Scientific received 3 complaints from customers that the tip of the NovaGold guidewire detached and/or became uncoiled.Because the tips are breaking and/or uncoiling at a rate that was not anticipated, it has been decided that these units should be removed from the field.Therefore, NeoMetrics does not object to Boston Scientific executing this product removal in accordance with Section 4.4 of the distribution agreement.For further questions please call (763) 559-4440. |  | 2512 | 1256 boxes ( 2 wires/box) | Worldwide Distribution - US and the countries of Austria, Belgium, Germany, Spain, France, Great Britain, Italy, Netherlands and Sweden. | Terminated | January 16, 2015 | 291 | 2512 | Undetermined
Z-1762-2014 | 68331 | No | HBE | Powered simple cranial drills, burrs, trephines, and their accessories. | Neurology | Neurology | 510(k) | Stryker 3.5mm Zyphr Elite Round Fluted Bur |  |  | May 22, 2014 | 2014 | Stryker Instruments Div. of Stryker Corporation | There may be missing notches, partial notches or no notches at all in some Cutting Accessories that are used with the Elite" Attachments. | Nonconforming Material/Component | Stryker Instruments sent an URGENT MEDICAL DEVICE RECALL notification letter dated May 22, 2014, to the consignees. Recall notification included affected product description and identification, reason for recall, risk to health, and actions to be taken by the customers. Customers will be instructed to review the notification, quarantine unused product and to return a signed Business Reply Form. They will be advised that a Stryker Representative will be sent to their facility to inspect any affected product on hand. Once product is inspected by the Stryker Representative, they can return conforming product back to stock, for normal use. Non-conforming product will be removed from the facility by a Stryker Representative following the inspection process. The Stryker Representative will send the product to the manufacturer for destruction. For questions regarding this recall please contact Stryker Instruments:Kara Spath269-389-4518 kara.spath@stryker.com |  | 39 | 39 | Worldwide Distribution - US (nationwide) | Terminated | May 06, 2015 | 349 | 13739 | Undetermined
Z-1782-2014 | 68317 | No | LKF | Obstetric-gynecologic specialized manual instrument. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | Cooper Surgical TAMPA CATHETER 5 French 33 cm. |  |  | May 13, 2014 | 2014 | CooperSurgical, Inc. | Sterility of the device may be compromised due to unsealed pouch | Process control | Cooper Surgical Inc notified consignees by letter dated 5/13/14 sent via Federal Express with confirmed delivery receipt. Consignees are requested to return for refund or exchange. If you have any further questions contact the firm at 203.601.5200. |  | 1180 | 1180 units | Distributed USA (nationwide) including the states of CO, CT, MO, VT, NY, OH, NJ, FL, NC, VA, AL, MA, PA, and CA, and the country of Canada. | Terminated | May 28, 2015 | 380 | 1180 | Undetermined
Z-1784-2014 | 64273 | No | LOX | Percutaneous Transluminal Coronary Angioplasty (PTCA) Catheter. | Cardiovascular | Cardiovascular | 510(k) | Cordis EMPIRA RX PTCA Dilatation Catheter Catalog # 85R30300S For cardiovascular use. |  |  | February 05, 2013 | 2013 | Cordis Corporation | Dilatation catheters could exhibit radial versus axial tears should they burst during inflation. | Process design | An Urgent Medical Device Recall letter and Acknowledgement form was sent overnight to mulitple contacts in each account February 7, 2013. A representative will follow-up as necassary to facility obtaining signature, faxing the acknowledgment form to Cordis, collecting and returning units. |  | 80 | 80 units | Worldwide distribution: US states: AZ, FL, IL, LA, MA, NJ, OR, SC, and TX. Armenia, Austria, Belgium, Colombia, Czech Republic, France, Hungary, Iran, Israel, India, Italy, Kuwait, Latvia, Lebanon, Luxembourg, Mexico, Morocco, Poland, Portugal, Russia, Saudi Arabia, South Africa, Spain, Switzerland, and United Arab Emirates. | Terminated | June 23, 2014 | 503 | 80 | Undetermined
Z-1786-2014 | 68332 | No | GBZ | Introduction/drainage catheter and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Taut Operative Cholangiogram Catheter, 4.5FR (1.5mm) x 76cm |  |  | May 16, 2014 | 2014 | Teleflex Medical | Several complaints were received for the metal support tube missing from the cholangiogram catheter. | Employee error | Teleflex sent an Urgent Medical Device Recall Notification dated May 16, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Customers were Instructed:1. Immediately discontinue distribution and quarantine any products with the catalog numbers and lot numbers listed above.2. Using the provided customer letter and Recall Acknowledgement Form templates, communicate this recall to any of your customers who have received product included within the scope of the recall.3. Have the customers return any affected product to you, together with a completed Recall Acknowledgement Form, for consolidation and return to Teleflex Medical. In the event that an alternative approach is needed, contact Teleflex Medical Customer Service for more information at 1-866-246-6990.4. To return product, complete the enclosed Recall Acknowledgement Form and fax it to 1- 866-804-9881, Attn: Customer Service. This will allow us to document the amount of product you have on hand for return. A customer service representative will contact you with a Return Goods Authorization (RGA) Number and will provide instructions for the return of product to Teleflex Medical.5. If you and your customers have no affected stock, please complete the enclosed Recall Acknowledgment Form and fax it to 1-866-804-9881, Attn: Customer Service. This will allow us to document your receipt of this letterTeleflex Medical is committed to providing high quality, safe and effective products. We sincerely apologize for any inconvenience this action may cause your operations. If you have any other questions, feel free to contact your local sales representative or Customer Service at 1-866-246-6990. |  | 450 | 450 | Worldwide Distribution - USA including the states of AL, AZ, AR, CA, CO, FL, IL, LA, ME, MI, MO, NY, OH, OR, TN, TX, VT and WI., and the country of Belgium. | Terminated | November 04, 2015 | 537 | 450 | Undetermined
Z-1789-2014 | 68226 | Yes | NPF | Apnea monitor. | Anesthesiology | Anesthesiology | 510(k) | Philips/Children's Medical Ventures Smart Monitor 2 Professional Series infant apnea monitor. |  |  | April 29, 2014 | 2014 | Childrens Medical Ventures | The battery pack wire harness is improperly assembled rendering the device inoperable. The wiring issue causes an error code to be displayed and results in the device sounding an alarm repeatedly at power up. The presence of the error code and the devices alarm prevents the device from functioning until the issue is addressed. | Component design/selection | Philips sent an Urgent Medical Device Correction letter dated April 29, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.This Customer Information is intended to inform you about:1.what the problem is and under what circumstances it can occur2.the actions that you as a customer can take to minimize the effect of the problem3.the actions planned by Philips to correct the problem.If you need any further information or support concerning this issue, please contact your local Philips representative: For further questions please call (412) 380-8881.Philips apologizes for any inconveniences caused by this problem. |  | 112 | 112 | Worldwide Distribution - USA including the states of MI, WI, SC, FL, TX, KS, MA, AZ, UT, WV, IN, MN, NV and VA., and the countries of Australia, Belgium and Canada. | Terminated | April 06, 2015 | 342 | 112 | Undetermined
Z-1798-2014 | 68373 | Yes | OAD |  |  | Cardiovascular | PMA | IntellaTip MiFi XP Temperature Ablation Catheter; 10m x 8F (2.67mm) Large Curve. |  |  | May 15, 2014 | 2014 | Boston Scientific Corporation | Some units of Intella Tip MiFi XP Temperature Ablation Catheters were not manufactured according to specification. | Use error | Customers were informed of the recall via overnight letter sent on May 15, 2014. |  | 968 | 968 units total all models | Nationwide Distribution. | Terminated | June 25, 2015 | 406 | 968 | Undetermined
Z-1831-2014 | 67058 | detergent | JCB | General purpose reagent. | Pathology | Pathology | 510(K) Exempt | Coulter DxH Cleaner 10L, Part No. 628023. |  |  | November 08, 2013 | 2013 | Beckman Coulter Inc. | Beckman Coulter is recalling the UniCel DxH800 and UniCel DxH600 Coulter Analysis Systems because the optical degradation may potentially cause a delay in reporting results for nucleated red blood cells (nRBC), automated white blood cell differential and the reticulocyte. | Under Investigation by firm | An Urgent Product Correction letter, dated 11/08/2013, was sent to all the customers who purchased the UniCel DxH 800 and UniCel DxH 600 Coulter Cellular Analysis Systems. The letter informs the customers of the problems identified and the actions to be taken. Customers are instructed to complete and return the enclosed Response Form within 10 days. Customers with questions are instructed to contact Beckman Coulter's Customer Support Center via website, http://www.beckmancoulter.com/customersupport/support, via phone at (800) 526-7694 in US and Canada. For outside of the US and Canada, customers are instructed to contact their local Beckman Coulter Representative.Beckman Coulter sent an update on 2/14/14, to inform customers that the recall was expanded to include specific lots of DxH Cleaner. |  | 12924 | 12,924 units total (9,733 units in US) | Worldwide Distribution -- US, Canada, Australia, China, Belgium, Brunei Darssalam, Colombia, Croatia, Czech Republic, France, Germany, Gibraltar, Greece, Hong Kong, Hungary, India, Ireland, Italy, Japan, Korea, Kuwait, Macao, Malaysia, Mexico, Myanmar, Netherlands, New Zealand, Norway, Oman, Philippines, Poland, Panama, Portugal, Puerto Rico, Qatar, Romania, Russia Federation, Saudi Arabia, Singapore, Slovakia, Slovenia, Spain, Sweden, Switzerland, Taiwan, Thailand, Trinidad and Tobago, Turkey, United Arab Emirates, United Kingdom, and Vietnam. | Terminated | December 25, 2014 | 412 | 12924 | Undetermined
Z-1805-2014 | 68286 | No | DKB | Clinical toxicology calibrator. | Toxicology | Toxicology | 510(K) Exempt | VITROS (R) Chemistry Products FS Calibrator 1, Product Code/ Catalog Number 6801873 -- For in vitro diagnostic use only --- VITROS Chemistry Products FS Calibrator 1 is used in conjunction with VITROS Chemistry Products Calibrator Kits to calibrate the VITROS 5,1 FS and 4600 Chemistry Systems and the VITROS 5600 Integrated System. |  |  | May 12, 2014 | 2014 | Ortho-Clinical Diagnostics | Calibrator values assigned for the calibrator kit lots were released using a reference method that was undergoing validation for a new lot of European Reference Material (ERM). The validation process was not complete prior to calibrator value assignment for VITROS (R) Chemistry Products hsCRP Reagent (Product Code 6801739), Generation (GEN) 28. | Process change control | Customer Product Correction Letters (CL2014-143, dated 5/12/14) were sent on 5/12/2014 via FedEx overnight courier and/or US Postal Service Priority Mail (for PO Boxes only) and /or ORTHO PLUS e-Communications to US consignees (OCD Direct, Drop-Ship and US Federal Government) to inform them of the issue and advise them to discontinue use of the affected calibrator lot(s), and inform their Laboratory Medical Director if they suspect that a previously reported result may have been affected. Distributor letter (DL2014-143) was also sent on 5/12/2014 via US Postal Service Priority Mail and/or FedEx overnight courier to US Distributor consignees and instructed to notify their consignees of the issue and the required actions. Foreign affiliates were pre-notified via e-mail on 5/08/2014 that the communication of this issue and required actions is imminent and were sent the finalized recall communication on 5/12/2014. |  | 825 | Domestic - 464 units; Foreign - 361 units | Worldwide Distribution-USA (nationwide) including Puerto Rico and the countries of Australia, Brazil, Canada, Chile, China, Colombia, England, France, Germany, India, Italy, Japan, Mexico, Panama, Singapore, Spain, and Venezuela. | Terminated | March 23, 2017 | 1046 | 1898 | Undetermined
Z-1810-2014 | 67740 | No | HQG | Prescription spectacle lens. | Ophthalmic | Ophthalmic | 510(K) Exempt | Safety glasses made with "Rage" frames manufactured between February 20, 2014 and February 25, 2014. The color of the frames of the recalled glasses were either white or black. |  |  | February 26, 2014 | 2014 | Wal Mart Stores, Inc | Fifteen pairs of prescription, safety glasses did not meet required, industrial lens thickness; nor did these glasses have the OSHA-required manufacturer markings. | Employee error | Wal-Mart sent an Urgent Medical Device Recall letter dated March 12, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Actions to be taken by the Customer/User:Do not use the eyewear in any situation where safety-impact lenses are required. Replacement eyewear has been manufactured and available for the patient to pick up at the vision center where the original eyewear was purchased.Type of Action by the Company:Walmart contacted all patients on February 27 and 28. Patients were notified that their eyewear was not made at safety thickness, and that the patient should not wear the eyewear in an environment where safety thickness lenses are needed. Walmart manufactured replacement orders for all affected patients.Should you have any questions regarding this notice please call (479)277-0720 |  | 15 | 15 safety glasses | US Distribution including the states of AR, AL, CA, FL, KS, LA, MN, NJ, NV,OH, TX, and SC. | Terminated | January 08, 2015 | 316 | 15 | Undetermined
Z-1811-2014 | 68526 | No | HXX | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Biomet Spine, Cypher MIS Screw Inserter, Catalog number 14-501669, instrument for Spinal Screw Fixation System. |  |  | May 15, 2014 | 2014 | Ebi, Llc | The Cypher MIS Screw Inserter may exhibit an increased rate of instrument tip failure. | Nonconforming Material/Component | Biomet sent an "Urgent Medical Device Recall Notice"/Fax Back Response Form dated May 15, 2014, via Fed Ex to all affected customers. The notice identified the product, the problem, and the action to be taken by the customer. Customers were instructed to inspect their inventory for the affected product and if located to remove the product from circulation; follow instructions on the attached Fax Back Response Form prior to return of the product; and return the product using the Fed Ex label included with the notification. Customers were instructed, if if they further distributed the product, they must notify their customers of the recall via the enclosed "Dear O.R. Manger" notice. If the customers had any questions they were asked to please contact Biomet, Monday through Friday (8 am to 5 pm) at 973.299.9300 xt. 2322. |  | 32 | 32 units (30 domestic and 2 international) | Worldwide Distribution - USA (nationwide) and the country The Netherlands. | Open, Classified |  |  | 32 | Undetermined
Z-1809-2014 | 68454 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Model Number(s): PB840 Ventilator (4-840120DIUU-XX). The product is a component of the Puritan Bennett 840 Ventilator: The Backlight Inverter Printed Circuit Board (BLI PCBA) for the 9.4 Display part number 4-079056-00. |  |  | June 04, 2014 | 2014 | Nellcor Puritan Bennett Inc. (dba Covidien LP) | In the case of a loss of GUI display due to a Backlight Inverter PCBA failure, the ventilator continues to provide uninterrupted ventilatory support at the programmed settings for the patient. However, there is a loss of display and thus there is a necessity to move the patient to another ventilator. | Component design/selection | Covidien will be issuing the Product Notification letter in 2014. Customers are requested to forward the letter to anyone they may have transferred or assigned for use one of the 9.4 display configuration ventilators. Service engineers will be replacing the BLI PCBA with a new design containing the 1.0 ampere fuse and then destroying the old PCBA. Ventilators replaced with newer models under the commercial program will be returned to Covidien for disposal.Should you have any questions regarding this letter or to report any issues with the Puritan Bennett" 840 ventilator contact our Technical Support Department at 1-800-255-6774, option 4, then option 1, to ensure proper device reporting procedures are followed. Thank you for your prompt attention to this matter. |  | 7371 | 7371 | Worldwide Distribution - USA (nationwide) including the states: AK, AL, AR, AZ, CA, CO, CT, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, TN, TX, UT, VA, VT, WA, WI, WV, WY and District of Columbia., and the countries of Canada and Japan. | Open, Classified |  |  | 7371 | Undetermined
Z-1808-2014 | 68100 | No | CCL | Oxygen gas analyzer. | Anesthesiology | Anesthesiology | 510(k) | GE Healthcare, Compact Airway Modules E-CO, E-COV, E-COVX, E-CAiO, E-CAiOV and E-CAiOVX. The Datex-Ohmeda S/5MT Compact Airway Module, E-CAiOVX family is used for monitoring hospital patients respiration, ventilation, and gas exchange status. |  |  | May 23, 2014 | 2014 | GE Healthcare, LLC | Potential loss of airway gas measurement in the Compact Airway Gas Modules. Loosening of the gas sampling pump hex screw in the module pump unit may lead to pump failure that may cause loss of respiratory airway gas parameter monitoring. If this issue occurs, the monitor will show an alarm note on monitor screen of either "Low gas sample flow" or "Check sample gas out". | Nonconforming Material/Component | GE Healthcare sent an "Urgent Medical Device Correction" letter dated May 23, 1014. The letter was addressed to Healthcare Administrator / Risk Manger, Chief of Nursing and Director of Biomedical Engineering. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction and Contact Information. For questions they should contact Technical support at 1-800-558-7044 or your local Service Representative. |  | 4961 | 4,961 (1,946 units US, 3,015 units OUS) | Worldwide Distribution - US (nationwide): All states in continental USA including DC except DE, ME, ND, RI &amp; VT. and the countries of: ARGENTINA, AUSTRALIA, AUSTRIA, BAHAMAS, BAHRAIN, BELGIUM, BOSNIA AND HERZEGOVINA, BRAZIL, CANADA, CHILE, CHINA, COLOMBIA, CZECH REPUBLIC, DENMARK, ECUADOR, EGYPT ESTONIA, FINLAND, FRANCE, GERMANY, GREECE, GHANA, HONG KONG, HUNGARY, INDIA, IRAQ, IRELAND, ISRAEL, ITALY, JAPAN, JORDAN, KAZAHKSTAN, KENYA, REPUBLIC OF KOREA, KUWAIT LEBANON, LITHUANIA, MACEDONIA, MAURITIUS,MEXICO, MOROCCO, NAMIBIA, NETHERLANDS, NEW ZEALAND, NORWAY,, PANAMA, PERU, POLAND, PORTUGAL, QATAR, RUSSIAN FEDERATION, SAUDI ARABIA, SERBIA, SINGAPORE, SLOVAKIA, SLOVENIA, SOUTH AFRICA, SOUTH KOREA, SPAIN, SWEDEN, SWITZERLAND, SYRIAN ARAB REPUBLIC, TAIWAN, THAILAND, TUNISIA, TURKEY, UNITED ARAB EMIRATES, UNITED KINGDOM, VENEZUELA, VIET NAM, and URUGUAY. | Terminated | October 21, 2015 | 516 | 4961 | Undetermined
Z-1824-2014 | 68374 | No | DQA | Oximeter. | Cardiovascular | Anesthesiology | 510(k) | Spacelabs Healthcare Ventilator Flexport Interface, Model 90436A-07 |  |  | May 09, 2014 | 2014 | Spacelabs Healthcare Inc | Spacelabs Healthcare is voluntarily recalling the Hamilton Galileo Ventilator Flexport, Model 90436A-07, where the monitored Minute Volumes (Vmin) has been reported at one time to reach ten times the actual value on the bedside monitor. | Component change control | Spacelabs changed their strategy for updating the software and sent the updated letter Urgent-Medical Device Correction- Hamilton Galileo Ventilator FlexportModel 90436A-07, dated 10 June 2015, to their consignees. The firm is going to send consignees new replacement flexport(s) with upgraded software and requested consignees to return old affected units via Return Goods Authorization (RGA) to Spacelabs and the firm will destroy affected units.Consignees are instructed to take the following actions:1. Please replace the affected devices and test for correct operation.2. Return your old device(s) using the RGA number and prepaid FedX shipping label.Use the Reference code: |  | 154 | 46 units distributed within the US; 108 units distributed outside the US. | Worldwide Distribution-USA (nationwide) including the states of AZ, IN, LA, GA, SC, FL, and NY, and 18 customers outside the US. | Terminated | November 16, 2015 | 556 | 154 | Undetermined
Z-1837-2014 | 68310 | No | PBI | Hip joint metal/polymer constrained cemented or uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | G7 Acetabular System Positioning Guide Rod, orthopedic surgical instrument for hip prosthesis PN 110003458. |  |  | May 07, 2014 | 2014 | Biomet, Inc. | Guide rod may fracture during impaction of the acetabular shell and be retained in the inserter or the patient and could result in a delay in surgery. | Use error | The firm issued the following notice on May 7, 2014: URGENT MEDICAL DEVICE RECALL NOTICE:This notification is to inform you of an Urgent Medical Device Recall initiated by Biomet Orthopedics ("Biomet") which involves Part Numbers: 110003458 G7 Positioning Guide Rod and 110003500 G7 Positioning Guide Post. These instruments have been consigned and/or invoiced to your account. Biomet has initiated this action following an investigation which identified that positioning guide rod may fracture during impaction of the acetabula shell. If the positioning guide rod fractures, then the positioning guide post may become stuck to the insertion handle. There is no adverse health outcome expected for the patient.This action requires the immediate location and discontinued use of the product and its return to Biomet.The following actions are REQUIRED:Immediately locate and remove the identified device(s) listed below from circulation.Carefully follow the instructions on the enclosed "FAX Back Response Form .Fax a copy of the Response Form to 574-372-1683 prior to return of product.Use priority carrier for your shipment.If you have further distributed this product, you MUST notify hospital personnel of this action via the enclosed "Dear Risk/Recall Manager" notice. This letter MUST be given to hospital personnel responsible for receiving recall notices. However, you are charged with the location and return of these products. |  | 1047 | 1047 total | Worldwide Distribution-USA (nationwide) and the countries of Canada, Netherlands, New Zealand, and Japan. | Terminated | October 30, 2015 | 541 | 1047 | Undetermined
Z-1834-2014 | 65274 | No | FOZ | Intravascular catheter. | General Hospital | General Hospital | 510(k) | NIPRO Safe Touch Safety Scalp Vein Set |  |  | May 04, 2013 | 2013 | Nipro Medical Corporation | There is a possibility of a crack in the CP luer connector. The position of the crack is at the connection site of the CP luer connector and the tubing. The crack could cause or contribute to leakage of fluids. | Component design/selection | Nipro sent a Recall Notification letter dated April 4, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.This recall does extend to the customer level. Please notify all customers of the recall and collect all affected product that is in their possession. When all products have been received from your customers call tags and RMAs will be issued in order for you to return all products to Nipro Medical warehouse in Memphis TN. You will then be reimbursed for all products received.Please place all products you have in stock on hold. Use the attached faxback to respond to this notification, with product codes, lots and quantities you have in all of your distribution facilities. One form should be completed for each distribution location. Return forms as soon as possible to Nipro Medical, fax: 305.592.4621. Any questions or concerns can be addressed by phone at 305.599.7174 ext. 249. |  | 176000 | 176,000 pieces | Worldwide Distribution - USA and the countries of Canada, Bolivia, Chile, Uruguay, Belize, Dominica, Honduras, Jamaica, Paraguay and Venezuela. | Terminated | March 02, 2015 | 667 | 176000 | Undetermined
Z-1844-2014 | 68444 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Puritan Bennett 840 Ventilator (4-840120DIUU-XX) |  |  | June 03, 2014 | 2014 | Nellcor Puritan Bennett Inc. (dba Covidien LP) | Covidien is conducting a medical device field correction for specific Puritan Bennett 840 Ventilators due to customer reports of loss of graphical user interface (GUI) display information while the ventilator continues to provide breath support. | Nonconforming Material/Component | Covidien sent an Urgent Medical Device Voluntary Field Correction letter dated 6/13/2014, to all affected customers. The letter identified the product the problem and and the action needed to be taken by the customer. Customers are requested to forward the recall letter to anyone they may have transferred or assigned for use one of the potentially affected ventilators. Service professionals will be identifying the affected PSUs and replacing them. The removed PSUs will be destroyed.To verify whether your devices are affected by this field safety corrective action and to initiate the process to receive the actions mentioned above, please complete the attached verification form and fax it to the Covidien contact indicated on the form. If you need further assistance, contact our Technical Support Department @ 1-800-255-6774, option 4, then option 1. |  | 104 | 104 | Worldwide Distribution - USA (nationwide) including Puerto Rico and the states of AK, AL, AR, AZ, CA, CO, CT, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, MI, MO, MS, NC, ND, NE, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV and WY., and the countries of Argentina, Australia, Belgium, Brazil, Canada, Chile, Colombia, Costa Rica, Egypt, Germany, Guatemala, Israel, Japan, Kazakhstan, Kenya, Lebanon, Libya, Macedonia, Mexico, Montenegro, Morocco, Panama, Peru, Russia, Saudi Arabia, Singapore, Switzerland, Taiwan, Tunisia, Turkey, Yugoslavia (Serbia &amp; Montenegro) Including GOVT/VA/Military | Open, Classified |  |  | 104 | Undetermined
Z-1956-2014 | 68564 | No | OVN | Absorbable surgical gut suture. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | Enforcement Discretion | McKesson Medi-Pak, Laceration Trays with Cotton O.R. Towels, Sterile, Single Use, Latex Free, 20 Each per Case, Rx only, Distributed by: McKesson Medical Surgical, Inc., Richmond, VA 23228 |  |  | June 12, 2014 | 2014 | Medical Action Industries Inc | Some of the sterile kits may have incomplete seals which might compromise the sterility of the kits. | Nonconforming Material/Component | Consignees were notified via E-Mails, Certified Letters and Telephone on 06/12/2014. |  | 474 | 474 cs (9,480 ea) total for all kits involved | Distributed in the states of AR, AZ, FL, GA, IA, IL, KY, MA, MN, NC, NY, OH, TN, TX, and VA. | Terminated | February 03, 2015 | 236 | 474 | Undetermined
Z-1977-2014 | 68490 | No | OYE | Automated differential cell counter. | Hematology | Hematology | 510(k) | BD FACS Canto and associated Fluidics Cart |  |  | June 06, 2014 | 2014 | BD Biosciences, Systems &amp; Reagents | Laser cooling fans and internal fan of the associated fluidics cart are not operational. | Nonconforming Material/Component | BD sent an Important Product Recall notification letters dated June 2014 to affected US customers by trackable mail. International affiliates were notified on June 6, 2014. The letter identified the affected product, problem and the actions to be taken. A Field Change notification has been established and a BD field Service Engineer will coordinate with each account to perform the necessary corrections. Questions or concerns may be directed to BD Customer Support at BD Biosciences, prompt 2. Out of US customers are to contact local BD rep with questions. |  | 35 | 35 installed devices | Worldwide Distribution - US Nationwide in the states of CA, OH, RI, GA, FL, MD, NY, WV, NJ, Ma, CT, IA, NM, MN and the countries of China and Belgium. | Terminated | January 27, 2015 | 235 | 35 | Undetermined
Z-1976-2014 | 68519 | No | HPP | Operating headlamp. | Ophthalmic | Ophthalmic | 510(K) Exempt | IntegraÂż LED Battery Charger, Single Bay. |  |  | June 11, 2014 | 2014 | Integra LifeSciences Corp. | LED battery chargers may prematurely fail and will not charge the LED battery as intended | Manufacturing material removal | Integra initiated a voluntary recall on June 11, 2014 to consignees that have been shipped affected product lots of IntegraÂż LED Battery Chargers, Catalogue No. 90523.The written notifications were delivered to US consignees by traceable courier service and non-US consignees by traceable email or facsimile. The consignee notification: Advises them of the nature of the issue and the potential for an adverse patientconsequence. Advises them to review their inventory and how to determine if they have affected battery chargers. Advises them to immediately stop the use or distribution of any affected product lot numbers they may have. Requests they identify and report to Integra if they do or do not have any unexpired affected product lots.Advises them replacements battery chargers will be shipped for any affected products they affirm are in their possession and how to return their affectedproduct. Questions regarding these instructions, please contact Integra Customer Service at 1-800-431-1123 or your Integra sales |  | 1250 | 1250 battery chargers | Worldwide Distribution - US Nationwide and the countries of Canada, Australia, Arab Emirates, Denmark, Germany, Great Britain, India, Ireland, Israel, Japan, Kuwait, Lebanon, Malta., Netherlands, Panama, Poland, Saudi Arabia, Singapore, South Africa, Spain, Sweden, and Tunisia. | Terminated | December 19, 2014 | 191 | 1250 | Undetermined
Z-2081-2014 | 68567 | No | IZL | Mobile x-ray system. | Radiology | Radiology | 510(k) | AGFA Digital Radiography X-Ray system DX-D 100 |  |  | June 06, 2014 | 2014 | AGFA Corp. | When liquid comes in contact with the DX-D 100 touch screen, the device may incorrectly recognize this as user input altering device settings. | Use error | Agfa Healthcare sent an "URGENT FIELD SAFETY NOTICE" dated June 6, 2014, to the US and Canadian customers. The letter described the safety alert and mitigation. Acknowledgment, via FAX-Back or email, that the information was received and understood was requested from the customers.Although medical staff are required to disinfect hands on a frequent basis, users must ensure that their hands are dry before using the touch screen of the DX-0100 because liquid residue may activate the action buttons on the touch screen. Customers were instructed:"Do not operate touch screen monitor with wet hands!"Do not let liquids come in contact with the touch screen while the DXD 100 is powered on!"Always double check your parameter settings prior to exposing the patient. We would like to remind you of this specific warning statement in the DX-0100 User Manual.Please complete the feedback form as soon as possible and return it to us by June 16, 2014.Should the above information not apply to your facility or should the device have been transferred to another organization, please be so kind as to indicate this on the attached feedback form and pass this Urgent Field Safety Notice to the organization where the device has been transferred.We thank you for your careful attention to this issue and your continued support.If you have any questions about this matter, please contact your local Agfa Healthcare organization: 1-877-777-2432, prompt 1 and reference PR1406020002 |  | 216 | 216 units | USA (nationwide) including the states of AL, CA, FL, GA, IL, IN, IA, KY, LA, MA, MI, MN, MS, MO, NE, NJ, NY, NC, OH, SC, TN, TX, VA, WV, WI, WY and Internationally to Canada. | Open, Classified |  |  | 216 | Undetermined
Z-2058-2014 | 68595 | No | CIZ | Androstenedione test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Coat-A-Count Direct Androstenedione, SMN 10381049, Catalog Number TKAN1. Siemens Healthcare Diagnostics |  |  | June 09, 2014 | 2014 | Siemens Healthcare Diagnostics Inc. | The assay is demonstrating a higher frequency of results greater than 10 ng/mL compared to alternate platforms. | Under Investigation by firm | A recall letter dated 6/9/14 was sent to all customers who purchased the Coat-A-Count Direct Androstenedione because the assay is demonstrating a higher frequency of results compared to alternate platforms. The letter informs the customers of the problems identified, risks to health, and actions to be taken. Customers are instructed to complete and return the Field Correction Effectiveness Check attached to the recall letter within 30 days. |  | 1164 | 1,164 units | Worldwide Distribution. US nationwide including CA, MI, NC, NE, NY, OH, PA, TX, and VA; Argentina, Austria, Belgium, Bolivia, Canada, Chile, Colombia, Finland, Germany, Great Britain, Greece, India, Ireland, Italy, Japan, Korea, Netherlands, Paraguay, Peru, South Africa, Spain, Sweden, Switzerland, and Uruguay | Terminated | July 24, 2015 | 410 | 1164 | Undetermined
Z-2067-2014 | 68781 | No | NKB | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | Excella II Standard Pedicle Screw, Catalog No. E2S47525 and E2S47535, 4.75 mm x 25 mm. For spinal fixation. Packaged in heat-sealed plastic bag. Also provided in reusable steam sterilization tray as implant set. |  |  | July 07, 2014 | 2014 | Innovasis, Inc | Innovasis is recalling the Excella II Standard Pedicle Screw due to mislabeling of the size and catalog number. | Process control | Distributors were notified by e-mail on 7/7/14, with attached copies of the written letter and forms to be printed, signed and returned to Innovasis. Distributors have been instructed to return all inventory in stock. |  | 32 | 32 | Distribution US nationwide, including TX, CA, KS, MN, and OK. Foreign distribution to Japan. | Terminated | August 11, 2014 | 35 | 32 | Undetermined
Z-2070-2014 | 68573 | No | LHI | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Excelsior Medical Sterile Triple Lead Tubing Set and Sterile Heavy Duty Triple Lead Tubing Set. To be used with the Pharm-Assist Dispensing Pump. Used for drug reconstitution and/or fluid transfer |  |  | June 05, 2014 | 2014 | Excelsior Medical Corp | There is a potential for a leak. A molding defect in the connector that joins the three leads to the main flow line may allow fluid to escape or air to enter the fluid path during transfer of fluids from one container to another. | Nonconforming Material/Component | Excelsior Medical issued "Urgent Medical Device Recall" notifications/"Customer Reply Forms" dated June 5, 2014 to affected customers via certified mail (return receipt request). The notice informed customers of the issue with the affected product; how to identify affected product; instructed the customers to discontinue use of tubing sets from the affected lots and return the attached form (Customer Reply Form) via fax to Excelsior to make arrangements for return of product. If product was further distributed by the customer the notification instructs them to identify the customer(s) that it was supplied to and notify them at once of the product recall. A phone number was provided in case customers had questions regarding the recall (1.800.487.4276 8:30 am - 5:00 pm EST) as well as an email address (recall12@excelsiormedical.com). |  | 331500 | 4250 tubing sets (PA03 - 78 cases; BX03 - 92 cases (each case contains 25 tubing sets)) | Worldwide Distribution - USA (nationwide) and Canada. | Terminated | August 19, 2015 | 440 | 331500 | Undetermined
Z-2077-2014 | 68611 | No | CBR | Nitrous oxide gas analyzer. | Anesthesiology | Anesthesiology | 510(k) | Spacelabs Healthcare Ultraview SL Multigas Module, Model 92518 (Software V1.00.00-09). Provides a means of monitoring a variety of gas concentrations and alert clinical personnel when the concentration of anesthetic agent(s), oxygen, carbon dioxide or nitrous oxide moves outside of user-defined limits. |  |  | June 18, 2014 | 2014 | Spacelabs Healthcare Inc | With this software version, the Minimum Alveolar Concentration (MAC) value displayed is extremely high (20.1) and physiologically invalid. | Process control | Spacelabs sent an Urgent Medical Device Correction letter dated June 20, 20/14 to all affected customers. The letter identified the affected product, problem and the actions to be taken. Customers are informed that Spacelabs Field Service personnel will be contacting their facility to schedule a convenient time to confirm, at no cost, a software update is needed to resolve this issue. Customers with questions are instructed to contact SpacelabsHealthcare at 1-800-522-7025 and select 2 for Technical Support. |  | 8 | 8 units total (3 units in US) | Worldwide Distribution - US in the state of Georgia and in the countries of : Malaysia, Paraguay, and France. | Terminated | December 08, 2014 | 173 | 8 | Undetermined
Z-2078-2014 | 68609 | No | EFA | Dental handpiece and accessories. | Dental | Dental | 510(k) | Titan T 5K Motor |  |  | June 13, 2014 | 2014 | Dental EZ Group Star Dental Division | The housing of 5K motor was laser marked incorrectly as a 20K motor.The recall was expanded in September 2014 to include additional lot #2251181. | Process control | Star Dental sent a Medical Device Recall letter, dated June 13, 2014 to affected customers. This letter identified the affected product, problem, and actions to be taken. For question call 866-383-4636 extension 4350.The recall was expanded in September 2014 to include additional lot #2251181. A letter, dated September 2014 was sent to customers who were impacted by the recall expansion. |  | 45 | 45 | US Distribution in the states of : IN, TN, WA, CA, UT, MN, and NY. | Terminated | February 24, 2015 | 256 | 45 | Undetermined
Z-2076-2014 | 68619 | No | JDQ | Bone fixation cerclage. | Orthopedic | Orthopedic | 510(k) | Tritium Sternal Plate System, Screw, 02.7mm X 12mm, 12 Pack, Sterile. |  |  | June 11, 2014 | 2014 | PIONEER SURGICAL TECHNOLOGY, INC. | Lack of Sterility Assurance | Employee error | On 6/11/14 and 6/12/14 the firm contacted their two consignees and surgeon via phone and explained that the product did not go through the sterilization process. The returns authorization number and shipping labels were discussed for the unused unit. On 6/18/14 and 6/20/14 an URGENT: MEDICAL DEVICE VOLUNTARY RECALL NOTIFICATION was sent to the customers as well as the one surgeon. The letter explained the sterility concerns and that the productis being recalled. Contact Dan Nelson, Manager of Quality, at 906-226-4489 if you have any questions or concerns. |  | 5 | 5 | US Nationwide Distribution in the states of TX and FL | Terminated | October 22, 2014 | 133 | 5 | Undetermined
Z-2079-2014 | 68437 | No | KHP | Lactic acid test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | RAPIDPoint 500 Measurement Clinical Chemistry System |  |  | May 23, 2014 | 2014 | Siemens Healthcare Diagnostics Inc | RAPIDPoint 500 Measurement Cartridges may have an error code one or more electrolyte parameters (e.g., Na+, Ca++, K+, Cl-) | Process design | Siemens Healthcare Diagnostics issued an Urgent Field Safety Notice by e-mail to all affected Siemens Healthcare Diagnostics Regional offices on May 27, 2014 both in the United States and elsewhere for communication with affected customers. This notice informs customers of the issue with the RAPIDPoint 500 Measurement cartridges. Questions contact: Siemens Customer Care Center or your local Siemens technical support representative at 877-229-3711. What steps can be taken to mitigate the issue:" Install a new RAPIDPoint 500 Measurement Cartridge.or" Install a 405 Measurement Cartridge - The RAPIDPoint 405 cartridge is available in four sizes:Since the Lactate parameter is not available on the RAPIDPoint 405 cartridge, however, you will need to arrange to get your lactate results from other instruments in your central laboratory.or" Use an alternative system to obtain your electrolyte values . |  | 20507 | 20,507 cartridges | NationwideForeign:U.A.E.AlbaniaAustriaAustraliaBosnia Herzeg.BelgiumBulgariaBahrainBrazilCanadaSwitzerlandP.R. ChinaColombiaCzech RepublicGermanyDenmarkAlgeriaEstoniaEgyptSpainFinlandFranceUnited KingdomGeorgiaGreeceHong KongCroatiaHungaryIndonesiaIrelandIsraelIndiaItalyJapanRepublic KoreaKuwaitLithuaniaLuxembourgLatviaMalaysiaMexicoNetherlandsNorwayNew ZealandPeruPolandPortugalQatarRomaniaSerbiaRussian Fed.Saudi ArabiaSwedenSingaporeSloveniaTurkeyTaiwan | Terminated | June 10, 2015 | 383 | 20507 | Undetermined
Z-2080-2014 | 68624 | No | LJH | Urological catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Disposable StrykeFlow Tip, Stryker Disposable Suction/I irrigator Tip; Sterile, Rx only, Latex Free. Stryker Suction Irrigation system, used to provide a clean and clear operative site through the evacuation and aspiration of irrigant solution. for endoscopic use. |  |  | June 20, 2014 | 2014 | Stryker Endoscopy | Internal review found that six lots of Disposable StrykeFlow Tips were shipped with a potentially deformed blister pouch which could impact sterility of the product. | Other | Stryker sent an Urgent Device Removal letter via certified mail on June 20, 2014 to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to return any unused product, inform individuals within their organization who need to aware of the device removal. Customers with questions were instructed to call 1-800-624-4422.For questions regarding this recall call 408-754-2664. |  | 496 | 496 boxes (6 tips per box) | Worldwide Distribution - USA (nationwide) and Internationally tAustralia, Canada, France, Greece, Netherlands, United kingdom and South Korea. | Terminated | October 15, 2014 | 117 | 496 | Undetermined
Z-2087-2014 | 68442 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Synthes Guide Blocks for the 2 Column Plate 6 Hole Head (Volar Distal Radius Plate System). Orthopedic surgical instrument. |  |  | May 29, 2014 | 2014 | Synthes, Inc. | The Guide Blocks for the 2 Column Plate 6 Hole Head (Volar Distal Radius Plate System) were found to be labeled incorrectly. Part no. 03.111.600 (Right) was found in the package labeled part no. 03.111.601 (Left). The laser etching denoting the orientation on the part is correct. | Process control | An urgent field correction notification, dated May 29, 2014, was sent to end users and sales consultants which described the product, problem, and actions to be taken. |  | 180 | 180 | Distribution US Nationwide and Bermuda. | Open, Classified |  |  | 180 | Undetermined
Z-2086-2014 | 68341 | No | LZO | Hip joint metal/ceramic/polymer semi-constrained cemented or nonporous uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | NOVATION Femoral Stem, Tapered Extended offset, plasma- coated, press fit. Catalog number 160-01-13 |  |  | March 07, 2014 | 2014 | Exactech, Inc. | The femoral stem is lacking specified plasma coating. | Component design/selection | Exactech sent an Important Product Recall Notice dated March 10, 2014, to all affected customers. The notice identified the product the problem and the action to be taken by the customer. Description of Issue: The femoral stem is lacking specified plasma spray coating.Clinical Impact: Without plasma spray coating, the stem may not remain stable and will not have the possibility for biological fixation.In order to comply with applicable regulations and Exactech policies customers were instructed to :" Immediately cease distribution or use of these products." Extend this information to your accounts that may have this product in their possession." Determine if they have any of the affected devices " Fax back the attached form. Complete the attached fax back form and return it to Exactech.Customers were instructed to contact their Exactech inventory representative within 5 business days to confirm quantities at their location. For any inventory restocking questions related to this issue, please call 1-800-392-2832.For questions regarding this recall call 800-392-2832. |  | 10 | 10 | Worldwide Distribution including Spain, Greece and Japan | Terminated | August 04, 2014 | 150 | 10 | Undetermined
Z-1978-2014 | 68500 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Stryker Orthopaedics Triathlon Tritanium Patella Inserter, . NON-STERILE |  |  | May 01, 2014 | 2014 | Stryker Howmedica Osteonics Corp. | Tritanium Patella Inserter Instrument fracture during implantation. . | Process control | Stryker Orthopaedics issued an "Urgent Medical Device Recall Notification" letters/Notification Acknowledgement forms dated May 7, 2014 to their customers via Fed-Ex on May 7, 2014. The notification informs the customers of the issue(s) with the affected product; potential hazards; and how to identify and handle any affected product. Customers were instructed to return any affected product to Stryker Orthopaedics and fill out and return the enclosed Notification Acknowledgement form. For questions call (1-201-972-2100). |  | 22 | 22 Units | US Nationwide Distribution - in the states of Arizona, Florida, Kentucky, Louisiana and New Jersey. | Terminated | June 09, 2017 | 1135 | 22 | Undetermined
Z-2089-2014 | 68795 | No | MAV | Angiographic injector and syringe. | Cardiovascular | Cardiovascular | 510(k) | Perouse Flamingo, Model 0218TA and 0218TB intended for use during cardiovascular procedure to create, maintain and monitor pressure in the balloon catheter. |  |  | May 10, 2012 | 2012 | Perouse Medical | Perouse Medical in France initiated a voluntary recall of certain batches of Perouse FLAMINGO Inflations Devices due to a potential water-tightness default. The potential default could cause an air entry when the plunger is positioned around 20 mL. | Process control | Perouse Medical initiated this recall by sending an e-mail notification to consignees on May 10, 2012. A formal written recall notification letter was sent to consignees on May 16, 2012. The recall letter, written in English, on company's letterhead, dated May 16, 2012, titled "VOLUNTARY BATCH RECALL Of FLAMINGO Inflations Devices", notified the consignee of the following: product information with codes and batch codes, product description, description of the potential default, reason for the recall action, instructions to the consignee, contact information and assistance. Response form "INVENTORY FORM". |  | 2496 | 2,496 units | US. | Terminated | September 09, 2014 | 852 | 2496 | Undetermined
Z-2091-2014 | 68622 | No | ITI | Powered wheelchair. | Physical Medicine | Physical Medicine | 510(k) | MK5 NX Controller--Wheelchair controller |  |  | June 24, 2014 | 2014 | Invacare Corporation | M41 Powered Wheelchairs received through Invacare Service Parts (ISP) have controllers that were programmed in a way this is not aligned to standard factory settings. The controller during different settings such as max forward speed, max forward acceleration, max reverse speed and max turning speed can result in momentary loss of control. | Process control | The controllers will be reporogrammed in the field upon installationInvacare Sent out a Technical Services Bulletin (Oct. 15, 2012) . Please call 800-832-4707, Monday-Friday, 8:00 am- 7:00 pm EST, for assistance. If you are aware of a consumer complaint regarding this controller, please contact the Invacare Customer Service Complaints Department at 877-374-8919, Monday-Friday, 8:00am- 7:00pm or cscomplaints@invacare.com |  | 819 | 819 | Nationwide Distriution: CA, SC, IL, UT, TN, KS, HI, NJ, LA, TX,FL, IN, AL, AZ, MN, WI, MN, MI, AR,MA, GA, MS, NJ, NM, OK, ME, PA, NY, NC, OK, MO, WA, IA, OR,KY, MD, NE, ID, VA, WA, ND, MT, OH, NH, CT. Also shipped to Canada | Terminated | October 15, 2014 | 113 | 819 | Undetermined
Z-2092-2014 | 68667 | No | NKB | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | Solanas 2.1mm FIXED DEPTH DRILL, 10mm, Part No. 63928-10. Component of the SOLANAS III Posterior Stabilization System, a spinal fixation system intended to improve stability of the cervical and thoracolumbar area of the spine (C1-T3). |  |  | June 26, 2014 | 2014 | Alphatec Spine, Inc. | Alphatec Spine is recalling the 2.1mm Fixed Depth Drill 10mm because they received a report that the instrument malfunctioned and golden stop slid from its fixed position during use. | Unknown/Undetermined by firm | Alphatec Spine, Inc. sent an Urgent Medical Device Notification letter dated June 20, 2014, June 26, 2014, to all affected customers. Alphatec Spine received a report of the 10mm fixed depth drill from this lot which malfunctioned and golden stop slid from its fixed position during use. The letter informs the customers of the problems identifed and the actions to be taken. Customers are instructed to confirm that they received the notification letter by completing the customer response form. Completed forms are to be sent to Alphatec Spine at 5818 El Camino Rd, Carlsbad, CA 92008. Customers may also return forms by fax at (760) 431-0289. Customers with questions were instructed to contact Alphatec Spine's Customer Service Department on how to return the affected product.For questions regarding this recall call 760-431-9286. |  | 14 | 14 units total (12 units in US) | Nationwide Distribution including CT, NC, AL, WV, IL, AR, and NY. | Terminated | September 03, 2014 | 69 | 14 | Undetermined
Z-2097-2014 | 68404 | No | DWF | Cardiopulmonary bypass vascular catheter, cannula, or tubing. | Cardiovascular | Cardiovascular | 510(k) | Sarns Malleable Venous Return Cannulae 24, 28, 30, 32, 34, 36 &amp; 40 Fr single stage with 3/8" &amp; 1/2 flare with or without Xcoating |  |  | May 23, 2014 | 2014 | Terumo Cardiovascular Systems Corporation | During an in-process inspection, Terumo Cardiovascular Systems (Terumo CVS) identified the presence ofloose fiber particulate that exceeded finished product specifications | Process control | Terumo sent an Urgent Medical Device Recall letters dated June 27, 2014 to their customers. The letter identified the affected product, problem and actions to be taken. The letter instructed customers to 1. Review this Medical Device Recall notice, 2. Assure that all users in your institution have received notice of this issue and remove any affected units from use, 3. Fill out the Product Tracking Sheet (enclosed), indicating the number of affected units at your institution and fax to 1-877-907-7509 or email to terumo4967@stericycle.com. 4. Complete the Packing Slip (enclosed) and follow the packing instructions to return the affected product, and 5. Terumo CVS will issue credit once the affected units are returned. For questions call Terumo CVS c/o Stericycle 1-866-629-6181. |  | 32880 | 32,880 units | Worldwide Distribution - US (nationwide) in the states of: AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, PR, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV &amp; WY and the countries of : Australia, Belgium, Canada, Chile, Japan, Malaysia, Mexico, Singapore, Taiwan, Thailand, United Arab Emirates (UAE) &amp; United Kingdom. | Terminated | February 24, 2015 | 277 | 376285 | Undetermined
Z-2107-2014 | 68731 | No | BTT | Respiratory gas humidifier. | Anesthesiology | Anesthesiology | 510(k) | Hudson RCI, Breathing Circuit, Adult, Single Limb w/Heated Wire, [Respiratory Gas Humidifier, (Direct Patient Interface)]. The Hudson Heated Wire Ventilator Circuit is intended as a conduit for respiratory gas between a patient and a ventilator, and includes heated wires for use with a Concha Column Humidifier. |  |  | July 09, 2014 | 2014 | Teleflex Medical | The temperature probe does not properly connect to the temperature port. | Nonconforming Material/Component | Teleflex Medical sent an Urgent Medical Device Recall Notification dated July 9, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Our records indicate that you have received products that are subject to this recall. We are now notifying our customers to take the following actions:1. If you have affected stock, immediately discontinue use and quarantine any products with the product and lot numbers listed above.2. To return product, complete the enclosed Recall Acknowledgement Form and fax it to 1-866- 804-9881, Attn: Customer Service. This will allow us to document the amount of product you have on hand for return. A customer service representative will contact you with a Return Goods Authorization (RGA) Number and will provide instructions for the return of product to Teleflex Medical.3. If you have no affected stock, please complete the enclosed Recall Acknowledgment Form and fax it to 1-866-804-9881, Attn: Customer Service. This will allow us to document your receipt of this letter.Teleflex Medical is committed to providing high quality, safe and effective products. We sincerely apologize for any inconvenience this action may cause your operations. If you have any other questions, feel free to contact your local sales representative or Customer Service at 1-866-246-6990. |  | 11679 | 11,679 units | Worldwide Distribution - US including the states of CA, FL, KS, KY, MA, MI, MO, NC, NE, NJ, NY, OH, PA, TN, TX and VA, and the countries of Argentina, Chile, Colombia, and the Philippines | Open, Classified |  |  | 11679 | Undetermined
Z-2118-2014 | 68615 | No | HAW | Stereotaxic instrument. | Neurology | Neurology | 510(k) | Medtronic Nexdrive Micropositioning Drive. Models Ml-1000 and Ml-2000. |  |  | June 11, 2014 | 2014 | Medtronic Neuromodulation | Potential for misalignment of the Z-stage scale. Using one of these devices for a procedure could result in the microelectrode being inserted to an incorrect target depth. | Nonconforming Material/Component | Medtronic sent an Urgent Medical Device Recall letter dated June 2014 on June 11, 2014, to all affected customers. The letter described the problem and the product involved in the recall. The letter described the "Action Requested" which included to stop using and return the recalled device and to complete and return the Reply Form. Customers with questions were instructed to contact Medtronic Neuromodulation Technical Services at 1-800-707-0933.EXPANDED 10/02/2014: Medtronic sent an Urgent Medical Device Recall letter beginning 10/02/2014. The letter explained that this letter is an update to the June 2014 action, to include 4 additional lot numbers. The letter explained the reason for recall identified all lots recalled and asked that devices stopped being used. A response form provided instructions on how to return product and the form was asked to be returned within 10 days. Questions should be directed to Medtronic Neuromodulation Technical Service at 1-800-707-0933, weekdays 7 am - 6 pm, Central time. |  | 3000 | 300 EXPANDED 10/02/2014 130 devices | Worldwide Distribution - USA including CA, CO, FL, IL, KS, KY, LA, MN, MT, NC, NE, OR, TX, VA, and UT, and Internationally to Canada, Spain, and United Kingdom. | Terminated | January 22, 2015 | 225 | 3000 | Undetermined
Z-2117-2014 | 68681 | Yes | CBP | Nonrebreathing valve. | Anesthesiology | Anesthesiology | 510(k) | CPR Shield With One Way Valve and Barrier Filter, Reorder No. 4921 --- Dynarex label, 100 per case --- Manufactured for: Dynarex Corporation - Made in China. |  |  | July 02, 2014 | 2014 | Dynarex Corporation | "THIS SIDE UP" is on the wrong side of the product. This error poses a potential health hazard if the wrong labeling is followed by caregivers. | Employee error | Voluntary Device Recall letters (dated 7/01/14) and recall verification forms were sent to customers via US Postal First Class Mail on 7/02/14. The letters instructed customers to check their inventory for the specific lots, discontinue distributing them, and quarantine them immediately. Customers were to promptly contact Dynarex Credit &amp; Return Representative Avi Celnick at 1-845-365-8200 ext. 6644 to arrange for product return and obtain an RGA #. For questions, please call Dynarex's toll free number 1-888-396-2739. |  | 250 | Domestic: 208 cases; Foreign: 42 cases | Worldwide Distribution -- USA, Canada, and Mexico. | Terminated | April 20, 2017 | 1023 | 250 | Undetermined
Z-2119-2014 | 68757 | No | CCW | Rigid laryngoscope. | Anesthesiology | Anesthesiology | 510(K) Exempt | Emerald Fiberoptic Blade, Macintosh 4, Teleflex Medical. A rigid laryngoscope is a device used to examine and visualize a patient's upper airway and aid placement of a tracheal tube. |  |  | July 09, 2014 | 2014 | Teleflex Medical | Packaging Error: Customers have reported receiving incorrect product. The box labeling indicates the correct product MAC 4 but the contents are incorrect, Miller 4 blades. | Other | Teleflex Medical sent an Urgent Medical Device Recall Notification dated July 9, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Customers were instructed to:1. If you have affected stock, immediately discontinue use and quarantine any products with the product and lot numbers listed above.2. To return product, complete the enclosed Recall Acknowledgement Form and fax it to 1-866- 804-9881, Attn: Customer Service. This will allow us to document the amount of product you have on hand for return. A customer service representative will contact you with a Return Goods Authorization (RGA) Number and will provide instructions for the return of product to Teleflex Medical.3. If you have no affected stock, please complete the enclosed Recall Acknowledgment Form and fax it to 1-866-804-9881, Attn: Customer Service. This will allow us to document your receipt of this letter.Teleflex Medical is committed to providing high quality, safe and effective products. We sincerely apologize for any inconvenience this action may cause your operations. If you have any other questions, feel free to contact your local sales representative or Customer Service at 1-866-246-6990. |  | 120 | 120 units | US Distribution including the states of FL, GA, KY, MD, MS, MT, PA, SC and TX. | Terminated | November 09, 2015 | 488 | 120 | Undetermined
Z-2112-2014 | 68556 | No | FRO |  |  | General &amp; Plastic Surgery | 510(k) | IPM Wound Gel Rx Wound Dressing |  |  | March 25, 2014 | 2014 | Wellspring Pharmaceutical Corporation | Potential for microbial contamination of the product. | Nonconforming Material/Component | WellSpring Pharmaceutical Corporation notified customers on March 25, 2014 via e-mail. The notification identified the affected product, the problem, and directed consignees to return all products. |  | 12540 | 3,135 cartons of 4/10g tubes | Worldwide Distribution -- USA, including the states of Alabama, Arkansas, Georgia, Illinois, Kentucky, Louisiana, Missouri, Mississippi, New Jersey, South Carolina, Tennessee, Texas, and Virginia; and the country of Canada. | Terminated | September 08, 2014 | 167 | 12540 | Undetermined
Z-2057-2014 | 68583 | No | MDB | Microbial growth monitor. | Microbiology | Microbiology | 510(k) | Thermo Scientific VersaTREK REDOX 1, Ref 7102-44, 80ml with Stir Bar (O2 aerobic), 50 bottles/box, IVD. The firm name on the label is Remel Inc., Lenexa, KS. |  |  | June 11, 2014 | 2014 | Remel Inc | Bottles did not provide positive signals within the time frame of the quality control specification for one of two strains of Streptococcus pneumoniae. | Employee error | The recalling firm made phone calls to their customers beginning 6/11/14 with follow-up letters dated 6/12/14 being issued via regular mail requesting the affected product be destroyed. Additional customers were later identified and additional phone calls were made on 6/16/14 and recall letters dated 6/16/14 were issued via regular mail. |  | 76 | 76/50-bottle boxes | Distribution was made to CO, IL, IN, NE, NJ, PA, SD, and a government facility located in VA. | Terminated | August 28, 2014 | 78 | 76 | Undetermined
Z-2055-2014 | 67813 | No | BRY | Anesthetic cabinet, table, or tray. | Anesthesiology | Anesthesiology | 510(K) Exempt | BAM (SHAPE) ARM, LEFT, an articulating, arm bracket that is attached to the service head of boom systems which are used for supporting and positioning equipment in operating rooms and patient care areas. The product comes as component of kit part number 0682000976, and the Arm part number in the kit is P17087. The kit, when assembled, supports a flat-panel video monitor that is attached to one of two types of Stryker Communications boom systems, EDS or FLEXiS. |  |  | March 17, 2014 | 2014 | Stryker Communications | During the use of the SHAPE Arm there is the potential for the monitor mount to come off of the SHAPE Arm if the lock washer and screw are not present. | Process control | Direct accounts in the US were notified by letter dated March 21, 2014, and asked to report any malfunctions to Stryker Communications immediately. In addition, Stryker Representatives will visit all US accounts to inspect all the product units on hand. If needed, Representatives will check that black, screw/lock washers supplied by Stryker are secured in place and will replace any that are not the typed furnished by Stryker. All units will be updated with a yellow, warning sticker on the monitor bracket which reads: "Failure to properly install the screw and lock washer below can result in the attached component/equipment falling." Stryker Communications' foreign distributors will oversee notification, inspection, and corrections of their accounts and report corrections and other findings to Stryker Communications. |  | 953 | 953 SHAPE Arms of all types | Worldwide Distribution. US nationwide including Puerto Rico, China, Philippines, Mexico, Korea, Japan, Colombia, Chile, Canada, Brazil, and Australia | Terminated | May 15, 2015 | 424 | 953 | Undetermined
Z-2051-2014 | 68581 | No | BTY | Predictive pulmonary-function value calculator. | Anesthesiology | Anesthesiology | 510(k) | Shape HF Cardiopulmonary System. Made up of Shape HF System Analyzer (PN 0001-9001) and Disposable Patient Interface (DPI), (PN 0004-4001). A pulmonary gas exchange testing system used during cardiopulmonary exercise testing where collection and review of gas exchange variables are indicated. The system will provide physiological data to physicians to aid in patient assessment. |  |  | May 28, 2014 | 2014 | Shape Medical Systems, Inc | Shape Medical has initiated a correction due to a mandatory software upgrade for the Shape HF Cardiopulmonary Testing System prior to use of the impacted DPI lot numbers. Use of the incorrect software version with the impacted DPIs could result in erroneous testing results potentially leading to incorrect diagnosis and incorrect treatment. | Under Investigation by firm | Shape Medical Systems sent an Urgent Medical Device Correction letter dated June 17, 2014, to all affected customers. Customers were informed that new orders of the Disposable Patient interface devices (DPI) will require a software upgrade for the Shape HF Cardiopulmonary Testing System prior to using the new DPIs. Impacted DPIs will be shipped with Instructions to conduct the software upgrade and Technical Service will contact customers shipped the impacted DPIs to confirm receipt and assist with the software upgrade. A Response Form was requested to be returned. Customers with questions were instructed to call Shape Technical Service at 1-888-906-6266 or 651-621-2990. |  | 800 | 800 | Nationwide Distribution | Terminated | May 22, 2015 | 359 | 800 | Undetermined
Z-2050-2014 | 68584 | No | DZE | Endosseous dental implant. | Dental | Dental | 510(k) | Zimmer Dental Tapered Screw-Vent Implant, Catalog TSVT6B13, Lot 62284006 |  |  | June 25, 2014 | 2014 | Zimmer, Inc. | Zimmer Dental is conducting a voluntary recall of a single lot of the Tapered Screw-Vent Implant, Catalog Number TSVT6B13, Lot number 62284006, due to a potential crack in the cap of the outer vial. | Under Investigation by firm | Recall notification letters sent to customers on 06/25/14.The recall letter, titled "URGENT DEVICE RECALL NOTICE", delivered to customers via a courier, with instructions for returning any unused affected product. The recall letter was accompanied with a response form and provided the following information to customers: product information with codes, reason for recall, risks, responsibilities, MedWatch Reporting and contacting information. |  | 129 | 129 units | Nationwide Distribution. | Terminated | October 30, 2014 | 127 | 129 | Undetermined
Z-1991-2014 | 68520 | No | MMI | Creatine phosphokinase/creatine kinase or isoenzymes test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | ADVIA CentaurÂż Systems TnI-Ultra assay; Catalog Number 02789602 (100 tests), SMN 10317708 and 02790309 (500 tests), SMN 10317790 |  |  | May 30, 2014 | 2014 | Siemens Healthcare Diagnostics, Inc | The solid phase reagent in some of the ADVIA CentaurÂż TnI -Ultra ReadyPacksÂż, lot 084, is darker in appearance, creating a potential for incorrect control and patient results. | Nonconforming Material/Component | An Urgent medical Device Correction (UMDC) was mailed, via FedEx, to all affected Siemens Healthcare Diagnostics customers in the United States on May 30, 2014 for customer receipt on June 2, 2014. All international customers were provided with a copy of the Urgent Field Safety Notice via hard copy, e-mail, and/or fax as determined by each country's local regulations and procedures. Each customer is asked to complete a fax-back form. This notice informed the customers of the potential for incorrect control and patient results due to solid phase which is darker in appearance. The notice instructs customers to place one TnI-Ultra ReadyPack on the system at a time and to calibrate and run controls on each ReadyPack to ensure accurate control and patient values. Patient TnI-Ultra results produced with this assay are acceptable if they follow acceptable calibration with valid quality control results. Upon receipt of kit lots ending in 086 and higher, calibration of each ReadyPack is not required. This notice should be reviewed with the facility's laboratory or medical director. |  | 34283 | 34,283 kits | Worldwide Distribution-USA (nationwide) including PR and the states of AK, AL, AR, AZ, CA, CO, CT, DE, FL, GA, HI, IA, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, MT, NC, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, and WY, and the countries of Canada, Mexico, Argentina, Austria, Australia. Bahrain, Belgium, Brazil, Bulgaria, Chile, China, Colombia, Costa Rica, Croatia, Cyprus, Czech Republic, Denmark, Ecuador, Egypt, Estonia, Finland, France, French Polynesia, Great Britain, Greece, Guadeloupe, Hong Kong, Iceland, India, Indonesia, Ireland, Israel, Italy, Japan, Latvia, Lebanon, Libya, Macedonia, Malaysia, ,Malta, Martinique, Netherlands, New Zealand, Norway, Pakistan, Paraguay, Peru, Philippines, Poland, Portugal, Qatar, Reunion, Republic of Korea, Romania, Russian Federation, Serbia, Saudi Arabia, Singapore, Slovakia, Slovenia, Spain, Sri Lanka, South Africa, Sweden, Switzerland, Taiwan, Thailand, Tunisia, Turkey, United Arab Emirates, Uruguay, and Vietnam. | Terminated | February 26, 2016 | 637 | 34283 | Undetermined
Z-1990-2014 | 68524 | No | LRG | Antimicrobial susceptibility test powder. | Microbiology | Microbiology | 510(k) | MicroScan WalkAway-40 plus Instrument and MicroScan Walkaway-96 plus instrument access door hinge. |  |  | June 02, 2014 | 2014 | Siemens Healthcare Diagnostics, Inc. | Springs contained in the access door hinge assembly on the Walk Away Plus instruments are becoming detached causing premature failure of the hinge assembly and a potential safety issue for the customers. | Component design/selection | Urgent Medical Device Recall notification letter was sent to all affected US customers on June 2nd by Federal Express.OUS notifications were also sent to regional Siemens organizations who distributed in-country notifications and relevant notifications. |  | 1027 | WA40: 560 US, 131 OUS, 9 already repaired; WA96: 81 US, 189 OUS, 57 repaired. | Worldwide Distribution-USA (nationwide) and the countries of Algeria Argentina Brazil Canada China Chile Columbia France Indonesia Libya Georgia Germany India Italy Japan Kazakhstan Lithuania Oman Pakistan Peru Poland Portugal Republic of Korea Romania Russia Saudi Arabia Spain Turkey. | Terminated | December 23, 2015 | 569 | 1027 | Undetermined
Z-1987-2014 | 68495 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Microline Surgical ReNew Laproscopic Fenestrated Grasper Forceps Tips (Reusable) |  |  | May 12, 2014 | 2014 | Microline Surgical, Inc. | Potential for the Fenestrated Grasper Tip jaw to break when force is applied to the jaw | Nonconforming Material/Component | Microline Surgical notified Distributors, Hospitals (Customers/End Users) and Suppliers via letter dated May 15, 2014 with labeling, pictures, and Recall Response Forms via UPS Overnight Air, email, and/or fax. Product is requested to be returned and a RMA number will be issued. Contact 978-867-1775 for questions. |  | 6290 | 6290 units | Worldwide Distribution-USA including the states of CA, HI, ID, MA, MI, MS, PA, TN, UT, and TX. and the countries of Belgium, France, Ireland, Germany, Indonesia, Korea, Norway, Saudi Arabia, Switzerland, Turkey, and UK. | Terminated | August 13, 2015 | 458 | 6440 | Undetermined
Z-1985-2014 | 68382 | No | IZL | Mobile x-ray system. | Radiology | Radiology | 510(k) | Siemens Mobilett XP Digital Mobile X-ray System |  |  | May 09, 2014 | 2014 | Siemens Medical Solutions USA, Inc | There is a potential fire hazard of certain Lithium-Ion batteries in the computers used in the original manufacturing or as spare parts between October 2010 and April 2011 for the Siemens Mobilett XP Digital mobile X-ray System, which might occur at any time with higher probability while the system is being charged and connected to the main power supply. | Component design/selection | Siemens sent a Safety Advisory Notice letter dated May 9, 2014, to affected customers. The letter identified the affected product, problem and actions to be taken. The letter stated that a Siemens service technician would be visiting their site to identify and service affected devices. |  | 10 | 10 | US Nationwide Distributions in the states of MO, GA, FL, NJ, CA, PA and including NM. | Terminated | December 15, 2014 | 220 | 10 | Undetermined
Z-1983-2014 | 68496 | No | HSZ | Surgical instrument motors and accessories/attachments. | General &amp; Plastic Surgery | Orthopedic | 510(K) Exempt | Anspach Pneumatic and Electric Motor Systems. |  |  | May 08, 2014 | 2014 | The Anspach Effort, Inc. | Several attachments and handpieces failed to meet the expected sterility assurance level (SAL) when sterilized by the method that is specified in the current Directions for Use (DFU) provided with the device. | Process control | Customers were notified 05/08/2014 by a Return Receipt Requested letter from Anspach with an URGENT: Medical Device Labeling Correction announcement, REGARDING: STERRAD 100S Hydrogen Peroxide Gas Plasma Technology no longer recommended for sterilization of Anspach products. The letter identified the problem and actions to be taken by the customer. Customers were to complete and return the attached reply form. If the products were distributed further, then those customers should have been notified as well. Customers were to contact Anspach Customer Support at 1-800-327-6887 if they want copies of the revised labeling. |  | 115988 | 115,988 units | Worldwide Distribution -- USA, including the states of AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX,UT, VA, VT, WA, WI, WV, WY, and Hawaii, and the territories of Puerto Rico, US Virgin Island and Guam; and, the countries of Argentina, Australia, Austria, Belgium, Brazil, Cameroon, Canada, Chile, Costa Rica, Czech Republic, Dominican Republic, Egypt, Finland, France, Germany, Grand Cayman, Hong Kong, India, Israel, Jamaica, Japan, Korea, Lithuania, Malaysia, Mexico, Nassau, Netherlands, Nigeria, Palestine, Panama, Poland, Portugal, Saudi Arabia, Singapore, South Africa, Spain, Sweden, Switzerland, Syria, Thailand, Turkey, United Arab Emirates, United Kingdom, and Venezuela. | Terminated | January 29, 2016 | 631 | 115988 | Undetermined
Z-1948-2014 | 68242 | Yes | BTM | Manual emergency ventilator. | Anesthesiology | Anesthesiology | 510(k) | Manual Resuscitator Bag Series: AF1000, AF2000, AF3000, AF4000, AF5000, BT2000, BT3000, BT4000, BT5000, BVM700, CPRM1000, CPRM2000, CPRM3000, PRO-1900, PRO-1000, PRO-2000, SC7000, SC8000, SC8020, SC8120, SC9000, SS3200, VN2000, VN3000, VN4000, VN5000, VT1000 , VN2102, VN2002. Emergency manual resuscitator for the Pediatric Population. |  |  | May 16, 2014 | 2014 | Ventlab LLC | The pop-off valve in the defective devices remain open and a squeeze of the bag may not generate enough force to force the duck bill valve open and therefore no air will be delivered to the victim. | Process change control | The firm, Ventlab, issued a press dated May 14, 2014, stating Ventlab, LLC has initiated a voluntary medical device removal of a limited number of Ventlab" Resuscitator Bags. The bag series and lot numbers were included in a table as well as manufacture dates. In addition, Ventlab, LLC sent an "URGENT: MEDICAL DEVICE PRODUCT REMOVAL IMMEDIATE ACTION REQUIRED" letter dated May 19, 2014 to its distributors and customers. The letter described the product, problem and actions to be taken. The customers were instructed to review and examine their inventory for affected product; stop using them and immediately contact Ventlab, LLC for further instructions on the return of these products; and if product was further distributed, identify their customers/consignees and notify them of this product removal. The customers were also instructed to complete and return the attached Return Response Form via fax at 1-800-400-8820, Attn: Quality Department, or email to: PFA@ventlab.com as soon as possible and inform all affected personnel of this removal.Ventlab will send new corrected replacement resuscitation bags to customers once you return the affected product.If you have any questions regarding this action, please call Ventlab at 1-800-237-5481 between the hours of 8:30 AM to 5:00 PM (EST) Monday through Friday, or 1-844-635-5326 or via e-mail at PFA@ventlab.com. |  | 251384 | 251,384 | US Nationwide Distribution in the states of OH, NY, TX, TN, FL, CT, and VA, | Terminated | November 17, 2015 | 550 | 251384 | Undetermined
Z-1981-2014 | 68489 | No | DQO | Diagnostic intravascular catheter. | Cardiovascular | Cardiovascular | 510(k) | Langston Dual Lumen Catheters, 6Fr, REF 5550 and 5540. The Vascular Solutions Langston Dual Lumen Pressure Monitoring catheter is used as a pressure measurement catheter and for delivery of contrast media during angiographic studies. |  |  | May 27, 2014 | 2014 | Vascular Solutions, Inc. | The inner catheter of some Langston V2 Dual Lumen Catheters have separated from the device hub during contrast pressure injections. This may require an intervention to retrieve the separated piece from the patient to prevent injury. | Component design/selection | Vascular Solutions sent an Urgent Medical Device Removal letter dated May 23, 2014, to all affected consignees.The letter described the problem and the product affected by the recall. Advised consignees to remove affected devices and place them in a secure area. Requested consignees to fill and return the "VSI Account Inventory Form" to VSI's Customer Service Department by email customerservice@vasc.com or fax 763-656-4251. VSI will replace all returned devices. Customers with questions were advised to contact their local Vascular Solutions Account Manager.For questions regarding this recall call 763-656-4300. |  | 8580 | 8,580 total (US 5714) | Worldwide Distribution - USA nationwide including DC and Puerto Rico, AUSTRIA, CANADA, GERMANY, ISRAEL, ITALY, NETHERLANDS, SPAIN, SWEDEN SWITZERLAND, and UNITED KINGDOM. | Terminated | January 30, 2015 | 248 | 8580 | Undetermined
Z-1992-2014 | 68488 | No | MGQ |  |  | General &amp; Plastic Surgery | 510(k) | Amerigel Wound Dressing 1 oz. tube |  |  | March 21, 2014 | 2014 | Kova Laboratories, Inc. | Product may not have been manufactured pursuant to a validated process protocol and CGMP. | Process control | Kova Laboratories contacted customers with an URGENT: Medical Device Recall message by phone and e-mail 3/21/2014 and 4/21/2014, of the affected products to be returned. |  | 122396 | 122,396 / 1oz. tubes | Distributed in the states of FL and NJ. | Open, Classified |  |  | 122399 | Undetermined
Z-1998-2014 | 68512 | No | DPW | Cardiovascular blood flowmeter. | Cardiovascular | Cardiovascular | 510(k) | 07-150-07 MIZUHO, MAL DISP DOPPLER PROBE BX 4, REF 138200 |  |  | May 28, 2014 | 2014 | Vascular Technology, Inc. | Report of the product penetrating the sterile barrier; thiscould render the product unsterile. | Process design | Vascular Technology sent an Urgent Medical Device Recall letter dated May 28, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.This recall has been initiated due to a report of the product penetrating the sterile barrier; this could render the product unsterile. Use of this product may lead to infection in the patient.Please immediately examine your inventory and return all product subject to recall to Vascular Technology at the above mentioned address. In addition, if you have further distributed this product, please identify your customers and notify them at once of this product recall.Your notification to your customer must include instructions telling your customer to return the recalled product if unused or to account for the recalled product if it has been used.For further questions please call (603) 594-9700. |  | 164 | 164 | Worldwide Distribution - US including California and the countries of Australia, United Kingdom and Taiwan. | Terminated | February 11, 2016 | 624 | 370 | Undetermined
Z-1996-2014 | 68552 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Henora REF: 70-4026-09V1 300psi I.V. Catheter Extension Set |  |  | March 28, 2014 | 2014 | RGI Medical Manufacturing, Inc. | Problem with low adhesion of the tubing to luer connection which may lead to a separation of the extension set while in use. | Process control | RGI Medical Manufacturing, Inc. phoned their direct account: Baxter Healthcare on February 28, 2014 and sent an Urgent: Device Recall letter dated March 21, 2014. The letter identified the affected product, problem and the actions to be taken. Contact 800-747-6870 to obtain instruction for returning the product to RGI |  | 19200 | 19,200 units. | US Nationwide Distribution in the states of AL, CA, DE, GA, IL, KY, MA, MD, ME, MI, MN, NJ, NM, NV, NY, OH, TN, TX, WA, and WI. | Open, Classified |  |  | 19200 | Undetermined
Z-1997-2014 | 68243 | No | CCX | Patient position support. | Anesthesiology | Anesthesiology | 510(K) Exempt | Devon Supine, Lithotomy and Trendelenburg Positioning Kit; Product Code: 31151090 |  |  | May 08, 2014 | 2014 | Covidien LLC | Straps may separate from the foam pad. | Process control | Covidien Sales Representatives notified customers in person by beginning May 8, 2014. The Sales Representative will provided a copy of the Recall letter (attachment 4), and collect any remaining product at the site. The notification and recovery of the product is expected to be complete no later than May 14, 2014. |  | 264 | 264 devices | US Nationwide Distribution in the states of NJ, CA, TX, IL, MO, PA, OH, IN, MI, MN, FL, IN, DC, PA, VA, MS, ME, NV, AR, KY, NY, KS, SC | Terminated | September 04, 2015 | 484 | 264 | Undetermined
Z-2020-2014 | 68553 | No | NAY | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | Intuitive Surgical Inc. Large Clip Applier instruments used in conjunction with da Vinci S and da Vinci SI Surgical Systems |  |  | June 13, 2014 | 2014 | Intuitive Surgical, Inc. | Manufacturing variation in the grip assembly of the Large Clip applier which can lead to the instrument jamming during insertion when used with the 8 mm cannula or with certain reducer accessories.. | Process control | Urgent Medical Device Correction letters were sent on June 18, 2014 to all affected sites via trackable mail. |  | 847734 | 5290 for part 420230-04; 1964 for part 420240-06 | Worldwide Distribution. | Terminated | December 01, 2014 | 171 | 847734 | Undetermined
Z-2022-2014 | 68546 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | STOPCOCK EXT. SET W/ 3-WAY STOPCOCK WITH INJ. SITE AND SPINLOCK CONNECTOR, 22 inches and 37 inches |  |  | May 29, 2014 | 2014 | B. Braun Medical, Inc. | Complaints were received that reported the stopcock of some Stopcock Extension Sets are assembled in the reverse position, resulting in the stopcock being closed when set in the open position and vise versa. | Process control | An urgent medical device correction notification, dated June 3, 2014, describing the product, problem, and actions to be taken, was sent to end users and distributors. |  | 10835 | 10,835 | USA (nationwide) including the states of PA, IL, NC, FL, IL, CA, CO, NE, LA, NC, LA, VA, TN, NE, TX, and KS. | Terminated | December 10, 2014 | 195 | 10835 | Undetermined
Z-2027-2014 | 68658 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | The Hospira MedNet Medication Management Suite (MMS). |  |  | June 06, 2014 | 2014 | Hospira Inc. | Hospira MedNet 5.5, 5.8.1, and 5.8.2 contains software defect where the dosing units of "nanog/kg/min" and "milliUnits/min" are not sent to a Plum A+ device (Version 13.40, 13.41, and 13.4.2) via Auto Programming when those dosing units are not properly established in the customized drug library. A dosing unit mismatch of "nanog/kg/min" or "milliUnits/min" on an unmatched medication will produce a | Under Investigation by firm | Hospira sent an URGENT MEDICAL DEVICE CORRECTION letter dated June 6, 2014 was sent to all affected customers. The letter included instructions for customers to: notify all potential users in their organization of the recall, and to complete and return the attached reply form. The letter further notified customers that they will be contacted by a member of Hospira's Clinical Services Team to make arrangements to review the drug library and implement the appropriate mitigation. Hospira will further work with the direct account to ensure Auto Programming installations receive Pharmacy consulting to support the implementation of the mitigation and that a correction for this issue will be addressed in a future version of the software. Reports of adverse events or product complaints can be made by contacting Hospira Global Complaint Management at 800-441-4100 (Monday - Friday, 8:00 AM - 5:00 PM CST) or via e-mail at ProductComplaintsPP@hospira.com. Additional information or technical assistance can be obtained by contacting the Hospira Advanced Knowledge Center at 800-241-4002, option 4 (24 hours a day, 7 days a week). |  | 13 | 13 units | US Nationwide in the states of FL, LA, PA, WA, WY and the country of Hong Kong | Open, Classified |  |  | 13 | Undetermined
Z-2039-2014 | 68646 | No | LMD | Medical image communications device. | Radiology | Radiology | 510(K) Exempt | Essential NEO, Audio/Visual Integration Unit |  |  | June 25, 2014 | 2014 | Karl Storz Endoscopy | KARL STORZ Endoscopy-America is recalling the Essential NEO because the connected power cable could potentially become overheated and result in smoke being generated and the cessation of function of the peripheral that the cable is connected to. | Nonconforming Material/Component | Storz sent an Safety Alert letter dated June 25, 2014, to all customers who purchased the Essential NEO. The letter informs the customers of the problems identified and the actions to be taken. Customers are instructed to complete the attached Acknowledgement/Response Form with contact information and return to KARL STORZ. Customers with questions are instructed to contact KARL STORZ' Technical Support Department at (800) 421-0837, option 3, then option 4. |  | 968 | 968 units | Nationwide Distribution including the states of: AL, AZ, CA, CO, CT, District of Columbia, DE, FL, GA, HI, IA, IL, IN, KY, LA, MX, MS, MD, MI, MA, MO, MN. NM, NV, NJ, NY, NC, NE, OK, OH, PA, SC, SD, TX, UT, VA, WI and WA., and Internationally to Canada. | Terminated | January 13, 2015 | 202 | 968 | Undetermined
Z-2032-2014 | 68547 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | Zimmer EDI PERIARTICULAR PLATING SYSTEM CANCELLOUS BONE SCREW, FULLY THREADED, HEX HEAD 4MM DIAMETER, 14 MM LENGTH, and EDI 4.0 MM CANCELLOUS SCREW, FULLY THREADED, 20 MM LENGTH. Temporary internal fixation devices, designed to stabilize fractures during the normal healing process. |  |  | June 12, 2014 | 2014 | Zimmer Manufacturing B.V. | Two lots of screws were commingled. Screws with etching and machining for lot number 62629432 were packaged in bags labeled for lot number 62628759; and screws with etching and machining for lot number 62628759 were packaged in bags labeled for lot number 62629432. The affected product was distributed between 3/6/14 and 4/21/14. | Process design | Zimmer sent an Urgent Medical Device Recall Lot specific letter dated June 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Customers were instructed:1. Please locate, quarantine, and return the affected product (part number 00-4840-020-00 lot number 62629432 and part number 00-2347-022-14 lot number 62628759) from all inventories.2. Please note that any hospitals that received direct shipments of this product from Zimmer will be notified via Fed Ex.3. Please review the list of hospital and surgeons to be provided to your territory via email.4. Provide the name, address, and phone number of any additional hospitals or surgeons that may have received or used the affected product, or confirm that the list provided is complete.5. Complete and return Attachment 1.6. Ensure that a copy of the Risk Managers and Surgeons letter is left with the facilities from which product is removed by providing them with a copy of the letter, if they are not identified as a facility receiving a letter directly from Zimmer.7. If after reviewing this notification you have questions or concerns, please contact Zimmer at 1-877-946-2761. |  | 496 | 496 | Worldwide Distribution - USA (nationwide) and the countries of India, Japan, China, Australia and Canada. | Terminated | June 24, 2015 | 377 | 496 | Undetermined
Z-2040-2014 | 68591 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | ColorSafe IV lines |  |  | August 17, 2011 | 2011 | Codan Us Corporation | CODAN US Corporation is recalling the ColorSafe IV (intravenous) product lines because they were manufactured and marketed prior to FDA approval. | Other | CODAN sent a recall letter dated August 17, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to return any of the affected product. If they sold the product they should provide Codan the name and address of the customer. so that they may be contacted to retrieve the shipped products. Customers were instructed to contact CODAN Customers Service for a Returned Goods Authorization. Customers with questions were instructed to call 714-430-1303. Customers were instructed to not send used devices.For questions regarding this recall call 714-545-2111. |  | 7250 | 7250 units | Worldwide Distribution - US including TX, CT, NY and Internationally to Belgium. | Terminated | July 10, 2014 | 1058 | 7250 | Undetermined
Z-2031-2014 | 68200 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | GemStar infusion pump, List numbers 13000 and 13100, GemStar Infusion Pump (pain management) List number 13150. Used for intravenous, arterial, subcutaneous, short-term epidural infusion and parenteral administration of general I.V. fluids, medications, nutritional fluids, and blood/blood products. |  |  | January 31, 2014 | 2014 | Hospira Inc. | The connection between the beeper subassembly and the pump may fail. The GemStar infusion pump will identify this failure during the "self-test" while powering up which will result in a Beeper Error ("code 10/001/000"). This Beeper Error (code 10/001/000) is a service alarm that places the pump in an inoperable mode and requires service before it can be returned to service. | Component design/selection | Hospira sent an URGENT MEDICAL DEVICE RECALL letter dated January 31, 2014 was sent to all affected customers. The letter included instructions for customers to: 1) inform potential users of these devices of the notification, 2) if the beeper error occurs during power up, remove the device from service immediately, 3) report any beeper errors to Hospira by calling 800-441-4100 (Monday - Friday, 8:00 AM - 5:00 PM CST) or e-mail ProductComplaintsPP@hospira.com, 4) complete the attached reply form and return it to the fax number of e-mail address on the form (even if there is no product in inventory), and 5) if the products were further distributed, notify those customers that received the products and ask them to contact Stericycle at 855-827-6578 (Monday - Friday, 8:00 AM - 5:00 PM EST) to receive a reply form. A subsequent letter was sent on February 18, 2014 that included a correction to the description of the error code generated when a beeper wire failure occurs. That letter included identical instructions as those found in the January 31, 2014 letter for recipients to follow. |  | 36881 | 36,881 infusion pumps | Worldwide Distribution - US: Nationwide in the state of AK, AL, AR, AZ, CA, CO, CT, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, WY including District Of Columbia, Guam, Puerto Rico; and the countries of Austria, Australia, Bahrain, Belgium, Canada, Chile, France, Germany, Greece, Guatemala, Hong Kong, Italy, Jordan, Kuwait, Lebanon, Libya, Luxembourg, Malaysia, Mexico, Netherlands, Oman, Portugal, Saudi Arabia, Slovenia, South Africa, Spain, Switzerland, Taiwan, Thailand, Turkey, United Arab Emirates, United Kingdom, Venezuela | Open, Classified |  |  | 36881 | Undetermined
Z-2038-2014 | 68424 | No | CAH | Breathing circuit bacterial filter. | Anesthesiology | General Hospital | 510(k) | PALL Medical Breathing Circuit Filter, REF BB50T and BB50T-BULK , Rx ONLY Single use, bi-directional bacterial/viral removal filter. |  |  | June 04, 2014 | 2014 | Pall Corporation | Possible leak from the filter allowing a small of amount of air to be released from the junction between the filter end cap (flat side of the filter) and the filter housing connection. | Other | Urgent Product Recall Letters (dated 6/03/2014) and Response Forms were sent to customers via fax and e-mail starting on 6/04/14. The letters instruct the direct accounts to examine their inventory and place the affect lots on hold in quarantine. For questions, or if a patient injury associated with this device has been reported to you, please call Karen Peterson directly at 516.801.9267 (office) or 516.644.6171 (cell). The letter for the bulk product also instructs customers to notify their customers of the recall and advise them not use the affected lots if the product had been distributed. |  | 24650 | 24,650 units domestically | Distributed in the states of MN, MO, OH, MA, NY, NJ, WY, AZ, CA, IL, LA, IN, CO, NC, and TX. | Terminated | April 01, 2015 | 301 | 24650 | Undetermined
Z-1737-2014 | 68309 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Ultrasonic Transducer Kits |  |  | May 14, 2014 | 2014 | Beckman Coulter Inc. | Beckman Coulter is initiating a field action because some Ultrasonic Transducer Kits werepotentially assembled with incorrect mounting screws. | Mixed-up of materials/components | Beckman Coulter sent an "URGENT MEDICAL DEVICE RECALL" letter dated May 8, 2014 via first class mail. The letter identified the reason for recall, product information, issue, impact, action, resolution, contact information, and customer response form. Customers are advised to review historical quality control (QC) data and System Check results, and determine whether a review of patient test results is warranted.For questions regarding this letter, please contact our Customer Support Center:From our website: http://www.beckmancoulter.com/customersupport/support, By phone: call 1-800-854-3633 in the United States and Canada, and Outside the United States and Canada, contact your local Beckman Coulter representative |  | 82 | 82 units | Worldwide Distribution and USA Nationwide - and the countries of Australia, Canada, China, France, Hong Kong, Italy, Japan, Mexico, Portugal, Spain, Switzerland, Taiwan,and Turkey. | Terminated | December 25, 2014 | 225 | 82 | Undetermined
Z-1731-2014 | 68267 | No | ETD | Tympanostomy tube. | Ear Nose &amp; Throat | Ear Nose &amp; Throat | 510(k) | Armstrong Grommet with wire Otological Ventilation Tube |  |  | May 12, 2014 | 2014 | Summit Medical, Inc. | One lot of ventilation tubes have mixed tubes with wire attached to the slanted flange and ventilation tubes with wire attached to the straight flange. This poses no health hazard to the user or patient. | Employee error | Consignees was sent a Summit Medical "Medical Device Recall" letter, dated May 9, 2014. The letter described the problem and the product involved in the recall. Requested consignees to examine their inventory and to quarantine the product. Consignees need to immediately examine their inventory quarantine any affected product. Also, they are to complete and return the "Recall Response Form." For questions they can contact Nicole Dove at (651) 789-3921 or ndove@summitmedicalusa.com. |  | 9 | 9 boxes | Worldwide Distribution -- USA, including CA, and GERMANY. | Terminated | June 19, 2014 | 38 | 9 | Undetermined
Z-1468-2014 | 67793 | No | JXG | Central nervous system fluid shunt and components. | Neurology | Neurology | 510(k) | CODMAN CERTAS Programmable Valves, central nervous system shunt, models as follows: |  |  | December 13, 2013 | 2013 | Codman &amp; Shurtleff, Inc. | Ensuring all customers have the tools to assist in verifying the setting of the valve on x-ray images and the location and orientation of implanted CODMAN CERTAS Programmable Valves. | Process control | Codman Neuro issued notification packages dated 12/13/13 to Domestic Hospitals and international accounts. Specific letters and associated inclusion of the additional aids, include A) a valve setting verification guide with instructions for use with x-ray images, and B) a plastic valve cut out overlay). A response reply form was included to be completed and returned to confirm effectiveness of the notification. |  | 9445 | 9,445 units | Nationwide Distribution. | Terminated | April 21, 2016 | 860 | 9445 | Undetermined
Z-1545-2014 | 67930 | No | KQO | Automated urinalysis system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | CLINITEK Status+ Analyzer (portable urine chemistry analyzer). Catalog number 10490946. |  |  | February 26, 2013 | 2013 | Siemens Healthcare Diagnostics Inc | Device with Sample Interference Notes (SIN) enabled not cleared for US marketing. | Reprocessing Controls | On February 26, 2013, a Siemens Field Engineer from Siemens Healthcare visited customer site and disabled the Sample Interference Notes feature on the instrument. |  | 1 | 1 unit | US Distribution in the state of: CA. | Terminated | May 06, 2014 | 434 | 1 | Undetermined
Z-1482-2014 | 67591 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Abbott Acclaim Infusion Pump |  |  | October 31, 2013 | 2013 | Hospira Inc. | Hospira has received customer reports of broken door assemblies on the Hospira/Abbott Acclaim Encore infusion pumps. If the door assembly breaks, it may prevent the door from closing properly and unrestricted flow may occur. If the door cannot be closed, the pump cannot be used which canresult in a delay in therapy. | Component design/selection | URGENT MEDICAL DEVICE CORRECTION letters, dated October 31, 2013, were sent to all direct accounts. The letters identified the affected product, as well as the reason for recall, risk to health, and instructions for device users. Users are to inspect each Hospira/Abbott Acclaim Encore infusion pump for door handle cracks prior to programming a therapy by following the provided instructions.The letter also included instructions for recipients to ensure that all potential users in the facility are made aware of the safety notification and the recommended actions; and complete the attached reply form and return it to the fax number or e-mail address on the form. Customers are to contact Stericycle at 866-891-0586 (Monday - Friday, 8:00 AM - 5:00 PM ET) to obtain additional reply forms if needed. If the devices were further distributed, those customers should be notified and asked to contact Stericycle at 866-891-0586 (Monday - Friday, 8:00 AM - 5:00 PM ET) to receive a reply form. The letter reminds customers that Hospira is retiring the Acclaim Encore infusion pump and will no longer support them as of December 31, 2014. |  | 13218 | 13,218 infusion pumps total | Nationwide Distribution -- AK, AL, AR, AZ, CA, CO, CR, CT, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, MT, NC, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, WA, WI, WV, and WY. | Terminated | August 16, 2017 | 1385 | 13218 | Undetermined
Z-1534-2014 | 67973 | No | NEW | Absorbable polydioxanone surgical suture. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | CAPIO, Monodeck, Violet Monofilament Polydioxanone Suture, 1 x 48 inches (122cm) Absorbable Surgical Suture, Rx Only, Teleflex Medical. Product Code: 833-137. |  |  | March 11, 2014 | 2014 | Teleflex Medical | The product did not meet minimum and/or average minimum Teleflex resorption strength requirements. | Other | Teleflex Medical sent an Urgent Medical Device Recall Notification dated March 11, 2014. The letter identified the affected product, problem and actions to be taken. Custmers were instructed to return all affected product to Teleflex Medical per the instructions on the Urgent Recall Notice. Customers were asked to complete the enclosed Recall Acknowledgement Form and fax to 1-866-804-9881, Attn: Customer Service. For questions contact your local sales representative or Customer Service at 1-866-246-6990 |  | 7380 | 7,380 ea (total) | Worldwide Distribution - US Nationwide in the states of CA, CO, GA, LA, IL, MA, MI, MN, MO, NC, and in the countries of Ireland and Singapore. | Terminated | March 29, 2016 | 749 | 7380 | Undetermined
Z-1533-2014 | 68141 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | GE Healthcare Discovery MR450, Discovery MR750, Optima MR450w, Optima MR450w equipped with the GEM option, Discovery MR750w, and Discovery MR750w equipped with the GEM option products. |  |  | April 23, 2014 | 2014 | GE Healthcare, LLC | GE Healthcare has recently become aware of a potential safety issue with certain GE MRI Products. Patient egress from the system can be delayed or impeded if the cradle release handle and/or cradle release block mechanism malfunctions or becomes misaligned. This could make the safe removal of the patient from the system difficult. | Process design | Consignees were sent on 4/23/2014 a GE Healthcare Ref # 60863 "Urgent Medical Device Correction" letter dated April 18, 2014. The letter was addressed to Hospital Administrators / Risk Mangers, Radiology Department Mangers and Radiologists. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction, and Contact Information. If customers encounter the stated problem with the device, they should contact their GE Healthcare Service Representative for an inspection and necessary repairs. GE will correct all affected devices at no cost to the customers. For questions customers should contact their local service representative. |  | 1548 | 1,548 units (704 US, 844 OUS). | Worldwide Distribution -- USA, except ND, RI and WY, and including DC, Guam, and Puerto Rico; and, the countries of ALGERIA, ARGENTINA, AUSTRALIA, AUSTRIA,AZERBAIJAN, BAHRAIN, BELARUS, BELGIUM, BRAZIL, CANADA, CAYMAN ISLANDS, CHINA, CROATIA, CZECH REPUBLIC, DENMARK, EGYPT, FINLAND, FRANCE, FRENCH POLYNESIA, GERMANY, HONG KONG, HUNGARY, INDIA, INDONESIA, IRELAND, ISRAEL, ITALY, JAPAN, KAZAKHSTAN, KOREA, KUWAIT, LEBANON, MARTINIQUE, MEXICO, MOROCCO, NETHERLANDS, NEW CALEDONIA, NEW ZEALAND, NORWAY, OMAN, POLAND, QATAR, ROMANIA, RUSSIA, SAUDI ARABIA, SERBIA, SINGAPORE, SOUTH AFRICA, SPAIN, SWEDEN, SWITZERLAND, TAIWAN, THAILAND, TUNISIA, T UREY, UNITED ARAB EMIRATES, UNITED KINGDOM, URUGUAY, UZBEKISTAN, and VENEZUELA. | Terminated | January 21, 2015 | 273 | 1548 | Undetermined
Z-1536-2014 | 67948 | No | MVU | Analyte specific reagents. | Pathology | Hematology | 510(K) Exempt | Anti EpCam PE - Analyte Specific Reagent. |  |  | April 04, 2014 | 2014 | BD Biosciences, Systems &amp; Reagents | Two lots of Anti EpCAM PE-Analyte Specific Reagent are mislabeled - the actual antibody present in the vial is CD9 PE. | Under Investigation by firm | Customer product recall notification letters were sent to all consignees on April 4, 2014 by certified mail. |  | 120 | 12 for lot 3295974; 10 of lot 3259822 | Worldwide Distribution-USA including the states of CA, TX, FL, MO, NY, MI, MO, and AZ and Seoul, Korea. | Terminated | September 28, 2015 | 542 | 120 | Undetermined
Z-1538-2014 | 68093 | No | DWS | Cardiovascular surgical instruments. | Cardiovascular | Cardiovascular | 510(K) Exempt | Joseph Lamelas Atrial Lift System, Model Number MI-ALS-001. A single surgical instrument or tool for retraction of the atrial wall tissue in cardiac surgery. |  |  | March 19, 2014 | 2014 | Biorep Technologies | The support arm attached with an adhesive to the body of the blade is detaching while attempting to engage the visor component. | Nonconforming Material/Component | Biorep Technologies sent a letter dated March 19, 2014, to its affected customer. It was advised to due to the high risk, according to our risk categories shown in Table 2, at this time, it is recommended that product be retrieved from its points of use.Any further questions please call (305) 687-8074. |  | 80 | 80 boxes (480 units) | US Distribution to a distributor in MN. | Terminated | August 07, 2014 | 141 | 80 | Undetermined
Z-1543-2014 | 67871 | No | FNK | Hydraulic adjustable hospital bed. | General Hospital | General Hospital | 510(K) Exempt | RetractaBed , Clinical Contour, Med - Mizer Inc..... Model SS 801, 802, 803, 803A: 80 " bed with 1, 2, or 3, motors and autocontour |  |  | April 18, 2014 | 2014 | Med-Mizer, Inc. | All configurations of the SS Retractabed, Clinical Contour, made prior to June 7, 2008 may have had insufficient welding on the foot cross tube. This tube is responsible for holding the two support rails together when using the foot side rails and keeps them from separating from the frame. If the weld on the tube breaks, it can lead to bed collapse or the rollers coming out of the channel on t | Process change control | On 5/1/14, Med-Mizer sent URGENT: MEDICAL DEVICE RECALL notification to their consignees. Notification included description of the affected product, reason for recall, risk to health, instructions on how to recognize how the device may fail, and instructions for actions to be completed by their consignees. Contact Information: 1-877-867-7365, sales@med-mizer.com, 8am-5pm EST. |  | 2980 | 2980 | USA (nationwide) | Open, Classified |  |  | 3172 | Undetermined
Z-1544-2014 | 68045 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | 12-Instrument Sterilization Tray; |  |  | April 16, 2014 | 2014 | Stryker Endoscopy | Gravity sterilization parameters provided in the 12 Instrument Sterilization reprocessing guide did not fully consider all worst-case scenarios during validation testing. | Reprocessing Controls | Stryker sent an Urgent Device Corrections letters dated April 18, 2014 to all US customers via certified mail. Stryker Sales representative and Stryker international distributors were notified with the letter and response card by e-mail. International sites are responsible for notifying customers. The letter identified the affected product, problem, potential hazards and actions to be taken. Customers were instructed to complete and sign the enclosed Acknowledgement of Receipt form and scan and e-mail to endorecall@stryker.com. For questions call 1-800-624-4422 option 3 or via email at endorecall@stryker.com. |  | 2591 | 2591 trays | Worldwide distribution | Terminated | August 18, 2014 | 124 | 2591 | Undetermined
Z-1540-2014 | 67910 | No | NBC | B-type natriuretic peptide test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Alere Triage BNP Test for the Beckman Coulter Access Family of Immunoassay Systems, Model Number 98200, Lot Numbers: 329599, 331265. The Alere Triage BNP test is intended to be used as an aid in the diagnosis of congestive heart failure (heart failure), as an aid in the assessment of severity of congestive heart failure, for the risk stratification of patients with acute coronary syndromes, and for the risk stratification of patients with heart failure. |  |  | March 20, 2014 | 2014 | Alere San Diego, Inc. | These lots may not recover within range when tested using certain commercially available controls. | Under Investigation by firm | Alere San Diego, Inc. initiated this recall by sending out notification letters via email and/or regular mail on March 20, 2014. The letter titled "URGENT MEDICAL DEVICE CORRECTION", dated "March 20, 2014, informed customers of the reason for recall, and potential impact, product description with lots numbers, customer/distributor required action, and contact information. The letter was accompanied by a "Customer/Distributor Verification Form". Customers were instructed that they may continue to use reagent lots 329599 and 331265." If they are using an alternate commercial control with Reagent Pack Lots 329599 or 331265, they can receive the Alere TriageÂż BNP QC Controls PN 98201 by requesting them on the Verification Form."Customers were instructed to share this information with their laboratory staff, including their laboratorys medical director and retain this notification as part of their laboratory Quality System documentation.For questions regarding this recall call 858-805-2000, ext 3015." If they have forwarded the product listed above to another laboratory, please provide a copy of this letter to them." Complete and FAX the enclosed Verification Form within 10 days to confirm your receipt of this notice. Only one Verification Form is required per facility.Customers with questions were instructed to call 877-308-8287 or email Triage.Support@alere.com. |  | 12592 | 12,592 kits | Worldwide Distribution - USA (nationwide) and Internationally to Australia, Canada, China, EU, Mexico, Russia, Singapore, South Africa, South Korea, Taiwan and United Arab Emirates. | Terminated | January 23, 2015 | 309 | 12592 | Undetermined
Z-1537-2014 | 68016 | No | JPD | Hemoglobin A[bdi2] assay. | Hematology | Hematology | 510(k) | Variant II beta-thalassemia Short Program Model Numbers 270-2103 (250 tests) and 270-2154 (500 tests), Bio-Rad Laboratories, Inc. in vitro diagnostic. The Variant II Hemoglobin testing system uses the principals of high performance liquid chromatography (HPLC) for the separation and determination of normal and abnormal hemoglobin. |  |  | April 08, 2014 | 2014 | Bio-Rad Laboratories, Inc. | There is an issue with the buffers which impacts the quantification of the Hemoglobin F peak. Calibration and Quality control are likely to fail if the affected lots are used. | Nonconforming Material/Component | BIO-RAD Laboratories sent an URGENT: MEDICAL DEVICE LETTER in the US on April 8. 2014 and worldwide with letters sent to customers and subsidiaries on April 14, 2014. Subsidiaries were asked to translate the letter and distribute to all affected customersCustormer were asked to complete a CUSTOMER MEDICAL DEVICE CORRECTION RESPONSE FORM and fax it to the Bio-Rad CSD Regulatory Affairs Department at the number provided to confirm that you havereceived this important information |  | 457 | 457 kits | Worldwide Distrubution ; USA INCLUDING CA, NY, GA, NC, AL, NJ, CT, WA, NY, LA, Il, MA and Internationally to China, Australia, Canada, France, Great Britain, Hong Kong, Thailand, Singapore,and Brazil. | Terminated | May 30, 2014 | 52 | 457 | Undetermined
Z-1562-2014 | 67986 | No | MQV | Resorbable calcium salt bone void filler device. | Orthopedic | Orthopedic | 510(k) | AlloFuse DBM Putty 5cc, Catalog No. 90038005 |  |  | March 25, 2014 | 2014 | AlloSource, Inc. | The donor was hemodiluted. | Process control | AlloSource sent letter via USPS Certified Mail Receipt to all affected customers dated March 25, 2014, and March 26, 2014. The letter identified the product the problem and the action needed to be taken by the customer. AlloSource requested that the distributor consignee notify subsequent consignees to determine disposition and to request return of unused product. Returned inventory will be quarantined physically and electronically upon receipt. Product initially quarantined and product returned will be destroyed following established procedures.Further questions please call (720) 873-4733 |  | 29 | 29 | Distribution US nationwide, including Michigan and a distributor in Colorado. | Terminated | August 11, 2014 | 139 | 29 | Undetermined
Z-1530-2014 | 67897 | No | CIT | Aspartate amino transferase (AST/SGOT) test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Alfa Wassermann Diagnostic Technologies, LLC. ACE AST Reagent. |  |  | March 05, 2014 | 2014 | Alfa Wassermann, Inc. | Extremely low absorbance readings. | Process control | Alfa Wassermann Diagnostic Technologies, LLC issued an "Important: Advisory Notice/"Product Response Form" dated March 5, 2014 and sent on March 6, 2014 to their customers for ACEÂż ALT Reagent Product reorder number SA1046 Lot F3604. They issued a follow-up "Important: Advisory Notice/"Product Response Form" dated March 21, 2014 and sent on March 26, 2014 to their customers for ACEÂż AST Reagent Product reorder numbers SA1047, Lot F3609, and RX1047, Lot F3610. The notification identified the product(s), problem, and actions to be taken by the customers. Customers were instructed to destroy any affected product and complete and submit the enclosed Product Response Form. For questions or concerns, contact the Alfa Wassermann Diagnostic Technologies, LLC Customer Solution Center, Toll Free: (866) 419-ALFA (2532). |  | 53 | SA1047 lot F3609: 146; RX1047 lot F3610: 53 | US Nationwide Distribution | Terminated | August 13, 2015 | 526 | 784 | Undetermined
Z-1554-2014 | 68095 | No | LYL | Soft (hydrophilic) contact lens care products. | Ophthalmic | Ophthalmic | 510(k) | Bausch &amp; Lomb, Biotrue, Multi-purpose contact lens solution, 10 fl.oz., Sterile. |  |  | April 21, 2014 | 2014 | Bausch &amp; Lomb Inc- Greenville Solutions Plant | Out of specification results for one of the two disinfectants. | Nonconforming Material/Component | Consignees were notified via registered mail starting 04/21/2014. |  | 256440 | 256,440 bottles | US Nationwide. | Terminated | January 09, 2015 | 263 | 256440 | Undetermined
Z-1563-2014 | 68114 | No | FIP | Water purification system for hemodialysis. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Mar Cor Purification, Central Water Platform (CWP), 102, 104 and 106. |  |  | April 24, 2014 | 2014 | Mar Cor Purification | Mar Cor Purification manufactured Central Water Plant (CWP) reverse osmosis system, models 102, 104 and 106. This correction has been initiated due to a potential situation with a specific lot of the water inlet solenoid valve. The potential exists that the water inlet solenoid valve could cease to function. If this occurs, CWP will stop producing reverse osmosis water. | Nonconforming Material/Component | Consignees were sent on 4/24/2014 a Mar Cor Purification "Urgent Medical Device Correction" letter. The letter described the problem and the product involved in the recall. Advised consignees to examine their CWP unit to determine if it is subject to this correction. If your unit is involved in the recall inform all staff of the potential issue. Contact Scott Nicholas, Project Coordinator / Service Trainer at 1-800-633-3038 to schedule replacement of the water inlet solenoid valve. This correction is to be carried to the user level. Consignees were requested to complete and return the "CWP Correction Notice Response Form". For questions they should contact Scott Nichols, Project Coordinator / Service Trainer at 1-800-633-3080. |  | 197 | 197 | US Nationwide Distribution - AL, AZ, AR, CA, CO, FL, GA, ID, IN, IA, KS, KY, LA, MA, MI, MN, MS, MO, MT, NH, NJ, NM, NY,NC, OH, OK, OR, PA, PR, SC, SD, TN, TX, UT, VA, WA, WV, WI. | Terminated | November 20, 2015 | 575 | 197 | Undetermined
Z-1555-2014 | 67711 | No | GAR | Nonabsorbable polyamide surgical suture. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Nonabsorbable Surgical Suture, U.S.P., Sterile - Do Not Resterilize, Deknatel. |  |  | March 11, 2014 | 2014 | Teleflex Medical | Product does not meet minimum needle attachment strength requirements; thus, the needle can become detached from the suture during use. | Employee error | Urgent Medical Device Recall Notification letters were issued on March 11, 2014 to consignees. The letter requested that customers discontinue use and quarantine any product on hand. The attached Recall Acknowledgment Form should be completed and returned to Teleflex regardless of whether affected product is on hand or not. Once received a customer service representative will contact customers with instructions on how to return the affected product. Questions should be directed to Customer Service at 1-866-246-6990. |  | 32271 | 32,271 ea. | Worldwide Distribution -- USA, including the states of AL, CA, FL, GA, IL, IN, KS, LA, ME, MA, MI, MN, MS, MO, NV, NJ, NM, NY, OH, PA, RI, SC, TN, TX, UT, VA, WV, and WI; and countries of Canada, Germany, and Ireland. | Terminated | August 14, 2017 | 1252 | 32271 | Undetermined
Z-1564-2014 | 67839 | No | CGN | Folic acid test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Siemens ADVIA Centaur Folate |  |  | March 25, 2014 | 2014 | Siemens Healthcare Diagnostics, Inc | Serum controls may go out of range low with the ADVIA Centaur ÂżSystems Folate assay and a shift in whole blood patient may result with kit lots ending in 219 and 223 | Nonconforming Material/Component | Siemens issued An Urgent Medical Device Recall (UMDR) to US customers on March 25, 2014. An Urgent Field Safety Notice (UFSN) was sent out for distribution to all customers outside the United States on March 24,2014. These notices inform the customer of reported controls failures with Folate ReadyPackÂż kit lots ending in 223. In addition the notices inform customers of a shift when moving from Folate ReadyPackÂż kit lots ending in 219 and kit lots ending in 222, 224, 225, 226 and 227. Customerswere instructed to discontinue use of ADVIA CentaurÂż Systems Folate kit lots ending in 219 and 223. ADVIA CentaurÂż systems Folate kit lots ending in 222, 224, 25, 226 and 227 are performing within specification and may continue to be used for folate interpretation. The notices should be reviewed with the facility's laboratory or medical director. Fax-back forms will be used to confirm customers received the notification. |  | 49123 | 49123 units | Worldwide Distribution-USA (nationwide) and the countries of Austria Belgium Bulgaria Cyprus Czech Republic Denmark Estonia Finland Finland France Fren.Polynesia Germany Great Britain Greece Guadeloupe Hungary Iceland Ireland Israel Italy Latvia Lithuania Luxembourg Malta Martinique Netherlands Norway Poland Portugal Qatar Reunion Romania Russian Fed. Serbia Slovakia Slovenia Spain Sweden Switzerland Turkey Unit.Arab Emir. | Terminated | April 30, 2015 | 401 | 49123 | Undetermined
Z-1568-2014 | 68031 | No | LKK |  |  | General Hospital | PMA | Refill Kit 8551/856X (includes Models 8561, 8562 and 8564) for use with Medtronic Implantable Programmable Infusion Pumps. |  |  | April 04, 2014 | 2014 | Medtronic Neuromodulation | Some refill kits in the affected lots may contain an extension tubing set that is occluded and healthcare professionals may be unable to aspirate or fill the SynchroMed pumps. There is no problem with the Lioresal Intrathecal (baclofen injection) contained in the affected drug refill kits. There have been no patient injuries beyond additional needle sticks associated with this issue. | Process control | Medtronic sent a "Medical Device Removal" letter dated April 2014 hand delivered, e-mail or faxed to affected customers. The letter described the problem and the product affected by the recall. The letter also contained the Action Required, Additional Information (Contact Medtronic Neuromodulation Technical Services at 1-800-707-0933 weekdays 7 am - 6pm Central Time) and requested consignees to complete and fax or e-mail the "Medical Device Removal Reply Form". |  | 13784 | 13,784 | Worldwide Distribution - USA Nationwide in the states of NM, ND, RI, and VT and the countries of Australia, Singapore, Thailand, Central and Eastern Europe, Western Europe, Middle East, &amp; Panama. | Terminated | November 05, 2014 | 215 | 13784 | Undetermined
Z-1546-2014 | 68038 | No | GJS | Prothrombin time test. | Hematology | Hematology | 510(k) | Alere INRatio2 PT/INR Professional Test Strips, Model Number: 99008G2. The Alere INRatio2 PT/INR Monitoring System (Professional Use), consisting of the Monitor and test strip, is used for quantitative determination of international normalized ratio (INR) in fresh, capillary whole blood to monitor the effect of warfarin therapy on clotting time by health care professionals. (in vitro diagnostic use). |  |  | April 16, 2014 | 2014 | Alere San Diego, Inc. | Test Strips may report an inaccurately low INR result. Several patients had a therapeutic or near-therapeutic INR with the Alere INRatio2 PT/INR Professional Test Strip but a significantly higher INR (outside of therapeutic range) when performed by a central laboratory. | Under Investigation by firm | Alere sent an Urgent: Medical Device Recall letter dated April 16, 2014 via fax, email or direct mail. The letter identified the affected product, problem and actions to be taken. Customers were instructed to discontinue use, return remaining inventory to Alere, and complete the provided verification form to acknowledge receipt of the recall notice. For questions contact Alere Technical Service by phone at 844-292-5373 or by E-mail at INRatio.Notification@alere.com. UPDATED: On 05/06/2014, Alere issued a press release on this recall to the consumer/public level. The press release, dated May 6, 2014, announced Alere Inc.s voluntary Class I recall in the U.S. of the Alere INRatio2 PT/INR Professional Test Strips (PN 99008G2). The press release also stated the following information: reason for recall, any adverse event reports, reason for the adverse event reports , root cause not yet determined, notified customers to immediately STOP using the product, product description with codes, contact information, FDA MedWatch information and about Alere and About the Alere INRation2 PT/INR Professional Monitoring System. |  | 99795 | Estimated 99,795 Total units | US Nationwide Distribution | Terminated | July 19, 2017 | 1190 | 99795 | Undetermined
Z-1570-2014 | 67720 | No | LKK |  |  | General Hospital | PMA | Medtronic SynchroMedÂż II Implantable Drug Infusion Pump, Model 8637-20, 8637-40. |  |  | February 26, 2014 | 2014 | Medtronic Neuromodulation | This recall provides important new information regarding overinfusion associated with the Medtronic SynchroMed II Implantable Pump. Overinfusion can result in a life-threatening overdose andcan also result in drug withdrawal due to premature emptying of the pump. Due to the low reported rate of occurrence of this issue and the inability to predict which pumps may be at risk,Medtronic is not re | Under Investigation by firm | Medtronic sent a "Urgent Medical Device Correction" letter dated March 2014. The letter was sent to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. The letter provided the Explanation of the Issue, Scope and Severity, Recommendations, and Important Guidelines.Customer visits were started by Medtronic field Representatives on February 26th, 2014.Medtronic is communicating this information to the appropriate regulatory agencies globally, including the U.S. Food and Drug Administration. We are committed to continuing to improve our product performance and services to enable you to manage your patients in a safe and effective manner. If you have questions, please contact Medtronic Neuromodulation Technical Services at 1-800-707-0933 weekdays 7am- 6pm CST. |  | 195198 | 195,198 pumps (146,435 US, 48,763 OUS) | Worldwide Distribution - All states in USA. OUS: List not provided at this time. | Open, Classified |  |  | 195198 | Undetermined
Z-1581-2014 | 68145 | No | HAW | Stereotaxic instrument. | Neurology | Neurology | 510(k) | DISP 9730950 STRL SPHERE, 1 PK INSERT |  |  | April 24, 2014 | 2014 | Medtronic Navigation, Inc. | Medtronic Navigation is recalling certain Medtronic single-use, disposable devices because there was not sufficient control to assure that returned, unused devices were appropriately examined for product and/or package damage to ensure sterility. | Nonconforming Material/Component | Domestic Urgent Medical Device Safety Notice letters were shipped on April 24 and 25, 2014. The letters indicated that a Medtronic representative will contact the hospital and arrange for a visit to collect affected products and arrange for no charge replacements or credit. Questions regarding this recall should be directed to Medtronic Navigation Technical Services at 1-800-595-9709. |  | 26674 | 26,674 total | Nationwide Distribution | Terminated | December 29, 2014 | 249 | 26674 | Undetermined
Z-1577-2014 | 68099 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Recalled product is a component of the Newport Medical Instruments e360 |  |  | April 18, 2014 | 2014 | Nellcor Puritan Bennett Inc. (dba Covidien LP) | Reports of the audible alarm not sounding when the ventilator is in an alarm state. The ventilator, as well as the visual and remote alarms, continue to operate normally. | Component design/selection | Newport Medical Instruments, Inc. (Newport), a wholly-owned subsidiary of Covidien is conducting a voluntary field corrective action via letters to be sent during the week of April 28, 2014.The letter titled "URGENT MEDICAL DEVICE FIELD SAFETY CORRECTIVE ACTION" informed customers of the recall by providing: reason for recall, product description, information on affected serial numbers regardless of software versions, issue identified, recommended actions to help mitigate risk to the patient, actions to eliminate the occurrence of this issue, important safety reminders, and contact information. Customers are to contact Technical Support Department to receive a software update kit, and then follow the instructions provided once the kit has been received. The attached verification and completion form should be completed and faxed. |  | 5765 | 5,765 | Worldwide Distribution -- USA, including the states of Arizona, California, Colorado, Connecticut, Florida, Georgia, Kentucky, Massachusetts, Maryland, Missouri, Nevada, Pennsylvania, Rhode Island, Texas, Virginia, and Washington; and, the countries of Algeria, Bangladesh, Belgium, Bolivia, Brazil, Bulgaria, Chile, China, Colombia, Dominican Republic, Ecuador, Egypt, England, France, Germany, Greece, Honduras, Hong Kong, India, Indonesia, Ireland, Israel, Japan, Kazakhstan, Kenya, Lebanon, Liechtenstein, Lithuania, Mexico, Morocco, Namibia, Nepal, Netherlands, Nicaragua, Pakistan, Panama, Paraguay, Peru, Philippines, Poland, Republic of Armenia, Republic of KOS, Russia, Saudi Arabia, Singapore, Slovenia, S. Africa, Sri Lanka, Taiwan, Thailand, Tunisia, Turkey, UAE, Uruguay, Vietnam, and Yemen. | Terminated | March 01, 2016 | 683 | 5765 | Undetermined
Z-1527-2014 | 67950 | No | ERL | Ear, nose, and throat electric or pneumatic surgical drill. | Ear Nose &amp; Throat | Ear Nose &amp; Throat | 510(k) | PolypVac 3.5mm Microdebrider; Catalog Number FG-000016; contents: 1 Microdebrider, 1 Stylet, 1 filter. Sterile. Used for the excision of polyps in the nasal passageways and in sinuses for which access has been previously established. |  |  | April 07, 2014 | 2014 | Laurimed LLC | The reciprocation mechanism in the PolypVac Microdebrider may turn off prematurely due to a component failure in the device handle. | Component design/selection | All direct accounts were notified on April 8, 2014 with an Urgent: Medical Device Recall of PolypVac Microdebrider letter, sent by first class mail. The letter identified the affected product, the reason for the recall, and stated that the firm is in the process of implementing corrections. The firm is requesting that all unused lots be returned to them. Replacement devices will sent to customers. In addition, customers are to complete the attached Lot Information sheet. Questions should be directed to bdubois@laurimed.com or 650-587-5296. |  | 58 | 58 | Distributed in the states of CA, VT, NC, DO, TX, FL, PA, OR, AZ, IL, ID, and VA. | Terminated | May 14, 2014 | 37 | 58 | Undetermined
Z-1531-2014 | 67211 | No | CIZ | Androstenedione test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | IMMULITE/IMMULITE 1000 Systems Androstenedione (Test Code AND), REF LKAO1, SMN 10381158, IVD, 100 test kit size --- Siemens Healthcare Diagnostics Products Ltd. Llanberis, Gwynedd, LL55 4EL, UK |  |  | December 20, 2013 | 2013 | Siemens Healthcare Diagnostics | Siemens Healthcare Diagnostics confirmed over-recovery of samples with concentrations greater than 5.5 ng/mL with the IMMULITE/IMMULITE 1000 and IMMULITE 2000/IMMULITE 2000 XPi Androstenedione assays (LKAO1, L2KAO2). Siemens has determined that these lots are linear up to a concentration of 5.5 ng/mL, instead of 10 ng/mL as stated in the Instructions for Use (IFU). Quality Control materials at con | Under Investigation by firm | The firm, Siemens, sent an "Urgent Medical Device Correction" letter dated December 19, 2013, to its customers informing them that Siemens Healthcare Diagnostics has confirmed over-recovery of samples with concentrations greater than 5.5 ng/mL with the IMMULITE/IMMULITE 1000 and IMMULITE 2000/IMMULITE 2000 XPi Androstenedione assays (LKAO1, L2KAO2). The distribution of the customer communication began on 12/20/13 via FedEx (US). The letter described the product, problem and actions to be taken. Customers were instructed to take the following actions: - Please discuss this communication with your medical director. - Siemens recommends the use of Quality Control materials with at least two levels having concentrations less than 5.5 ng/mL. - Patient values of &gt; 5.5 ng/mL should be verified using an alternate method. - and complete and return the Effectiveness Check questionnaire attached to the Urgent Field Safety Notice within 30 days and fax to: CUSTOMER CARE CENTER at 302-631-7597.Field service personnel were sent a Service Knowledge Base (SKB) describing the issue and instructing them how to deal with customer questions.If you have any questions, please call 1-914-524-3074 or contact your Siemens Customer Care Center or your local Siemens technical support representative. |  | 892 | 892 | Worldwide distribution: US (nationwide) and countries of: Argentina, Australia, Austria, Belgium, Brazil, Canada, China (People's Republic of), Czech.Rebuplik, Denmark, Estonia, France, Germany, Great Britain, Italy, New Zealand, Portugal, Spain, Switzerland, Turkey and Vatican City. | Open, Classified |  |  | 8633 | Undetermined
Z-1733-2014 | 68076 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Siemens Healthcare Diagnostics Vista Air Compressor. |  |  | April 14, 2014 | 2014 | Siemens Healthcare Diagnostics, Inc. | A small number of Vista Air Compressors were miswired, which if installed and powered up, could cause a fire, or become an electrical shock hazard. | Nonconforming Material/Component | Siemens sent an Urgent Medical Device Correction letter dated April 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customers.Actions to be Taken by the Customer1. If an Air Compressor has already been delivered to your laboratory, or isdelivered in the future for installation by your Siemens Healthcare CustomerService Engineer, please do not make any attempt to install the Air Compressorby yourself.2. Please notify your local Siemens Healthcare Diagnostics Customer Care Center when an Air Compressor is at your site so arrangements may be made for installation by a Siemens Healthcare Diagnostics Customer Service Engineer. Please discuss the content of this letter with your Medical Director.Please complete and return the Field Correction Effectiveness Check Form attached to this letter. Please retain this letter with your laboratory records and forward this letter to those who may have received this product.We apologize for the inconvenience this situation has caused. If you have anyquestions, please contact Siemens Customer Care Center - Technical Solutions at (800-441-9250) or your local Siemens technical support representative. |  | 4 | 4 | Worldwide Distribution: US Distribution to states of: IL, OH and TN; and country of: Italy. | Terminated | December 15, 2014 | 245 | 4 | Undetermined
Z-1497-2014 | 67895 | No | GAQ | Stainless steel suture. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Monofilament and Multifilament Non-absorbable Steel Suture; Product Codes: 0520197-2, 29-727, 29-7270M4, 29-7281M4, E29-5553M4, E29-7270M4, E29-7272M4, E29-727A, EP4013P, EP4110N, X-2235, X29-7271M2, X-5667M8 and X-6730M5. |  |  | March 11, 2014 | 2014 | Teleflex Medical | The products are being recalled because they did not meet minimum diameter requirements. | Other | Teleflex sent an Urgent Medical Device Recall Notification letter dated March 11, 2014 to all affected constomers. The letter identified the affected product, problem and actions to be taken. Customers were instructed if they have affected stock immediately discontinue use, quarantine, and return all affected product to Teleflex Medical per the instructions provided in the letter. Customers were asked to complete the enclosed Recall Acknowledgement Form and Fax it to 1-508-964-6078, Attn: Customer Service. For questions contact your local sales representative or Customer Service at 1-508-964-6042. |  | 3816 | 3,816 ea | Worldwide Distribution - CA, CO, FL, IN, KY, MA, NM, PA, Canada, Ireland and Germany. | Terminated | December 17, 2015 | 646 | 3816 | Undetermined
Z-1475-2014 | 67820 | No | LIB | Microorganism differentiation and identification device. | Microbiology | Microbiology | 510(K) Exempt | ZYM B Reagent (REF 70493). |  |  | March 06, 2014 | 2014 | BioMerieux SA | bioMerieux identified a visual defect and activity issue on the ZYM B reagent (REF 70493) and FB reagent (REF 70562). The activity issue lead to a false negative result of some Rapid ID 32 strips and API biochemical tests using the ZYM B reagent or FB reagent | Other | Letters were issued to customers on March 10, 2014. The letter contained instructions to be used if they still were using the product. The letter also included a report form which was to be returned. |  | 33853 | 33,853 | Worldwide Distribution-USA (nationwide) including DC and Puerto Rico, and the states of AL, AK, AZ, AR, CA, CO, CT, DE, FL, GA, HI, ID, IL, IN, IA, KS, KY, LA, ME, MD, MA, MI, MN, MS, MO, MT, NE, NV, NH, NJ, NM, NY, NC, ND, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VT, VA, WV, WA, WI, and WY and the country of Algeria, Angola, Argentina, Australia, Austria, Bahrain, Bangladesh, Belarus, Benin, Bolivia, Bosnia, Brazil, Brunel, Bulgaria, Burkina Faso, Cambodia, Cameroon, Canada, Central African Republic, Chile, China, Colombia, Costa Rica, Croatia, Czech Republic, Dutch Antilles, Ecuador, Egypt, El Salvador, Estonia, Erythrea, France, Gabon, Georgia, Germany, Ghana, Greece, Guatemala, Haiti, Herzegovina, Honduras, Hungary, India, Indonesia, Iraq, Israel, Italy, Japan, Jamaica, Jordan, Kazakhstan, Kenya, Korea, Kosovo, Kuwait, Latvia, Lebanon, Lithuania, Macedonia, Malaysia, Mali, Mexico, Moldavia, Morocco, Mozambique, Myanmar, Netherlands, Nigeria, Oman, Pakistan, Panama, Paraguay, Peru, Philippines, Poland, Portugal, Qatar, Russia, Saudi Arabia, Senegal, Serbia, Singapore, Slovenia, South Africa, Spain, Sweden, Thailand, Trinidad &amp; Tobago, Tunisia, Turkey, Ukraine, United Arab Emirates, United Kingdom, Uruguay, Venezuela, and Vietnam. | Terminated | March 14, 2017 | 1104 | 33853 | Undetermined
Z-1473-2014 | 67777 | No | LIB | Microorganism differentiation and identification device. | Microbiology | Microbiology | 510(K) Exempt | Oxoid Legionella Latex Test, DR0800M, containing 50 tests per box. The box contains a variety of components, one of which is Oxoid L. pneumophila Group 2-14 Latex Reagent, DR0802. The responsible firm on the label is Oxoid Ltd., Basingstoke, Hants, England. |  |  | March 17, 2014 | 2014 | Remel Inc | A reagent contained within the product may return false negative results. | Under Investigation by firm | The recalling firm initiated a recall by issuing letters dated 3/17/14 via regular mail to their customers. |  | 72 | 72/50-test boxes | Distribution was made to CA, FL, IL, MA, ME, NJ, PA, TN, TX, VA, and VT. There was no foreign/military/government distribution. | Terminated | December 11, 2014 | 269 | 74 | Undetermined
Z-1480-2014 | 67913 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | GE Healthcare, Discovery MR450, Discovery MR750, Optima MR450w, Optima MR 450w GEM, Discovery MR750w, and Discovery MR750w GEM products. Magnetic Resonance Diagnostic Device. |  |  | April 02, 2014 | 2014 | GE Healthcare, LLC | Potential safety issue with gradient coil electromechanical connections associated with GE MR Products.Gradient coil electromechanical connections may loosen causing an increase in resistance. The increased resistance can lead to increasing heat during scanning which, in turn, can cause material in the Rear Endbell enclosure to overheat, generate smoke in the scan room, and cause some localize | Process change control | Consignees were sent on 4/2/2014 a GE Healthcare "Urgent Medical Device Correction" letter dated April 2, 2014. The letter was addressed to Hospital Administrators / Risk Mangers, Radiology Department Mangers and Radiologists. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product correction and Contact Information (Fore questions, contact your local service representative ). |  | 1546 | 1,546 (704 US, 842 OUS). | Worldwide Distribution-US (nationwide) except the states of ND and RI, including DC, GU, PR, and the countries of ALGERIA, ARGENTINA, AUSTRALIA, AUSTRIA,AZERBAIJAN, BAHRAIN, BELARUS, BELGIUM, BRAZIL, CANADA, CAYMAN ISLANDS, CHILE, CHINA, CZECH REPUBLIC, DENMARK, EGYPT, FINLAND, FRANCE, FRENCH POLYNESIA, GERMANY, GREECE, HONG KONG, HUNGARY, INDIA,INDONESIA, IRELAND, ISRAEL, ITALY, JAPAN, KAZAKHSTAN, KOREA, KUWAIT, LEBANON, MARTINIQUE, MEXICO, MOROCCO, NETHERLANDS, NEW CALEDONIA, NEW ZEALAND, NORWAY, OMAN, POLAND, PUERTO RICO, QATAR, ROMANIA, RUSSIA, SAUDI ARABIA, SERBIA, SINGAPORE, SOUTH AFRICA, SPAIN, SWEDEN, SWITZERLAND, TAIWAN, THAILAND, TUNISIA, TURKEY, UNITED ARAB EMIRATES, UNITED KINGDOM, URUGUAY, UZBEKISTAN, VENEZUELA. | Terminated | April 08, 2015 | 371 | 1546 | Undetermined
Z-1479-2014 | 67722 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | LifeShield LATEX-FREE HemoSet 100 mL Burette I.V. PlumSet, Nonvented, 105 inch with Prepierced Reseals, OPTION-LOK and Inline Dual Channel Cassette. |  |  | February 25, 2014 | 2014 | Hospira Inc. | Hospira identified an incorrect lower lid (set component) had been supplied and was being used during the manufacturing process of list number 11241-03, a HemoSet containing a burette and blood filter, intended for infusion of blood/blood products. In a gravity delivery, the correct lower lid dispenses 15 drops per mL and the incorrect lower lid found dispenses 10 drops per mL. | Mixed-up of materials/components | Hospira sent a URGENT MEDICAL DEVICE RECALL letter dated February 25, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. 1) check their inventory and immediately quarantine any affected product; 2) complete the attached reply form and return it to the fax number or e-mail address on the form; 3) inform potential users within the direct account's organization of this product of the recall; 4) return affected products to Stericycle using labels provided with the notification; and 5) follow the instructions on the return label for returning products. Additional labels and guidance on returning products can be obtained by calling Stericycle at 1-888-240-4282 (Monday - Friday, 8:00 AM - 5:00 PM ET). If the recalled products were further distributed, the recall notification documents should be forwarded to those accounts. Those accounts can contact Stericycle at 1-888-240-4282 (Monday - Friday, 8:00 AM - 5:00 PM ET) to obtain a reply form and guidance on the return process. |  | 3893 | 3,893 sets | US Distribution including the states of FL, IA,IN, KS, KY, ME, MO, MT, PA, TN, TX, UT, VA and WA. | Open, Classified |  |  | 3893 | Undetermined
Z-1485-2014 | 68041 | No | DQX | Catheter guide wire. | Cardiovascular | Cardiovascular | 510(k) | Merit Laureate Hydrophilic Guide Wire, Catalog No. LWSTFS35260EX, Straight Tip Stiff Shaft, 0.035" (0.89mm), 260 cm (102"). |  |  | April 14, 2014 | 2014 | Merit Medical Systems, Inc. | Discrepancy between the carton and unit labeling for the Merit Laureate Hydrophilic 0.035" 260 cm Guide Wire, catalog number LWSTFS35260EX, which has a straight tip and a stiff shaft. This lot was partially mixed with another lot of Laureate0.035" 260 cm Guide Wire, catalog number LWSTDA35260EX, which has an angled tip, standard shaft and Chinese unit labeling. | Process control | Customers were notified on 4/14/14 via certified letter with instructions to identify product, discontinue use, and return products. |  | 173 | 173 | Worldwide Distribution: USA Nationwide and China. No military/govt/VA consignees. | Terminated | July 17, 2014 | 94 | 359 | Undetermined
Z-1487-2014 | 67959 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | iLab Ultrasound Imaging System; iLab 3.0 (Polaris 1.0) software. Used for ultrasound examinations of intravascular pathology. |  |  | April 03, 2014 | 2014 | Boston Scientific Corporation | Specific iLab Ultrasound Imaging Systems using newest 3.0 (Polaris 1.0) software change the pullback speed from the default setting of 0.5 mm/sec to 1.0 mm/sec, the Longview distance measurement may display an inaccurate value that is double the actual distance. | Nonconforming Material/Component | Boston Scientific sent an Urgent Medical Device Field Correction letter to all affected customers via hand delivery by BSC sales representatives. The letter identifed the affected product, problem and actions to be taken. Customers were asked to immediately post the Field Correction information in a visible location near the system(s) to ensure it is easily accessible to all users of the device. For question contact your local Sales Representative. |  | 9 | 9 units | USA Nationwide Distribution in the states of NY, NC, DC, FL, PA | Terminated | March 18, 2015 | 349 | 9 | Undetermined
Z-1471-2014 | 67466 | No | DSQ |  |  | Cardiovascular | PMA | Heartware HVAD Pump Implant Kit (Heartware Ventricular Assist Device) Catalog Numbers: 1100, 1101, 1102, 1103, 1104, and 1205. For use as a bridge-to-cardiac transplantation in patients who are at risk of death from refractory end-stage left ventricular heart failure. The HeartWare VAS is designed for in-hospital and out-of-hospital settings, including transportation via fix-wing aircraft or helicopter. |  |  | December 06, 2013 | 2013 | HeartWare Inc | The housing to the Pump's driveline connector became partially or fully separated from the front portion of the driveline connector. | Process control | HeartWare, Inc. sent an Urgent Medical Device Correction letter to all their hospitals and doctors on December 12, 2013 instructing that at each routine clinic visit, they should inspect the patient's driveline connector as described in the attached Technical Bulletin. For already implanted patients, arrange a follow up visit at the earliest convenience to check patient's driveline connector. They were asked to sign and return the Acknowledgement Form to HeartWare. For questions, contact your HeartWare Representative or HeartWare Customer Service at 1-877-367-4823. |  | 4586 | 4,586 | Worldwide Distribution - USA Nationwide in the states of MN, IL, FL, WI, MO, MA, NC, CA, GA, AL, CO, TX, OH, NY, VA, UT, KY, MD, AZ, SC, IN, PA, DE, NJ, LA, OR, WA, MI, IA, TN and DC and countries: Argentina, Australia, Austria, Belgium, Brazil, Canada, Chile, Croatia, Czech Republic, France, Germany, Greece, Hong Kong, India, Israel, Italy, Japan, Lebanon, Lithuania, Luxembourg, Malaysia, Netherlands, New Zealand, Norway, Poland, Singapore, Slovakia, South Africa, Spain, Switzerland, Turkey and United Kingdom. | Terminated | November 03, 2016 | 1063 | 4586 | Undetermined
Z-1490-2014 | 67509 | No | JFY | Creatinine test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | VITROS Chemistry Products CDM PROM (Calibration Data Module) No. 0168, Catalog Number/REF 199 9077, used in conjunction with: VITROS Chemistry Products CRSC DT Slides, Generation (GEN) 82, REF 6802721, and |  |  | January 29, 2014 | 2014 | Ortho-Clinical Diagnostics | Customers were unable to calibrate, or use previously calibrated, VITROS DT CRSC slides (REF/Product Code 6802721), GEN 82, on the VITROS DT II System, when using Calibration Data Module (CDM) Rev. 168 (REF/Product Code 199 9077), as CDM 168 does not contain calibration parameters for DT CRSC. | Process control | Ortho Clinical Diagnostics sent an Urgent Product Correction Notification Letters dated January 29, 2014, via FedEx overnight courier and/or US Postal Service Priority Mail (for PO Boxes only) to US consignees to inform them of the issue. Customers were advised to Inspect their inventory to determine if they had VITROS CRSC DT Slides, GEN 82. - To use their current inventory of GEN 82 until the replacement order arrives providing that they have CDM PROM 0167 installed on their system and calibrate using VITROS DT Calibrator Kit, Lot 42. - If product replacement is required, discard their remaining inventory of VITROS CRSC DT Slides, GEN 82 upon receipt of their replacement order. - Complete and return the Confirmation of Receipt Form by February 7, 2014, by fax to 1-888-557-3759 or 1-585-453-04110, or email to: ConfirmationAdmin@its.jnj.com. Indicate on the form if they have any inventory of the affected lot that will require replacement. - Forward this information if they have distributed this product outside of your facility. Distributors were also instructed to identify all customers and any other branches who were shipped VITROS CRSC DT Slides, GEN 82 from their facility and send a copy of the Urgent Product Correction Notification Letter to inform them of the issue. Foreign affiliates were informed by e-mail on January 22, 2014, of the issue and instructed to notify their consignees of the issue and required actions. |  | 5056 | US: 284 units, Ex-US: 4772 units | Worldwide Distribution - US Distribution including the states of AR, FL, IA, IL, KY, MI, MN, ND, NY and VA, and the countries of Australia, Brazil, Canada, Chile, China, Colombia, England, France, Germany, India, Italy, Japan, Mexico, Panama, Singapore, Spain, and Venezuela. | Open, Classified |  |  | 5056 | Undetermined
Z-1505-2014 | 67568 | No | KPR | Stationary x-ray system. | Radiology | Radiology | 510(k) | MULTIX FUSION. |  |  | February 13, 2014 | 2014 | Siemens Medical Solutions USA, Inc | It was observed that the corrugated hose holder above the 3D rail on the Multix Fusion system became loose. According to the Installation Instructions the corrugated hose holder shall be afixed with transversal carriage, with Loctite 243 and Torque 10nm during system installation. Siemens initiated the update instruction to verify and secure the hose holder with Loctite 243. A PCB D802 Rev. 00 b | Process change control | Affected customers were initially notified with a recall letter on/about 2/12/14 and were sent Update Instructions to affected customers. |  | 12 | 12 | Distributed in the states of AK, PA, KY, OH, WI, MI, FL, and NY. | Terminated | March 16, 2015 | 396 | 12 | Undetermined
Z-1493-2014 | 67746 | No | FYD | Air-handling apparatus for a surgical operating room. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Stryker PenAdapt"; Rx Only, Sterile. |  |  | March 26, 2014 | 2014 | Stryker Instruments Div. of Stryker Corporation | During packaging verification testing, a failure occurred due to cuts in the packaging potentially causing a breach in the sterile barrier. Additional testing identified microscopic aluminum metal shavings and foreign particles estimated in size to be between 0.01inches x 0.015 inches to 0.03 inches x 0.04 inches that were introduced during the manufacturing process and were sterilized within the | Material/Component Contamination | On March 26, 2014, Stryker Instruments sent URGENT MEDICAL DEVICE RECALL NOTIFICATION to their consignees via mail and email. Notification included product description, reason for recall, risk to health, and actions to be taken by the Customer/User. For questions regarding this recall please contact Stryker Instruments:Monday-Friday 8am-5pm (EST)Angela Ragainis/Kelly Jo Whipple269-389-4354 / 269-389-2921angela.ragainis@stryker.com / kellyjo.whipple@stryker.comCustomers will be asked to sign the Business Reply Form stating they have read the Notification and followed all instructions within. Customers will return affected product to Stericycle for destruction using a return label provided by Stryker or Stericycle. Upon receipt of affected product a credit will be issued to customer account.In addition, facilities are being asked to research their records for patient issues that may have been associated with the use of a recalled device, such as infection and/or inflammation. |  | 35310 | 35,310 unexpired units | Worldwide Distribution-USA (nationwide) including DC and the states of AK, AL, AR, AZ, CA, CO, CT, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV,and WY, and the countries of Austria, Canada, Sweden, England, UAE, Netherlands, Spain, and Hong Kong. | Terminated | December 11, 2014 | 260 | 35310 | Undetermined
Z-1522-2014 | 67902 | No | LFL |  |  | General &amp; Plastic Surgery | 510(k) | Integra CUSA Excel Disposable Wrench. Catalog Number C5601. The CUSA Excel+ Sterile 23 kHz Torque Wrench is an accessory used with the CUSA Excel+ System ultrasonic surgical instrument. The torque wrench is used to properly secure the CUSA tip to the CUSA handpiece in the sterile field. |  |  | March 10, 2014 | 2014 | Integra LifeSciences Corp. | Some 23 kHz Torque Wrenches were incorrectly assembled and contained a 36 kHz Hex Wrench Insert instead of the 23 kHz Hex Wrench Insert. | Under Investigation by firm | Integra sent an Urgent: Voluntary Medical Device Recall letter dated March 10, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.The written notifications were delivered to US consignees by traceable courier service and non-US consignees by traceable email or facsimile. Integra is asking that you do the following:1. Review your unused inventory of 23 kHz CUSA Sterile Torque Wrenches and determine if they are on the affected lot number list on the next page and if they are still in the box.2. Complete the attached form.- If you do not have an affected lot number on the list (see page 2), YOU CAN CONTINUE TO USE THE WRENCHES. Check the box I do not have any affected products on the list- If you do have product(s) on the affected lot number list, AND THEY ARE STILL IN THE BOX, determine if the box also has a green dot sticker as in the photo below. If so, YOU CAN CONTINUE TO USE THE WRENCHES. Check the box I do not have any affected products on the list on page 2.3. Complete the other information on the form if you do or do not have affected products, return the form by email or fax as indicated on the form. Keep a copy of the form for your records.When your form is received, if you have noted you have affected products Customer Service will contact you and provide an RMA number and directions to return the product, as well as input an order to replace the quantity you indicated on the form.Should you have any questions regarding these instructions, please contact Customer Service at 1-855-532-1723.We apologize for any inconvenience this may cause and thank you for your cooperation in this effort. |  | 23904 | 3,984 devices packaged as 6 devices / box (664 boxes) | Worldwide Distribution - USA (nationwide) including the countries of Australia, Canada, and Switzerland. | Terminated | February 13, 2015 | 340 | 23904 | Undetermined
Z-1496-2014 | 67795 | No | LPH | Hip joint metal/polymer/metal semi-constrained porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | TRABECULAR METAL MODULAR ACETABULAR SYSTEM, Shell with Cluster Holes, Porous 60mm. Catalog number 00-6202-060-22; Hip implant component. |  |  | April 01, 2014 | 2014 | Zimmer, Inc. | The shell was missing the locking ring. | Nonconforming Material/Component | The firm, Zimmer, sent an "URGENT MEDICAL DEVICE RECALL-LOT SPECIFIC" notification dated April 2014 to their consignees. The notification included product description, product issue, health risks, and consignee responsibilities (including directions on how to return the affected product). The consignees were instructed to locate all products from lot 62490588 and quarantine them immediately; carry out a physical count of all affected product; complete and return the Inventory Return Certification Form via email to: corporatequality.postmarket@zimmer.com; return the recalled product, and notify Zimmer of any hospitals that you have further distributed the affected product and of surgeons that have implanted the product. Consignees with questions pertaining to the recall are directed to call the customer call center at 1-877-946-2761 between 8:00 am and 5:00pm EST. |  | 17 | 17 units | Worldwide Distribution: US (nationwide) in states of: US: GA, WI, NJ, CA, NY, NC, PA, FL, CT, ME, TX, and IL; and countries of: Australia and Canada. | Terminated | September 11, 2014 | 163 | 17 | Undetermined
Z-1503-2014 | 67859 | No | MBQ |  |  | Physical Medicine | Contact ODE | Digital Diagnost, stationary radiographic system; Models: 712020/22/25/26/50/52/54/55/57/60/62/80/81/82/85 |  |  | March 17, 2014 | 2014 | Philips Medical Systems, Inc. | The footboard of Patient support for stitching can be fixed in upper position. When the fix clamp (hook) is not correctly adjusted into its correct position, the footboard is not securely held. | Under Investigation by firm | On March 17, 2014, consignees were notified of the recall via letter. The letter explains what the issue is, actions firms can take to avoid or minimize the occurrence of the issue, and the actions planned by Philips to correct the issue.For further information or support concerning this issue, please contact your local Philips Representative. 1-800-722-9377 |  | 985 | 985 devices | Nationwide: US: AL, AR, AZ, CA, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, MA, MD, ME, MN, MO, MS, MT, NC, ND, NH, NJ, NM, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VT, WY / Worldwide including Canada &amp; Mexico | Open, Classified |  |  | 985 | Undetermined
Z-1506-2014 | 67648 | No | KNT | Gastrointestinal tube and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Sengstaken Tube and Sengstaken-Blakemore Tube. Teleflex product codes 204800, 204802, and 204803, a gastrointestinal tube and accessories, consisting of flexible or semi-rigid tubing used for instilling fluids into, withdrawing fluids from, splinting, or suppressing bleeding of the alimentary tract. |  |  | March 04, 2014 | 2014 | Teleflex Medical | Sterility of the product cannot be guaranteed. | Other | Letters were issued on 3/4/2014 to consignees asking them to immediately discontinue use and quarantine any products with the recalled catalog numbers. Distributors were also asked to conduct a sub-recall. The recall letters included a Recall Acknowledgment Form which was to be faxed back to Teleflex. |  | 117801 | 117,801 ea. | Worldwide Distribution-USA (nationwide) including the states of AL, AK, AR, CA, CO, FL, GA, ID, IL, IA, KY, LA, MD, MA, MI, MN, MS, MO, NE, NH, NJ, NM, NY, NC, OH, PA, SC, TN, TX, UT, VT, VA, WA, WV, and WI, and the countries of Algeria, Andorra, Armenia, Argentina, Austria, Australia, Azerbaijan, Belarus, Bermuda, Bahamas, Boliva, Brazil, Canada, Chile, China, Costa Rica, Cyprus, Czechoslovakia, Germany Guatemala, Denmark, Estonia, Egypt, Finland, France, Great Britain, Georgia, Guadeloupe, Greece, Guyana, Curacao, Hungary, Israel, India, Italy, Jordan, Japan, Kuwait, Kazakhstan, Luxembourg, Latvia, Libya, Madagascar, Macedonia, Martinique, Malta, Mauritius, Mexico, Nicaragua, Nigeria, Netherlands, Norway, Oman, Peru, French Polynesia, Poland, Portugal, San Marino, Singapore, Spain, Suriname, San Salvador, Thailand, Tunisia, Turkey, Uruguay, Uzbekistan, Wallis &amp; Futuna Islands, Zambia. | Terminated | January 28, 2016 | 695 | 117801 | Undetermined
Z-1509-2014 | 67972 | No | NEW | Absorbable polydioxanone surgical suture. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Monodek Violet Synthetic Absorbable Surgical Sutures, MF 0 TC43/HR26 48 |  |  | March 11, 2014 | 2014 | Teleflex Medical | Product does not meet minimum knot tensile strength requirements. | Under Investigation by firm | Consignees were notified by an Urgent Medical Device Recall Notification letter, dated 3/11/2014. The letter identified the affected product and the reason for recall. Customers were instructed to immediately discontinue use of and quarantine any affected product in stock. The affected product is to be returned; and the Recall Acknowledgement Form should be completed and faxed to the number provided regardless of whether customers have affected product in stock. Questions should be directed to a local sales rep or Customer Service at 1-866-246-6990. |  | 3072 | 3,072 ea. | Worldwide Distribution -- USA, including the state of MA, and the country of Germany | Open, Classified |  |  | 3072 | Undetermined
Z-1516-2014 | 67885 | No | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | The Acumed Tension Band Pin is a fixation pin which is .0625" in diameter and is available in lengths between 35mm and 90mm. It is to be used in conjunction with commonly used orthopedic wire. |  |  | March 26, 2014 | 2014 | Acumed LLC | Certain lots of Acumed Tension Band Pins are being recalled because the implantable product was manufactured with a grade of stainless steel that is not within specifications. | Component change control | Acumed sent an Urgent Medical Device Recall letter dated March 26, 2014 to all affected customers. The letter identified the affected products, problem and actions to be taken. Customers are instructed to immediately stop using and/or distributing all lots identified in the letter. Return affected products as instructed in the letter. For questions contact Acumed Business Service via email at BusinessService@acumed.net or via phone at 877-627-9957. |  | 6128 | 2544 unit in the US and 3584 units outside the US | US Nationwide Distribution - in the states of CA, CO, FL, GA, LA, MI, MN, MO, OH, OR, PA, PR, TX, UT, VA, WA and WY.Product was distributed outside the US to: Australia, China, Figi, France, Great Britain, Italy, Japan, Korea, Maylasia, Sweden, South Africa, Spain. | Terminated | February 04, 2015 | 315 | 6563 | Undetermined
Z-1514-2014 | 68088 | No | JAI | Powered radiation therapy patient support assembly. | Radiology | Radiology | 510(k) | HexaPOD evo RT Couchtop. |  |  | April 17, 2014 | 2014 | Elekta, Inc. | If screws are not properly tightened with a torque wrench they may become loose over time. | Component design/selection | Corrective Action #1 - Important Field Safety Notice 618 01 302 002 was issued April 2014. The Notice informs end users about the problem and the potential clinical impact. It also give a recommendation how the risk can be mitigated - inspect the screws and if loose contact Elekta Service to tighten the screws with 20 Nm. The notice also advises the customer that a modification kit will be released that introduces an improved screw solution. The solution will be implemented on all affected products when available. Corrective Action #2 Permanent Solution - Replace Current Screws - An improved screw has been designed and will be available for Elekta Field Service organization to replace on all affected products at the end of April. The new design solution provides a very long used thread depth, exceeding the standard recommended thread-depths by far. This design will prevent the screws from loosening. |  | 266 | 266 | Worldwide Distribution-USA including DC and the states of AZ, CA, CT, FL, GA, IL, HI, KY, MD, MA, MI, MS, MO, NJ, NY, NH, NC, ND, OH, OK, OR, PA, SC, TX, VA, WA, WI, and WY, and the countries of Algeria, Australia, Austria, Belgium, Brazil, Canada, China, Colombia, Czech Republic, Denmark, Finland, France, Germany, Greece, Hong Kong, India, Israel, Italy, Japan, Kazakhstan, Mexico, Netherlands, New Zealand, Peru, Poland, Portugal, Russia, Saudi Arabia, South Korea, Span, Sweden, Switzerland, Taiwan, Thailand, Turkey, and United Kingdom. | Terminated | January 29, 2015 | 287 | 266 | Undetermined
Z-1518-2014 | 67210 | No | JHR | Immunoglobulins A, G, M, D, and E immunological test system. | Immunology | Immunology | 510(k) | IMMULITE 2000 /IMMULITE 2000 XPi Systems TIE Total IgE, REF L2KIE2, SMN 10380873, and L2KIE6, SMN 10380872, IVD --- Siemens Healthcare Diagnostics Products Ltd. |  |  | December 20, 2013 | 2013 | Siemens Healthcare Diagnostics | Siemens confirmed an overall average positive bias of 23% against the WHO 2nd IRP 75/502 with the IMMULITE Systems Total IgE assays. | Under Investigation by firm | Urgent Medical Device Correction Letters (dated 12/20/13) and effectiveness check response forms were sent to customers on 12/20/13 via FedEx to inform them of the issue. Customers were requested to complete and return the Effectiveness Check questionnaire attached to the Urgent Medical Device Correction letter within thirty (30) days. Field service personnel were sent a support document describing the issue and instructing them how to deal with customer questions. |  | 18489 | 18,489 units | Worldwide distribution. | Open, Classified |  |  | 24442 | Undetermined
Z-1521-2014 | 67786 | applicator, absorbent tipped, sterile | KXG | Absorbent tipped applicator. | General Hospital | General Hospital | 510(K) Exempt | Fisherbrand Sterile Swabs, Calcium Alginate Fiber Tipped Aluminum Applicator Swab, .025 and .035 in Diameter, all lots, 100 per pack |  |  | February 05, 2014 | 2014 | Fisher Scientific Co | A customer reported the presence of an organism in a lot of Calcium Alginate Fiber Tipped Wood Applicator Swab. | Process control | Fisher Scientific sent an Urgent: Medical Device Recall/Return letter dated March 18, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. ACTIONS TO BE TAKEN BY THE USER OR DISTRIBUTOR:Fisher Scientific requests that you discontinue use and distribution of all products referenced above. We also request that you return all product (opened and unopened) as soon as possible. A credit will be issued for returned product.User: Fisher Scientific requests that you complete the Acknowledgment Form enclosed with this letter, even if you do not have the above mentioned product(s) in your inventory, within three (3) working days.Distributor: Our records indicate that you may have purchased one or more of the above products for re-sale to your customers. You should complete the attached Acknowledgment Form for inventory you have received and/or which is still in stock. Fisher Scientific requests that you complete and return the Acknowledgment Form enclosed with this letter, even if you do not have the above mentioned catalog numbers in your inventory, within three (3) working days. In addition, please contact your affected customer base to advise them of the expanded recall, by providing them with a copy of this letter. You should insert your contact information, email and fax numbers in the Acknowledgment Form and request that your customer return the Acknowledgment Form to you, and you notify us of the information conveyed in such Acknowledgment Form(s).ACTION BY FISHER SCIENTIFIC:Fisher Scientific has no inventory of the affected products as it was discontinued in May 2013. However, we do have alternative products that we can recommend as replacements. Please contact customer service for information on product alternatives.We appreciate your immediate attention to this expanded voluntary recall. By returning the attached Acknowledgment Form you wil |  | 7430 | 7430 | Nationwide Distribution | Terminated | August 10, 2015 | 551 | 7430 | Undetermined
Z-1520-2014 | 67883 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | VITROS Chemistry Products FS Diluent Pack 3, REF 680 1754, 3 x 45 mL, (D1)/15 mL (D2), IVD --- Ortho-Clinical Diagnostics, Inc. Rochester, NY, 14626 |  |  | March 26, 2014 | 2014 | Ortho-Clinical Diagnostics | Ortho Clinical Diagnostics (OCD) confirmed the potential for lower than expected C-Reactive Protein (CRP) results only to be generated when using VITROS Chemistry Products Specialty Diluent (Lot F3168) or VITROS Chemistry Products FS Diluent Pack 3 (Lot 01-3266). The VITROS Instructions for Use states that if the if a C-reactive protein concentration exceeds the systems measuring (reportable or dynamic) range, dilute the sample with VITROS Specialty Diluent or a patient sample containing a low concentration of CRP. An initial threefold dilution is recommended. Internal testing determined that results using VITROS CRP Slides have the potential for approximately 51% bias using a 3x dilution factor. | Other | On 3/26/2014, Urgent Product Correction Notification Letters were sent via FedEx overnight courier and/or US Postal Service Priority Mail (for PO Boxes only) to US consignees (OCD Direct, Drop-Ship and US Federal Government) to inform them of the issue and advise them to discontinue use of the affected lot(s), discard the affected lots and to inform their Laboratory Medical Director if they suspect that a previously reported result may have been affected. Replacement product will be sent for the product discarded. J&J; Foreign affiliate consignees were notified by e-mail informing them of the issue on 3/26/2014. |  | 1042 | 1042 distributed (220 went to warehouse in France) | Nationwide, Puerto Rico, and foreign distribution to Australia, Brazil, Canada, Chile, China, Colombia, England, France, Germany, India, Italy, Japan, Mexico, Panama, Singapore, Spain and Venezuela. | Terminated | March 27, 2017 | 1097 | 2698 | Undetermined
Z-1608-2014 | 68026 | No | BSP | Anesthesia conduction needle. | Anesthesiology | Anesthesiology | 510(k) | Ultrasound needles, 22G (0.7 mm) x 50mm, packaged individually, 10 needles/box, item number REF USC050-22. Needles for peripheral nerve blocking under ultrasound guidance. |  |  | March 19, 2014 | 2014 | RM Temena GmbH | Ultrasound needles may be defective and leak at the needle hub. | Process design | Temena sent an Urgent Recall letter dated March 24, 2014, to the affected customers. The letter identified the product the problem and the action needed to be taken by the customer.IMPORTANT INSTRUCTIONS:1. Ensure that the devices questions will not be applied to humans.2. Opened packages shall not be used any more.3. Please send us back the request of products (including opened packages)4. Inform your distribution partners / further resellers about this recall to achieve a 100% recall of all unused products.5.Please add a copy of our invoice / respectively of our delivery notes, to carry out a cost reimbursement or receive compensation delivery or a credit note.6. Please use attached form to answer Temena SAS.Further questions please call: +33(0)1.30.86.05.30We sincerely apologize for the inconvenience this may have caused by this product recall we thank you in advance for your support |  | 7370 | 7370 needles | US Distribution including the states of MD and IL. | Completed |  |  | 7370 | Undetermined
Z-1612-2014 | 66775 | No | BYZ | Vacuum-powered body fluid suction apparatus. | General Hospital | General Hospital | 510(K) Exempt | Mucus Specimen Trap 40 cc, Specimen Trap 40 cc with additional transport cap; Product number DYND44140 (pack of 50), Product number DYND44140H (individual). Mucus Specimen Trap 80 cc, Specimen Trap 80 cc w/ additional transport cap; Product number DYND44180 (pack of 50). Used to collect mucus specimens during suction of fluid from the oral cavity, nose-throat area and/or bronchi of a patient. |  |  | November 04, 2013 | 2013 | Medline Industries Inc | There is a potential that the specified mucus specimen traps are not sterile. Use of the affected products may result in the contamination of patient specimens. | Process design | Medline sent a Recall Notification Letters dated November 4, 2013 to all affected customers.The letters included instructions for customers to: check their stock of the affected products; complete and return the enclosed response form even if none of the recalled products are in inventory; Medline will provide customers with a Return Goods Authorization upon receipt of the response form and credit will be issued upon receipt of the returned recalled products; and, notify any customers who may have received the recalled products. Customers who have any questions can contact Medline Industries at 866-359-1704. |  | 1610 | 1,610 mucus specimen traps | US Nationwide Distribution in the states of CA, MA, MD, MI, MN, NH, NJ, NY, OH, PA, TX, and UT. | Completed |  |  | 1610 | Undetermined
Z-1610-2014 | 68029 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | PS-C Femoral Nonporous Rt Sz 6, Model Number 161-1604 Component of the Balanced Knee System and Balanced Knee Revision System, for total knee arthroplasty procedures. |  |  | April 01, 2014 | 2014 | Ortho Development Corporation | A femoral component containing pegs was found in a box for the pegless version. | Process control | Ortho notified customers via email and phone calls on April 1, 2014. Customers were asked to return product in inventory. |  | 5 | 5 | US Nationwide Distribution in the states of NV, and AZ. | Terminated | July 14, 2014 | 104 | 5 | Undetermined
Z-1714-2014 | 68363 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | Toshiba Infinix X-Ray Interventional System INFX-8000V/UU. An X-Ray system that is capable of radiographic and fluoroscopic studies used in an interventional setting. |  |  | May 12, 2014 | 2014 | Toshiba American Medical Systems Inc | Although the Flat Panel Detector (FPD) touch sensor is activated during use, the error message is displayed and the audible signal continues sounding during the table movement of longitudinal tilt, the movement of longitudinal tilt does not stop. | Other | Toshiba America Medical Systems, Inc. initiated this recall by sending recall notifications via USPS return receipt mail.The recall notification letter, dated May 12, 2014, titled "URGENT: MEDICAL DEVICE CORRECTION", informed customers of the following: device information, serial number, SID, reason for reason, purpose of the notification, corrective action, request to customers, contact information, and customer reply form.Customers with questions regarding this letter were instructed to call (800) 421-1968 or contact their local Toshiba Representative at (800) 521-1968.For questions regarding this recall call 714-730-5000. |  | 2 | 2 | Nationwide Distribution including CA and IN | Terminated | October 02, 2014 | 143 | 2 | Undetermined
Z-1675-2014 | 68159 | No | GWM | Intracranial pressure monitoring device. | Neurology | Neurology | 510(k) | Integra Licox Test Kit REF BC10R Rx Only |  |  | April 16, 2014 | 2014 | Integra LifeSciences Corp. | Integra LifeSciences has identified through a small number of complaints that some lot numbers of Licox Test Kits (Catalog Number BC10R) may provide erroneous "calibration test failure" message when they are used with the Licox PtO2 Monitor (LCSO2, or may provide an erroneous out of range result with the Licox CMP Brain Tissue Oxygen and Temperature Monitor (AC31). | Under Investigation by firm | Integra sent an "Urgent Voluntary Medical Device Recall/Return Response form notifications dated April 16, 2014 to US customers via traceable courrier service and to non-US customers via traceable email or facsimile. The customer notification advises them of the nature of the issue and the potential for adverse patient consequences; advises them to review their inventory and if they have BC10R Test Kit Resistors with an affected lot number to immediately stop using it; requests they identify and report to Integra if they do or do not have any affected BC10R affected lot numbers; and advises them how to return any affected products in their possession. Customers with questions were instructed to call 1-855-532-1723. |  | 49 | 49 test probes | Worldwide Distribution - USA including Illinois, Florida, Maryland, Massachusetts, North Carolina, Pennsylvania, South Dakota and Internationally to Australia, Austria, Belgium, Chile, Finland, Germany, Great Britain, Greece, Israel, Italy and Russia. | Terminated | October 24, 2014 | 191 | 49 | Undetermined
Z-1673-2014 | 68293 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Medrad Continuum MR Infusion system- non-wireless system |  |  | April 03, 2014 | 2014 | Bayer Corp | There is a potential safety risk associated with potentially damaged locking pins within the bracket of non-wireless versions of Continuum systems shipped prior to March 2008. | Component design/selection | Bayer sent an Urgent Medical Device Recall and Removal Letters dated April 3, 2014, and April 30, 2014, were sent to direct accounts to notify them of the bracket issue and to provide additional instructions for testing and/or returning the brackets.As a critical medical needs site currently using Continuum with a recalled bracket you have two options:1. Immediately return your Continuum(s) for refund.2. Take actions to ensure continued safe use of your system associated with the bracket recall. You will still be eligible for a refund when you return the system no later than June 30, 2015.When you are ready to return your Continuum(s), please refer to:Âż Enclosure 2: Disassembly Instructions.Âż Enclosure 3: Continuum Return Form.We sincerely regret the inconvenience and impact of the recalls and subsequent removal of Continuum from the market. Should you have questions regarding the pro-rated refund program, please call 877-229-3767. |  | 1347 | 1347 | Worldwide Distribution - USA (nationwide) including the countries of Australia, Brazil, Belgium, Canada,Germany, Singapore, South Africa, Poland, Mexico and UAE. | Terminated | April 16, 2015 | 378 | 1347 | Undetermined
Z-1681-2014 | 68006 | No | MQP | Spinal intervertebral body fixation orthosis. | Orthopedic | Orthopedic | 510(k) | Synthes XRL Medium Trial - Central Body, an instrument used as part of the Synthes XRL System, a vertebral body replacement device. For spinal surgery. The Synthes XRL device is intended to be used with the Synthes supplemental internal fixation systems (e.g., USS, including MATRIX, Pangea, and TSLP). |  |  | April 08, 2014 | 2014 | Synthes, Inc. | Two lots of the Synthes XRL Medium Trial-Central Body device were incorrectly etched with the incorrect dimensions for height. | Process control | Synthes sent an Urgent Notice: Medical Device Recall letter dated April 8, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Action Required:Synthes has record that your facility has the product(s) subject to this recall. Synthes asks that you review your inventory and immediately remove the affected lots from stock.If you DO have any of the identified devices, please take the following steps: Contact your local Synthes Spine Sales Consultant to return any inventory located within your facility.Complete the Verification Section at the end of this letter by checking the appropriate box indicating affected product has been located. Also, please indicate the number of devices found. Please include your name, title, telephone number and signature in the spaces provided.Return the Verification Form (page 4 of this letter) to: Fax: 877-270-9246 orScan/email: Synthes3041@stericycle.comIf you DO NOT have the identified product, please take the following steps:Complete the attached Verification Section at the end of this letter by checking the appropriate box indication that no affected product has been located. Please include your name, title, telephone number and signature in the spaces provided this return documentation acknowledges your receipt of medical device removal information.Return the documents to Synthes by: Fax: 877-270-9246 or Scan/email: Synthes3041@stericycle.comNote: If the Verification Form is answered on behalf of more than one facility and/or individual, please clearly indicate the name and address of the facility and/or individual on page 4 of the notification.If you have any questions, please call 610-719-5450 or contact your Synthes Sales Consultant. |  | 3 | 3 | US Distribution including the states of FL, NJ and MN. | Terminated | September 10, 2015 | 520 | 3 | Undetermined
Z-1699-2014 | 68080 | No | LXB | Hearing Aid. | Ear Nose &amp; Throat | Ear Nose &amp; Throat | 510(k) | Cochlear Baha Soft tissue gauge 6mm. |  |  | April 09, 2014 | 2014 | Cochlear Americas Inc. | Cochlear America is recalling Cochlear Baha Soft Tissue Gauge due to possible manufacturing imperfection in the smooth surface of the metal that could potentially leave residual metal on the patient's skin. | Process control | Customers were notified via letter on 4/17/14. This communication included an explanation of the recall and instructions to quarantine and return products. |  | 288 | 288 | Nationwide Distribution. | Terminated | October 06, 2014 | 180 | 288 | Undetermined
Z-1680-2014 | 68305 | No | OEI | Powered suction pump. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Wound Drain - Label reads in part "Elutia ***Hubless Coated Silicone Surgical Wound Drain***" |  |  | May 15, 2014 | 2014 | Bacterin International, Inc. | The recalling firm reported there was inadvertent omission of quarterly (Q4 - 2013 and Q1 - 2014) dose audits/verifications on the affected products. All of the affected products were appropriately sterilized | Use error | Bacterin sent e-mail and FedEx letter with return receipt request to it's single consignee on May 15, 2014. Consignee is asked to separate the recalled product in a secure location for return to Bacterin International, Inc. Consignee is asked to complete tracking/verification form included in the letter/e-mail and call 1-888-886-9354 from 8:00AM to 5:00PM (Mountain Time) or send an e-mail to CS@Bacterin.com to obtain a Return Authorization (RA) number. |  | 390 | 390 | Distributed to one foreign consignee in Lebanon. | Terminated | June 02, 2014 | 18 | 390 | Undetermined
Z-1702-2014 | 67453 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | LifeCare 5000 Pump (Plum 1.6) infusion pump, indicated for volumetric infusion of critical drugs and fluids by intravenous (I.V.) administration. |  |  | January 24, 2014 | 2014 | Hospira Inc. | There is the potential for the door roller assembly on the Plum LifeCare 5000 Series and Plum XL families of infusers to break. In the event of a broken door roller pin, the door cannot appropriately lock the cassette in the right position. Depending on the conditions of a broken door roller pin, a number of events can occur. The door cannot be closed or a cassette/door alarm or a false occlusion | Component design/selection | Hospira sent an URGENT MEDICAL DEVICE RECALL letter dated January 24, 2014, to all affected customers via traceable UPS. The letters included instructions for recipients to: inspect the door roller assemblies; if a door roller or pin appears loose, broken or missing, remove the device from clinical service until the door assembly is replaced with a corrected door assembly and contact Hospira at 800-441-4100 (Monday - Friday, 8:00 AM - 5:00 PM CST) to report the issue. All devices requiring repair will be serviced and replaced with the existing design door and/or pin. Hospira is in the process of designing a metal door roller pin to replace the plastic on currently available on the Plum XL and once that design is complete, any door roller/pin serviced on a Plum XL will be replaced with the new metal design. Recipients are further instructed to: inform potential users in their organization of the recall notification; complete the attached reply form and return it to the fax number or e-mail address on the form; and if the pumps were further distributed, to notify the accounts that received the devices and ask them to contact Stericycle at 877-597-9581 (Monday - Friday, 8:00 AM - 5:00 PM EST) to receive a reply form. The URGENT MEDICAL DEVICE RECALL letter also reports that Hospira is in the process of retiring the Plum LifeCare 5000 and Plum XL families of infusion devices as noted in the IMPORTANT PLUM INFUSION PUMP DISCONTINUATION AND TRANSITION NOTICE issued May 1, 2013. As of June 30, 2015, Hospira will consider the products within the Plum LifeCare 5000 and Plum XL Infusion Systems families retired and will no longer support them.For questions regarding this recall call 224-212-2000. |  | 132889 | List Number 02507: 132,889 pumps | Worldwide Distribution - USA nationwide including District of Columbia, Puerto Rico, and US Virgin Islands; Australia, Argentina, Bahrain, Belgium, Bosnia and Herzegovina, Brazil, Bulgaria, Canada, Chile, Colombia, Croatia, Czech Republic, Egypt, Finland, France, Germany, Greece, Hong Kong, Hungary, Israel, Italy, Jordan, Kuwait, Luxembourg, Mexico, Oman, Peru, Poland, Portugal, Qatar, Russia, San Marino, Saudi Arabia, South Africa, Spain, Sweden, Switzerland, Syria, Taiwan, Turkey, United Arab Emirates, United Kingdom, and Yemen. | Open, Classified |  |  | 337666 | Undetermined
Z-1708-2014 | 68153 | No | NEW | Absorbable polydioxanone surgical suture. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | MonoswiftÂż PGCL [poly (glycolide-co-caprolactone)] Absorbable Surgical Suture; USP size 3/0 PGCL; part number L936. Product is labeled sterile. The product is sealed inside an inner foil pouch with an outside TyvekÂż pouch. It is an undyed, monofilament synthetic absorbable suture with needle, size 3/0 (2.0 metric), PS-1 Needle |  |  | April 23, 2014 | 2014 | C P Medical Inc. | CP Medical, Inc., is recalling one lot of Monoswift PGCL [poly (glycolide-co-caprolactone)] Absorbable Sutures for failing to meet the USP minimum average value for knot pull (tensile) strength. | Employee error | CP Medical, Inc., sent a Recall Notice letter dated April 28, 2014 to affected customers. The letter identified the affect product, problem and actions to be taken. The letter instructed customers to immediately discontinue selling or using the recalled product and return all existing stock to CP Medical. They also request their customers to contact downstream customers and request that all their existing stock be returned. CP Medical, Inc. asks all customers to provide them with information as to the quantity of recalled product they are able to return and for all stock they are not able to recover, customers are asked to provide quantity information and location of the product for accountability purposes. The firm requests that all customers contact their customer service representative at 1-800-950-2763 for a Return Goods Authorization (RGA) # and shipping account number. All shipping costs, including those from customers, will be paid for by CP Medical. Replacement product (at no cost) will be provided. |  | 528 | US distribution:4 boxes/ 12 sutures per box (48 total); Outside US: 44 boxes/12 sutures per box. (528 total). | Worldwide Distribution - US Nationwide | Terminated | September 12, 2014 | 142 | 528 | Undetermined
Z-1700-2014 | 68298 | No | MUE | Full-field digital mammography system. | Radiology | Radiology | 510(k) | Hologic Selenia Full Field Digital Mammography System: SEL-00002, SEL-00013, SEL-00030, SEL-00039. |  |  | May 13, 2014 | 2014 | Hologic, Inc. | Selenia Full Field Digital Mammography System defective relay switch, if there is a need to use the emergency off switch, the system may not turn off which could result in patient injury | Component change control | Hologic notified customers by letter dated 5/13/14 via certified mail or Federal Express to inform customers of the problem and provide acknowledgement of the notification and confirmation that their systems have been correctedA Dealer notification dated 5/13/14 sent by certified mail or Federal Express to inform Dealers of the need to notify their customers of the problem and to correct the identified Selenia system using Hologic Technical Bulletin TB-00806. On 5/16/14 the firm expanded the recall to an additional units Mammography System M-IV (2 units) and Selenia Mammography Systems (4) by letter dated 5/14/14.If there are any questions regarding this field correction, please call Hologic Customer Support at 877 371-4372. |  | 11 | 11 units | Distributed in Puerto Rico, and the states of CA, NY, and TX, and the country of Republic of Korea. | Terminated | August 18, 2014 | 97 | 14 | Undetermined
Z-1718-2014 | 68358 | No | NAY | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | Patient Side Cart assembly, 4-arm, for use with Intuitive Surgical, Inc. da Vinci Surgical System IS 4000. The patient-side cart is positioned at the operating room table and contains four arms that are positioned over the target patient anatomy. The endoscope attaches onto any arm. The patient-side assistant attaches/detaches the endoscope and instrument intra-operatively. |  |  | May 09, 2014 | 2014 | Intuitive Surgical, Inc. | Torque wrenches used to tighten certain PSC bolts were out of calibration. This could lead to over- or under-tightening of bolts that may result in system faults, mechanical sound (e.g. scraping, squeaking, rattling), cracks or vibration in the system. | Nonconforming Material/Component | Intuitive sent an Urgent Medical Device Correction letter dated May 13, 2014 to affected customers. The letter identified the affected product, problem and actions to be taken. All personnel who use the device should be notified of the issue and the acknowledgement form completed and returned.Intuitive reps will visit each affect sites to install a replacement PSC and ensure that it is properly integrated with the Da Vinci.For questions contact Intuitive Surgical at 800-876-1310 Option 3 (Monday - Friday 6 am to 5 pm PST). |  | 2 | 2 | Medical facilities in TX and FL | Terminated | June 04, 2014 | 26 | 2 | Undetermined
Z-1719-2014 | 68193 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | MACROLYTE Premie Dispersive Electrode, REF 440-2400, CONMED CORPORATION. The device is applied to the patient during electrosurgical procedures to provide a path for RF current to leave the patient and return to the generator. |  |  | May 08, 2014 | 2014 | ConMed Corporation | ConMed Corporation has received complaints regarding sparking, no output and burning at the cord set connection. The complaints have been confirmed by ConMed. | Nonconforming Material/Component | ConMed sent an Urgent Medical Device Recall letter dated May 8, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Corrective Action: All returned devices will be received into the facility and the lot number and quantity recorded on to an electronic spreadsheet. The returned product will then be stored in a secure quarantined area. At this time, management has not determined as to how the devices will be dispositioned. A protocol(s) will be written to address the destruction of the returned devices. This will be submitted to NYK-DO for approval prior to any action being taken with the returned devices. CAPA 153136 was opened. Fixtures were revised to allow for placement of the foil/gel substrate from the bottom of the foam. This is to maintain consistent dimensions on the product and not allow for float/movement of the substrate from being offset.For further questions please call (315) 624-3024 or fax to (315) 624-3030. |  | 56880 | Domestic - 12,600 units, Foreign - 44,280 units | Worldwide Distribution - USA (nationwide) and the countries of Australia, Belgium, Brazil, Canada, Ecuador, El Salvador, Finland, France, Great Britain, Honduras, India, Indonesia, Israel, Italy, Japan, Latvia, Lebanon, Malaysia, Mexico, Panama, Philippines, Portugal, Saudi Arabia, South Africa, Spain, Thailand, Trinidad, United Arab Emirates and Uruguay. | Terminated | April 11, 2017 | 1069 | 56880 | Undetermined
Z-1615-2014 | 68144 | No | KWI | Elbow joint radial (hemi-elbow) polymer prosthesis. | Orthopedic | Orthopedic | 510(k) | ExploR 7x26mm Modular Radial Stem, Catalog number 11-210062. Elbow prosthesis orthopedic implant. |  |  | April 22, 2014 | 2014 | Biomet, Inc. | This lot of the ExploR Modular Radial Stem may be missing the wedge ramp threads. If the wedge ramp threads are missing from the stem, the likelihood of the set screw loosening from the radial stem intraoperatively is increased. If the set screw loosens, a revision surgery may be necessary. | Process control | Biomet sent an URGENT MEDICAL DEVICE SAFETY ALERT dated April 22, 2014 to affected customers. The letter identified the affected product, problem and actions to be taken. Questions related to this notice should be directed to (574) 372-1570, Monday through Friday, 8 a.m. to 5 p.m. |  | 8 | 8 devices | US Nationwide Distribution in the states of NY, MD, NC, SC, MI, OR, and NV. | Terminated | January 13, 2015 | 266 | 8 | Undetermined
Z-1713-2014 | 68073 | No | DYB | Catheter introducer. | Cardiovascular | Cardiovascular | 510(k) | ArrowÂż Percutaneous Sheath Introducer Kits |  |  | April 14, 2014 | 2014 | Arrow International Inc | Arrow is recalling the ArrowÂż Percutaneous Sheath Introducer Kits due to dilator tip damage noted during internal inspection. Dilator tip damage has the potential to result in vesseldamage. | Process control | Teleflex/Arrow International sent an Urgent Medical Device Recall Notification letter dated April 17, 2014, to distributors and end users. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to immediately discontinue use and quarantine any affected product and to complete the enclosed Recall Acknowledgement Form and fax it to 1-800-343-2935 to return product. Customers would be contacted by a customer service representative for a Return Goods Authorization (RGA) Numbers and would be provided instructions for the return of the product. Customers should complete the form whether they have affected product or not. Customers with questions were instructed to call 1-800-343-2935.For questions regarding this recall call 610-378-0131. |  | 15424 | 15424 | Worldwide Distribution - USA (nationwide) and Canada. | Open, Classified |  |  | 15424 | Undetermined
Z-1722-2014 | 68245 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Philips Medical Systems Computed Tomography X-ray systems (Brilliance CT series, iCT, iCT SP, Ingenuity CT, Ingenuity Core, Ingenuity Core128, Ingenuity Flex, and the MX8000 Dual v. EXP scanners) |  |  | April 01, 2014 | 2014 | Philips Medical Systems (Cleveland) Inc | A customer reported that the table top had become free floating. A Field Service Engineer (FSE) confirmed that the service latch had become disengaged. | Component design/selection | Philips Healthcare sent an Urgent Medical Device Correction letter dated April 1, 2014, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed If they experience a horizontal, free-floating couch motion, to contact their Field Service Engineer immediately. Customers were instructed to retain a copy of this Field Safety Notice with the equipment Instructions forUSER Use (IFU). Customers with questions or needing further information were instructed to contact their local Philips representataive or local Philips Healthcare office. For North America, contact the Customer Care Solutions Center at 1-800-722-9377, Option 5.For questions regarding this recall call 440-483-3032. |  | 6839 | 6, 839 units | Worldwide Distribution: US nationwide including PR and DC, Afghanistan, Albania, Algeria, Argentina, Australia, Austria, Azerbaijan, Bahrain, Bangladesh, Belarus, Belgium, Bhutan, Bolivia, Bosnia, Brazil, Brunei Darussalam, Bulgaria, Canada, Cambodia, Cape Verde, Chile, China, Colombia, Croatia, Cyprus, Czech Republic, Denmark, Dubai, Dominican Republic, Ecuador, Egypt, Ethiopia, Finland, France, French Polynesia, Gabon, Georgia, Germany, Ghana, Greece, Greenland, Guadeloupe, Herzegovina, Hong Kong, Hungary, Iceland, India, Indonesia, Iran, Iraq, Ireland, Israel, Italy, Jamaica, Japan, Jordan, Kazakhstan, Kenya, Kuwait, Kyrgyzstan, LAO People's Dem, Rep., Latvia, Lebanon, Libya, Libyan Arab Jamahiriya, Lithuania, Luxembourg, Malawi, Malaysia, Malta, Martinique, Mauritius, Mexico, Monaco, Mongolia, Morocco, Mozambique, Myanmar, Nahar, Nepal, Netherlands, New Caledonia, New Zealand, Nigeria, Norway, Oman, Pakistan, Palestine, Panama, Peru, Philippines, Poland, Portugal, Qatar, Reunion, Romania, Russia, Saudi Arabia, Senegal, Serbia, Singapore, Slovakia, Slovenia, South Africa, South Korea, Spain, Sri Lanka, Sudan, Sweden, Switzerland, Syrian Arab Republic, Taiwan, Tanzania, Thailand, Tunisia, Turkey, Turkmenistan, Ukraine, United Arab Emirates, United Kingdom, Uzbekistan, Venezuela, Viet Nam &amp; Yemen. | Open, Classified |  |  | 6839 | Undetermined
Z-1723-2014 | 68295 | No | LKK |  |  | General Hospital | PMA | Medtronic Model 8780 Ascenda Intrathecal Catheter. |  |  | May 02, 2014 | 2014 | Medtronic Neuromodulation | Medtronic has determined that an Ascenda catheter kit may have had endotoxin in excess of USP limits. | Release of Material/Component prior to receiving test results | There was only one consignee affected by the recall. The consignee was contacted by phone on 5/2/14 and the information was provided verbally. General Manager &amp; Vice President, Pain Stimulation/Targeted Drug Delivery spoke via phone with the physician managing the patient. The following information was provided:"Medtronic identified the potential for endotoxin levels exceeding USP requirements on product recently implanted (Model 8780, serial number N457062003)."There may be the potential for serious patient injury due to endotoxin within the first hours after implant."Medtronic recommends observing the patient for symptoms of meningitis. Also to be aware that CSF would not test with a positive culture because endotoxin induced meningitis results in an aseptic meningitis.The physician expressed confidence in the management of the patient and stated he was calling the nursing care facility to increase monitoring of the patient. Subsequent information was received regarding the patient indicating there were no signs of an endotoxin reaction. |  | 1 | 1 | Distributed in the state of CT. | Terminated | October 07, 2014 | 158 | 1 | Undetermined
Z-1721-2014 | 67967 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Siemens SOMATOM Emotion 16, Emotion 6, and SOMATOM Spirit Computed Tomography Systems |  |  | March 21, 2014 | 2014 | Siemens Medical Solutions USA, Inc | The rubber damper may degrade and break, causing loosening of the motor mounting plates, which could lead to a series of events that could place the operator and patient at risk of injury. | Component design/selection | Siemens sent a Customer Safety Advisory Notice/ Field Safety Notice letter dated March 21, 2014, to affected customers. The letter identified the affected product, problem and actions to be taken. The letter informed customers that the system modification would be implemented by Siemens service specialists in April 2014 to correct the problem. |  | 345 | 345 | US Nationwide Distribution in the states of IA, TX, FL, NE, NY, OH, NC, KS, IL, CA, OK including PR. | Terminated | March 16, 2015 | 360 | 345 | Undetermined
Z-1717-2014 | 68312 | No | KPR | Stationary x-ray system. | Radiology | Radiology | 510(k) | Siemens Ysio Systems |  |  | May 09, 2014 | 2014 | Siemens Medical Solutions USA, Inc | Excessive usage of cleaning liquid on Siemens Ysio Systems may cause infiltration of the generator console, which may lead to a short circuit causing an unintended release of radiation. | Use error | Siemens sent an Safety Advisory Notice letter dated May 9, 2014, to affected customer. The letter identified the affected product, problem and actions to be taken. For question call +49 (913) 84-9908. |  | 433 | 433 | US Nationwide | Terminated | August 18, 2015 | 466 | 433 | Undetermined
Z-1716-2014 | 68261 | No | HAW | Stereotaxic instrument. | Neurology | Neurology | 510(k) | The NAV3i Platform Power Box supplies the electrical components with power (e.g. Stryker Nav3 Graphite Camera, SPC-3 Computer, IO-Tablet (touch screen) and the monitor). |  |  | May 13, 2014 | 2014 | Stryker Instruments Div. of Stryker Corporation | The Stryker NAV3i Platform Power Boxes produced between December 9, 2013 and March 7, 2014 may have been assembled without electrical insulation of the soldering joints | Process change control | On May 13, 2014, Stryker Instruments sent URGENT MEDICAL DEVICE RECALL notification to their consignees. Notification included the description of the affected product, reason for the recall, risk to health, and actions to be taken by the Customer/User. Once Stryker receives the response from the customer, a Stryker Representative will contact the customer to set up a time to perform the upgrade to the affected power box. Old Power Boxes will be returned to Stryker and placed in an industrial compactor/recycler for destruction. Questions regarding the recall are directed to Kelly Jo Whipple M-F 8am-5pm @ 269-389-2921 or kellyjo.whipple@stryker.com |  | 20 | 20 | Distributed in the states of AR, CA, IN, LA, MA, MD, MI, MT, NV, and TX. | Terminated | September 15, 2014 | 125 | 20 | Undetermined
Z-1730-2014 | 68142 | No | FSY | Surgical lamp. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Maquet PowerLED 700. |  |  | April 16, 2014 | 2014 | Maquet Medical Systems USA | Maquet has identified an issue with the Maquet PowerLED 700 (certain serial numbers). One or several plastic fixation tabs can break, which can cause the ambient light module to partially detach. | Other | Maquet issued "Urgent-Medical Device Field Correction" notices, dated April 16, 2014 to affected customers via Federal Express. The notice identified the product, related issue and the action needed to be taken by the customer. The corrective action will include inspection of the fixation tabs of the Ambient Light Module adn replacement of the Ambient Light Module if tabs are found to be defective. The correction will be performed by an authorized MAQUET Service Technician. MAQUET Service Team representatives will schedule on-site service with customers. If customers have any questions or require additional information, the notice instructs them to contact their local MAQUET representative, or MAQUETCustomer Service at 1-888-627-8383 (option 3), Monday through Friday, between the hours of 8:00 a.m. and 6:00 p.m. EST. |  | 1169 | United States 1169 units | US Nationwide Distribution. | Terminated | April 06, 2016 | 721 | 1169 | Undetermined
Z-1727-2014 | 68230 | No | GHH | Fibrinogen/fibrin degradation products assay. | Hematology | Hematology | 510(k) | Alere TriageÂż D-Dimer Test |  |  | May 02, 2014 | 2014 | Alere San Diego, Inc. | Alere initiated this recall because a limited number of Alere TriageÂż D-Dimer devices from PN 98100 Lot W53884B were incorrectly manufactured. As a result, when testing patient sample these affected devices will either generate an error code (e.g. E4 error), or they may generate an incorrect result. | Under Investigation by firm | Alere sent an Urgent Medical Device Recall letter dated May 2, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Customers are instructed to stop using all product from the affected lot, lot W53884B, and to discard any unused product. The customers were also instructed to complete and return the verification form indicatingany credit or replacement product requested. They were also instructed to contact their primary consignees and advise them of the notice.Should you have any questions about the information contained in this notification, please call (877) 308-8287. and /or email:Triage.Support@alere.com |  | 373 | 373 kits | Nationwide Distribution only. | Terminated | February 27, 2015 | 301 | 373 | Undetermined
Z-1726-2014 | 66274 | No | OUE | Instrumentation for clinical multiplex test systems. | Clinical Chemistry | Immunology | 510(K) Exempt | Leica Ceiling Mount MS-1F &amp; MC-1F for use with the M500 N operating microscope. Article Number 10446657 utilizes a supply voltage of 110 V. Article Number 10446658 utilizes a supply voltage of 220 V. |  |  | September 11, 2013 | 2013 | Leica Microsystems, Inc. | The ceiling mount supplier encountered an event in which the drive screw in the motor unit, which lifts and lowers the lower lift arm, broke. As a result, the lift arm fell down approx. 60 cm (23.6 inches). | Nonconforming Material/Component | A Medical Device Correction letter, dated September 11, 2013, was sent to all direct accounts. The letters advised customers that a Leica Microsystems representative will contact them and make arrangements to conduct an inspection of the ceiling mount(s) and to provide additional preventative maintenance and inspection protocols for the devices. Customers are asked to complete and return the enclosed Acknowledgement Form and send it back to Leica Microsystems within 10 days of receiving the Medical Device Correction letter. Customers with questions should contact their Leica Microsystems representative. |  | 103 | 103 total | Worldwide Distribution -- USA, including the states of AZ, IL, MA, MO, NC, OR, TX, VA, and WI; and, the countries of Australia, Austria, Chile, China, Colombia, Czech Republic, Greece, Israel, Italy, Japan, Malaysia, Netherlands, Norway, Oman, Saudi Arabia, South Korea, Spain, Sweden, Switzerland, and United Kingdom. | Terminated | April 24, 2017 | 1321 | 103 | Undetermined
Z-1674-2014 | 68150 | Yes | MUQ |  |  | Cardiovascular | PMA | BG3515-5 US 5mL BioGlue Surgical Adhesive. |  |  | April 25, 2014 | 2014 | CryoLife, Inc. | Serum albumin component monomer failed to meet internally established end of shelf-life specification. | Process control | All consignees were notified of the voluntary actions by issuance of Field Correction/Removal Notice on April 24, 2014. This notices states which lot numbers are to be returned, and which lot numbers can be relabeled with a corrected expiration date. |  | 6190 | 6190 units | Distributed Nationwide including Puerto Rico and the states of AL, AK,AZ, AR, CA, CO, CT, DE, FL, GA, HI, ID, IL, IN, IA, KS, KY, LA, MD, MA, MI, MN, MS, MO, MT, NV, NH, NJ, NM , NY, NC, ND, OH, OK, OR, PA, SC, SD, TN, TX, UT, VA, WA, WV, and WI. | Terminated | November 10, 2014 | 199 | 6190 | Undetermined
Z-1669-2014 | 68190 | No | GFG | Surgical instrument motors and accessories/attachments. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Synthes ProDisc-C 2.0mm Milling Bit, Synthes 2.0mm Milling Bit HXC, Synthes 2.0mm Milling Bit with 2.35mm x 33.5mm Shaft, and Synthes 2.0mm Milling Bit with Flat Coupling. Spinal surgical instrument: The Synthes ProDisc-C 2.0mm Milling Bits are used as a part of the milling option for keel preparation of the vertebral bodies. The milling bits are coupled to a power tool and interface with the milling guide, which is connected to a trial implant for the keel preparation step. |  |  | April 29, 2014 | 2014 | Synthes, Inc. | Certain lots of the ProDisc-C 2.0mm Milling Bits have an increased potential to break during use. | Process control | DePuy Synthes sent a Recall letter, dated April 29, 2014, to affected customers. The letter identified the affected customers, problem and actions to be taken. The letter instructed customers to discontinue use and distribution of the device and to return affected devices to Synthes. For questions call 610-719-5450. |  | 6690 | 6690 | US Nationwide Distribution | Open, Classified |  |  | 6690 | Undetermined
Z-1671-2014 | 67527 | No | JIT | Calibrator. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | ARCHITECT Intact PTH Calibrators consists of 6 bottles (4.0 mL each). Calibrator A contains Bis Tris Propane buffer with protein (bovine) stabilizer. Calibrators B-F contain PTH (synthetic peptide) in Bis Tris Propane buffer with protein (bovine) stabilizer. Preservatives: sodium azide and ProClin 300. |  |  | February 12, 2014 | 2014 | Abbott Laboratories | Abbott has confirmed that a performance shift in the ARCHITECT Intact PTH assay has the potential to generate falsely elevated results on patient samples. Results generated with impacted lots may demonstrate a positive shift relative to those generated with previous calibrator and/or control lots. The issue may also impact your established ARCHITECT Intact PTH reference ranges. Abbott ARCHITECT In | Under Investigation by firm | Product Information letters dated August 23, 2013 were sent to all direct accounts alerting the direct accounts to a positive bias observed in test results using some of the affected kits. The letter advised the direct accounts to take this information into account when interpreting patient results. Customers with questions were advised to contact Customer Service at 877-4ABBOTT (877-422-2688) (24 hours a day, 7 days a week) and to refer to Product Information Letter PI23AUG2013. Product Recall letters dated February 12, 2014 were sent to all direct accounts. The letter included instructions to: 1) immediately discontinue the use of, and destroy, any remaining inventory of ARCHITECT Intact PTH assay according to each customer's laboratory practices; 2) to continue Intact PTH testing at your laboratory, Abbott recommends that customers identify an alternate method for testing patient samples; 3) if further guidance is needed, please contact the customers's Abbott sales representative; 4) review the letter with the treating physician or healthcare provider to ensure they are aware of the issue and can determine whether re-evaluation of the patient is necessary, and that Abbott included a physician letter that you can use to notify your ordering physicians of this issue; 5) If the customer forwarded any product to other laboratories, please inform them of this Product Recall and provide a copy of the recall letter, customer reply, and physician letter to them; and, 6) complete and return the Customer Reply. Customers with questions were advised to contact Customer Service at 877-4ABBOTT (877-422-2688) (24 hours a day, 7 days a week). |  | 7780 | 7,780 calibrator kits | Worldwide Distribution - US Nationwide including District of Columbia and Puerto Rico; Algeria, Andorra, Angola, Argentina, Armenia, Austria, Azerbaijan, Bahamas, Bangladesh, Barbados, Belarus, Belgium, Bosnia &amp; Herzegovina, Brazil, Brunei, Bulgaria, Cameroon, Canada, Cayman Islands, Chile, China, Colombia, Costa Rica, Croatia, Cyprus, Czech Republic, Denmark, Dominican Republic, Ecuador, Egypt, El Salvador, Ethiopia, Finland, France, Gaza and Jericho, Georgia, Germany, Greece, Guatemala, Honduras, Hong Kong, Hungary, India, Indonesia, Iran, Iraq, Ireland, Israel, ITALY, Jamaica, Japan, Jordan, Kenya, Korea (South Korea), Kuwait, Kyrgyzstan, Latvia, Lebanon, Libya, Lithuania, Luxembourg, Macedonia, Malaysia, Maldives, Mali, Mauritania, Mexico, Montenegro, Morocco, Netherlands, New Zealand, Norway, Oman, Pakistan, Palestine, Panama, PanamÂż, Paraguay, Peru, Philippines, Poland, Portugal, Qatar, Romania, Russia, San Marino, Saudi, Senegal, Serbia, Singapore, Slovakia, Slovenia, South Africa, SPAIN, Sri Lanka, Sweden, Switzerland, Taiwan, Thailand, Trinidad and Tobago, Tunisia, Turkey, United Arab Emirates, United Kingdom, Ukraine, Uruguay, Venezuela, Vietnam, and Yemen | Terminated | May 22, 2015 | 464 | 17243 | Undetermined
Z-1666-2014 | 68040 | No | BSP | Anesthesia conduction needle. | Anesthesiology | Anesthesiology | 510(k) | Sprotte Spinal Tray / Sprotte Spinal Needle 24 gauge with Introducer |  |  | April 03, 2014 | 2014 | B. Braun Medical, Inc. | B.Braun Medical, Inc. (BBMI) has received reports of 24 gauge Sprotte Needles that are missing the accompanying introducer needles. | Nonconforming Material/Component | B/Braun sent an Urgent Medical Device Recall Notification letter dated April 3, 2014, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to review the Device Recall Notification in its entirety and ensure that all users in their organization and other concerned persons are informed about this voluntary product recall. The letter should also be forwarded to their customers. Customers were instructed to determine their current inventory of the affected lots within their facility, DO NOT DESTROY ANY AFFECTED PRODUCT. Further use of the affected product should be discontinued immediately and quarantined. Customers should use the attached Product Removal Acknowledgement form to record the total number of inidividual units. If cusotmers have no remaining inventory they should check the box indicating zero inventory. Return the completed Product Removal Acknowledgement form to BBMI Quality Assurance department by fax to 610-849-1197 or email to PA_QualityAssurance.BBMUS_Service@bbraun.com within 2 weeks of receipt, even if the inventory is zero (0). If customers have any affected product they should call BBMI Customer Support Department at 800-227-2862 to arrange for return and replacement product. Customers with questions were instructed to call the Clinical and Technical Support Department at 1-800-854-6851.For questions regarding this recall call 610-596-2870. |  | 7400 | 7400 | Worldwide Distribution - USA (nationwide) and to Canada. | Terminated | December 02, 2014 | 243 | 7400 | Undetermined
Z-1607-2014 | 67684 | No | DSQ |  |  | Cardiovascular | PMA | HeartWare HVAD Pump Implant Kit (Heartware Ventricular Assist Device) - HeartWare Battery |  |  | April 16, 2014 | 2014 | HeartWare Inc | Patients and caregivers are instructed that if a battery does not provide two hours of support or behaves unusually, that battery should no longer be used and should be replaced. | Under Investigation by firm | HeartWare sent an Urgent Medical Device Correction letter dated April 30, 2014, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were advised to familiarize themselves with the letter and review their Patient Manual about proper power management. Discuss with their VAD Coordinator any questions they may have about the HeartWare System or the letter.Always keep two (2) power sources connected to your controller; never leave your controller connected to only one (1) power source except while briefly switching to another power source." Review the recommended practices for power management provided later in this letter. If a battery shows any of the abnormal battery behaviors described, stop using that battery and contact your VAD Coordinator immediately to replace it.Customers were instructed to confirm that they have received and understood the information. Sign and return the enclosed Confirmation form as soon as they can to HeartWare in the prepaid self-addressed envelope or by fax to 1-305-364-2665. For questions customers were instructed to contact their doctor or VAD coordinator.For questions regarding this recall call 305-364-1575. |  | 32349 | 32,349 | Worldwide Distribution - USA including AL, AZ, CA, CO, CT, DE, FL, GA, HI, IL, IN, KY, LA, MA, MD, MI, MN, MO, NJ, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, WY, District of Columbia and Internationally to Canada, Dominican Republic, Australia, Belgium, Italy, Germany, New Zeland, South Afica, Japan, Venezuela, Czech Republic, France, Switzerland, Brazil, UK, Saui Arabia, Taiwan, Slovakia, Luxemburg, Hong Kong, Lebano, India, Chile, Oldenburg, Israel, Netherlands, Croatia, Egypt, Turkey, Malaysia, Norway, Greece, Spain, Lithuania, Korea, Singapore, Poland, and Argentina. | Open, Classified |  |  | 32349 | Undetermined
Z-1620-2014 | 68030 | No | FMF | Piston syringe. | General Hospital | General Hospital | 510(k) | RAPIDFILL Syringe Strip, For use with RAPIDFILL Automated Syringe Filler (ASF): H93890200 RAPIDFILL Syringe Strip, White. H93890221 RAPIDFILL Syringe Strip, Yellow. H93890222 RAPIDFILL Syringe Strip, Orange. H93890223 RAPIDFILL Syringe Strip, Red. H93890225 RAPIDFILL Syringe Strip, Blue. H93890227 RAPIDFILL Syringe Strip, Salmon. H93890229 RAPIDFILL Syringe Strip, Violet. H93890231 RAPIDFILL Syringe Strip, Gray. H93890232 RAPIDFILL Syringe Strip, Green. Sterile 10 ml syringes with twist-off luer sealing tip caps banded together with label material. |  |  | April 16, 2014 | 2014 | Baxter Corporation Englewood | Possible breach in sterile barrier for inner and outer bags of theproduct. | Process control | Baxter sent an Urgent Product Recall letter dated April 30, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Baxter is requesting that you take the following actions:1. Locate and remove all affected product from your facility (the product code and lot number can be found on the individual product package or shipping carton).2. To return the affected product, receive credit and order replacement, please contact Baxter Healthcare Center for Service at 1-888-229-0001, Monday through Friday, between the hours of 7:00am and 6:00pm Central Time. Please have your Baxter eight-digit ship-to account number ready when calling.3. Complete the enclosed customer reply form, and return it to Baxter by either fax or scanned e-mail.4. If you are a dealer, wholesaler, or distributor/reseller that distributed any product to other facilities, please conduct a recall with your end-user customers in accordance with your customary procedures.If you have questions regarding the content of this communication, please call The Center for One Baxter at 1-800-422-9837 during the hours of8:00am to 5:00 pm Central Time, Monday through Friday. |  | 774600 | 774, 600 | Worldwide Distribution - US Distribution including the states of AL, AZ, CT, CA, FL, GA, OH, OR, IL, KS, MA, MI, MN, MO, NC, PA, HI, TN, VA, NV and WI and the country of Canada. | Terminated | March 23, 2015 | 341 | 774600 | Undetermined
Z-1621-2014 | 67964 | No | OZT |  | Cardiovascular | Cardiovascular | 510(k) | InterValve V8 Balloon Aortic Valvuloplasty Catheter, model 172212C110, 192412C110, 212612C110, 232812C110. Used in Cardiovascular procedures for Balloon Aortic Valvuloplasty. |  |  | March 14, 2014 | 2014 | InterValve Inc | Leak issues were discovered through accelerated age testing. | Nonconforming Material/Component | Since this recall involved only 2 units at one site, the devices were picked up by the sales representative. There was no written communication regarding this.The site was verbally informed that there was a problem with the device that manifests itself during prepping and the devices need to be retrieved for retest. |  | 2 | 2 | Distribution to TN only | Terminated | June 16, 2014 | 94 | 2 | Undetermined
Z-1622-2014 | 68107 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | syngo Imaging XS |  |  | April 03, 2014 | 2014 | Siemens Medical Solutions USA, Inc | A complaint reported that the archiving method at one customer site was changed from implicit to explicit, and the cleanup procedure enabled by the explicit archiving method may have deleted data from archive itself. | Use error | Siemens sent a Customer Safety Advisory Notice dated April 3, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Our service organization already checked all syngo Imaging sites and did not find any other occurrence of the wrongly configured link.Additionally our service organization will include the check for such wrongly configured links into the existing proactive site monitoring procedures.We appreciate your understanding and cooperation with this safety advisory and ask you to immediately instruct your personnel accordingly. Please include this safety advisory notice in your operator's manual. In the interests of safety, we ask that you perform the above preventive measures and inform all affected personnel immediately.If you have sold this device/equipment and it is no longer in your possession, we kindly ask that you forward this safety notice to the new owner of this device/equipment. Please inform us about the new owner of the device/equipment.We regret any inconvenience that this may cause, and we thank you in advance for your understanding. For further questions please call (610) 219-6300.. |  | 64 | 64 | Nationwide Distribution | Terminated | December 15, 2014 | 256 | 64 | Undetermined
Z-1629-2014 | 68021 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Anspach Irrigation Clip 40 |  |  | December 03, 2013 | 2013 | The Anspach Effort, Inc. | The Irrigation Clips were assembled with the incorrect (smaller) clips. | Mixed-up of materials/components | The firm, Anspach, sent an URGENT: Medical Device Removal letter dated December 3, 2013 to its customer. The letter described the product, problem and actions to be taken. Should you have any queries, please do not hesitate to contact Aspach Customer Support at (800) 327-6887. |  | 150 | 150 | US Nationwide Distribution in the state of Mississippi. | Terminated | April 18, 2016 | 867 | 150 | Undetermined
Z-1623-2014 | 68129 | No | LJH | Urological catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | LATEX-FREE; EXTENSION SET; For Bladder Irrigation or Urinary Drainage, Compatible with T-U-R FLOW-POUCH RESERVOIR, list 6542, 71 inch (180 cm) Nominal Length; List No. 4693-01; single use; Hospira |  |  | March 25, 2014 | 2014 | Hospira Inc. | The Tyvek cover may not be completely sealed to the rest of the package allowing openings at the edge of the sterile package for the following products: List No. 04693-01, Latex Free Extension Set; List No. 06543-01, Nonvented Y-Set for Transurethral Resection; List No. 06544-01 Cystoscopy / Irrigation Set; List No. 06599-01, Large Bore Y-Irrigation Set (Latex-Free); and, List No. 15239-01, 4-Lea | Process design | Hospira sent an URGENT MEDICAL DEVICE RECALL letter dated March 25, 2014 to all affected customers. The letter included instructions for the direct accounts to: 1) check their inventories and quarantine any of the affected products; 2) complete and return the attached reply form to the fax number or e-mail address on the form; 3) inform potential users of these products of this recall; 4) return the affected products to Stericycle using the labels provided with the letter (additional labels or reply forms can be obtained by calling Stericycle at 877-377-5128, Monday - Friday, 8:00 AM - 5:00 PM EST); and, 5) if the products were further distributed, notify those customers and request that they contact Stericycle at 877-377-5128 (Monday - Friday, 8:00 AM - 5:00 PM EST). An updated URGENT MEDICAL DEVICE RECALL letter dated March 28, 2014 was sent to all affected customers. The letter contained instructions similar to those found in the March 25, 2014 letter. |  | 42292 | 42,292 sets (US only) | Worldwide Distribution - US: Nationwide in the states of (AK, AL, AR, AZ, CA, CO, CT, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, UT, WA, WI, WV, WY) including District of Columbia and Puerto Rico; *** the countries of Barbados, Canada, Costa Rica, Dominican Republic, Hong Kong, Kuwait, Libya, Saudi Arabia, United Arab Emirates | Open, Classified |  |  | 2962342 | Undetermined
Z-1628-2014 | 67923 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Lithium Ion Battery M4605A and M4607A for use with IntelliVue Patient Monitors. The monitor is used for monitoring and recording of, and to generate alarms for, multiple physiological parameters. |  |  | April 04, 2014 | 2014 | Philips Medical Systems, Inc. | The risk of battery failure increases with age, when a battery remains in use longer than 3 years after date of manufacture or 500 charge-discharge cycles. Such failure can result in overheating that in rare cases can cause the battery to ignite or explode. | Unknown/Undetermined by firm | Philips sent an Urgent: Medical Device Correction letter dated April 4, 2014 to all affected customers. The letter identified the affected product, problem, action to be taken and instructions for Use Addendum. Customers were instructed to review this information with all staff members who are involved in the operation and battery management of the IntelliVue patient Monitors and need to be aware of the contents of this communication. The addendum should be stored with the Monitor Instructions for Use. Please contact your local Philips representative or call us at 1-800-722-9377 with questions or concerns about this correction. |  | 188520 | 188,520 devices | Worldwide Distribution - Including US Nationwide, Canada, Czech Republic, Denmark, Netherlands, Estonia, Finland, France, Germany, Greece, Hungary, Italy, Japan, Kazakhstan, Norway, Poland, Romania, Russia, China, Slovakia, Sweden, and Turkey. | Open, Classified |  |  | 188520 | Undetermined
Z-1631-2014 | 67924 | No | OSI | AC-powered adjustable hospital bed. | General Hospital | General Hospital | 510(K) Exempt | ECATS E 850 Bariatric Bed, True Air Technologies, Inc. |  |  | March 18, 2014 | 2014 | SCM True Air Technologies LLC | The firm is recalling the products due to lack of design controls, lack of a Device Master Record, and failure to approve specific vendors. | Nonconforming Material/Component | SCM True Air Technologies LLC sent an Urgent Medical Device Recall letter to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to remove the affected product from use and return to True Air Technologies immediately. The units should be decontaminated and bagged with a certification of disinfection before return to True Air Technologies. True Air will provide replacement units to replace the returned units. Customers were instructed to refer to the Medical Device Recall Return Response Form attached for acknowledgement of receipt and contact information for this voluntary recall. Customers with questions should call 800-682-7163. |  | 48 | 48 beds | Nationwide Distribution including GA, KY, and MI. | Completed |  |  | 155 | Undetermined
Z-1634-2014 | 68240 | No | DRY | Cardiopulmonary bypass on-line blood gas monitor. | Cardiovascular | Cardiovascular | 510(k) | CDI Blood Parameter Monitoring System; Catalog Numbers or Part Numbers - 145891, 500AVHCT; Serial Numbers - BPM SN B015552 (installed into CDI 500 Monitor SN 3535) |  |  | May 09, 2014 | 2014 | Terumo Cardiovascular Systems Corporation | There is a potential for inaccuracies in blood gas readings/measurements. There is a potential risk of respiratory alkalosis or cerebral vasoconstriction with low paCO2. If the Hemoglobin and Hematocrit measurements are not accurate, there is a potential risk of unnecesary blood transfusion with resulting risk of infection or inflammatory reaction. BPM shipped without meeting manufacturing accep | Release of Material/Component prior to receiving test results | On 5/9/14, TCVS placed a phone call to the one affected customer regarding the recall. The phone call explained the reason for the recall, identification of affected product, potential hazard, and instructions for return of affected product. |  | 1 | 1 | one customer in IA | Terminated | July 11, 2014 | 63 | 1 | Undetermined
Z-1635-2014 | 68012 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Microline Renew Fenestrated Control Tip Grasper, Single Patient Use, Product Number: 3222 |  |  | April 14, 2014 | 2014 | Microline Surgical | Grasper jaw may break when force is applied to the jaw | Nonconforming Material/Component | Microline Surgical sent an Recall Notice dated April 14, 2014, via overnight mail to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Immediately examine your inventory and quarantine product subject to this recall. In addition, if you have further distributed this product, please identify your customers and notify them at once of this product recall. Your notification to your customers may be enhanced by including a copy of this recall notification letter.Please contact Microline Customer Service ( 978) 922-9810 for a RGA number to return all unused product. This recall should be carried out to the hospital level. Please complete and return the enclosed response form as soon as possible. |  | 540 | 540 units | Worldwide Distribution - US including the states of TX, IL, TN and NY, and the country of Chile. | Terminated | September 02, 2015 | 506 | 540 | Undetermined
Z-1636-2014 | 68043 | No | JDS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | TRIGEN INTERTAN NAIL: |  |  | April 04, 2014 | 2014 | Smith &amp; Nephew Inc | A manufacturing error resulted in the set screw being inserted too deep in the affected nails. | Process control | The firm initiated the recall via an "Urgent - Product Recall 1st Notification - Urgent" letter and e-mail on 04/04/2014, identifying the product, explaining the problem, and the potential risk associated with using the product. Customers are to inspect their inventory, and locate and quarantine any affected product. Additionally, the Inventory Return Certification form should be completed and returned along with any affected product identified, once a Return Authorization number has been obtained from Smith &amp; Nephew's Safety Affairs Department. |  | 1253 | 1,253 units total | Worldwide Distribution -- USA, France, Germany, Sweden, Switzerland, UK, Canada, Australia, Japan, China, India, Dubai, South Africa, South Korea, Mexico, and Venezuela. | Terminated | July 01, 2017 | 1184 | 1253 | Undetermined
Z-1642-2014 | 67958 | No | MUE | Full-field digital mammography system. | Radiology | Radiology | 510(k) | MXR-35 Diagnostic X-ray High Voltage Generator manufactured by DRGEM Corporation. |  |  | June 29, 2012 | 2012 | Fujifilm Medical System USA, Inc. | The MXR-35 X-ray Generator's labeling did not include the certification statement as required by 21 CFR Part 10 10.2. | Radiation Control for Health and Safety Act | Fujifilm will replace the violative labels with replacement labels provided by DRGEM, to bring the X-ray generator into compliance. You stated that you expect for all nine installed units to be corrected by the end of May 2014. The seven units in the Stamford, CT warehouse will be corrected when the generators are delivered to customer sites. Under 21 CFR 1003.31(c)(2), this exemption is granted subject to the following conditions:1. You must correct the missing certification statement for all units shipped into the US within the timeline you specified, by the end of May 2014.2. You must ensure that all future units that will be delivered and installed in the US will be in compliance with the labeling requirement, per 21 CFR 1010.2. |  | 14 | 14 | Worldwide Distribution - US Distribution including the states of TX and KY and Internationally to China. | Open, Classified |  |  | 14 | Undetermined
Z-1649-2014 | 67858 | bandage, elastic | FQM | Elastic bandage. | General Hospital | General Hospital | 510(K) Exempt | Esmark Elastic Bandage (Sterile), 4" x 9'. Item DYNJ05116A, Individually wrapped, 20 bandages per case. Packaged by Medline Industries Inc. This product is used as an elastic bandage to support and compress a part of a patient's body. It is also used as a tourniquet to restrict blood flow to a part of a patient's body. |  |  | March 20, 2014 | 2014 | Medline Industries Inc | This lot did not go through the correct sterilization procedures. This product may potentially be non-sterile. | Employee error | Medline Industries issued an Immediate Action Required letter dated March 20, 2014 to all affected customers. The letters included instructions to: 1) immediately check inventories for the recalled products and quarantine the recalled products; 2) complete and return the enclosed destruction form listing the quantity of destroyed product (credit will only be issued if the completed form is received); and 3) If the customer is a distributor, promptly notify the distributor's customers that may have received the recalled products about this recall. Direct accounts and their customers that have any questions can contact Medline Industries at 866-359-1704. |  | 2860 | 2,860 bandages | Nationwide Distribution including AZ, CA, IA, IL, IN, KY, LA, MI, MS, NJ, NY, OH, PA, TX, and WA. | Open, Classified |  |  | 2860 | Undetermined
Z-1643-2014 | 67996 | Yes | KDI | High permeability hemodialysis system. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Polyflux Revaclear, Capillary Dialyzer, Model Number 110633; |  |  | April 16, 2014 | 2014 | Gambro Renal Products, Incorporated | The firm received several complaints for internal blood leaks involving various lot numbers produced in late 2013. | Storage | The recall was initiated by an Urgent: Medical Device Recall letter delivered by UPS Overnight Delivery on April 16, 2014. The letter identified the affected product and reason for the recall. The firm is asking customers to remove and quarantine any remaining affected product from their inventory for collection and replacement. Customers are to also complete and return the Customer Reply Form to the Gambro Regulatory Affairs Department. Customers with affected product are to contact Gambro Customer Support or their Gambro representative to return the product. Questions should be directed to Customer Support at 1-800-651-2623. |  | 2461824 | 2,461,824 units total | Worldwide Distribution -- USA, including the territory of Puerto Rico, and the countries of Bermuda and Australia. | Terminated | December 08, 2015 | 601 | 2461824 | Undetermined
Z-1648-2014 | 67643 | No | KLI | Total thyroxine test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | VITROS TT4 Calibrator |  |  | February 20, 2014 | 2014 | Ortho-Clinical Diagnostics | Total T4 Calibrators and Reagent Packs may have calibration failures or low Quality Control results. | Other | Ortho Clinical Diagnostics sent an Urgent Product Correction Notifications dated February 19, 2014, to all affected customers via Fed Ex and/or US Postal Service Priority Mail. The letter identified the product the problem and the action needed to be taken by the customer.Customers are advised to immediately discontinue using and discard all remaining affected product. Customers should review their Quality Control results that were generated using the affected lot. Customers should consult their Laboratory Medical Director if they suspect that patient results were affected. Complete the confirmation of receipt form and return by March 4, 2014. Forward a copy of the Product Correction Notice if the product was distributed outside of your facility. Fax to :1.888.557.3759 or 1.585.453.4110 |  | 583 | 583 units Total (357 units domestically &amp; 226 units internationally) | Worldwide Distribution - USA (nationwide) and the countries of Australia, Brazil, Canada, China, Colombia, India, Japan, Mexico, Republic of Panama, Singapore, Venezuela, England, Germany, Italy and Spain. | Terminated | April 19, 2017 | 1154 | 583 | Undetermined
Z-1523-2014 | 67815 | No | DSP | Intra-aortic balloon and control system. | Cardiovascular | Cardiovascular | 510(k) | Datascope System 98/98xt, CS 100/CS100i and CS 300 Intra-Aortic Balloon Pump. The Intra-Aortic Balloon Pump (IABP) is an electromechanical system used to inflate and deflate intra-aortic balloons in cardiovascular procedures. |  |  | March 21, 2014 | 2014 | Maquet Datascope Corp - Cardiac Assist Division | Potential mechanical failure of the fan assembly associated with the power supply. A fan assembly failure could result in the power supply overheating and cause the IABP to shut down without warning. | Nonconforming Material/Component | Maquet Inc. sent an Urgent Medical Device Field Correction letter dated March 21, 2014, to all affected customers. The letter identified the affected product, problem, and actions to be taken.. Customers were instructed that their Service Representatiave would contact them to schedule the replacement of the fan assembly at no cost to the customer. Upon completion customers were requested to sign a service repair order to verify satisfactory completion of the work. Customers with questions within Puerto Rico, were instructed to contact Technical Support in Mexico at +52-55-9000-8970, ext 104.For questions regarding this recall call 973-709-7652.A field correction letter was sent out on April 15, 2014 to additional consignees that did not receive the March 15, 2014 letter.Maquet issued press on May 9, 2014. |  | 11881 | 11,881 units (US 5183 and OUS 6698) | Worldwide Distribution - USA (nationwide) and approximately 104 countries. | Open, Classified |  |  | 11881 | Undetermined
Z-1639-2014 | 67922 | Yes | KPO | Hemodialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Naturalyte Liquid Bicarbonate Concentrate, 6.4 Liter Bottle |  |  | April 10, 2014 | 2014 | Fresenius Medical Care Holdings, Inc. | Naturalyte Liquid Bicarbonate maybe contaminated | Under Investigation by firm | Fresenius Medical Care North America issued an Urgent Medical Device Recall Letter dated April 10, 2014, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to examine their stock immediately and determine whether they have any of the affected parts on hand. If customers have the affected product, they were instructed to discontinue use and place all Naturalyte Liquid Bicarbonate Concentrate in a secure area for return to Fresenius Medical Care Renal Therapies Group, LLC (FMC-RTG). Customers will be instructed to contact their Fresenius Medical Care Technical Services Team for instructions on how to return the recalled product. Firm will also conduct 100% telephone calls. If you have any additional questions, please contact your FMCNA Customer Service Team at 1-800-323-5188.Fresenius Medical Care sent an Urgent Expanded Medical Device Recall letter on May 1, 2014 for 9 additional lots. Customers were instructed to complete and return the attached FAX BACK FORM to confirm that they have received the notice and to indicate whether or not they have the affected product in their possession.For questions regarding this recall call 800-662-1237. Firm issued Press on 5/21/14. |  | 672784 | 672,784 units | Worldwide Distribution - USA (nationwide) and Canada. | Open, Classified |  |  | 672784 | Undetermined
Z-1650-2014 | 68094 | No | DWJ | Thermal regulating system. | Cardiovascular | Cardiovascular | 510(k) | SW-2008 Snuggle Warm Sterile Cardiac Connective Warming Blanket; used with the Equator Convective Warmer, are intended for thermal regulation of a patient's temperture and are designed to prevent hypothermia and/ or reduce |  |  | April 22, 2014 | 2014 | Smiths Medical ASD, Inc. | The hose could separate from the blanket. | Process change control | On April 24, 2014, consignees were notified of the recall via letter. They were instructed to do the following: Inspect inventory and segregate any unused Blankets listed in the notification, complete and return the attached Confirmation Form by fax or e-mail within 7 days of receipt of this notice, and destroy the product. |  | 2120 | 2120 devices | Worldwide Distribution Including the States of CA, IL, KY, MS, NC, TN, TX, WA, WI and the Countries of BH, DE, NO, RU | Terminated | February 27, 2015 | 311 | 2120 | Undetermined
Z-1651-2014 | 68223 | No | LBH | Cavity varnish. | Dental | Dental | 510(k) | Gluma Desensitizer Power Gel |  |  | April 23, 2014 | 2014 | Heraeus Kulzer, LLC. | The firm discovered that the Gluma Desensitizer Power Gel can be inadvertently extruded from the application syringe all at once when it is supposed to extrude on the desired area a small amount at atime. If extruded all at once, product could reach the dental patients mucosa, face or eyes or the dental professional could be exposed. This would cause irritation, burning or blistering. | Nonconforming Material/Component | Heraeus Kulzer sent a Voluntary Recall letter dated April 22, 2014 to all affected doctors and supply centers. The letter identified the affected product, problem and actions to be taken. The letter instructed customers to:1. Inspect your inventory, both in storage and in use, and identify the lot number on your kits and syringes. Refer to attached images.2. If you identify any Gluma Desensitizer PowerGel with the aforementioned kit or syringe lot numbers, immediately cease usage of this product and call Technical Services at 1 (574) 299-54093. Please advise if you forwarded this product to any other end users. We will need their contact information inClusive of name, address and phone number.For questions regarding this notification, please call our Technical Service team at (574) 299-5409. |  | 1278 | 5,112 syringes or 1,278 kits (1 Kit=4 syringes) | US Nationwide Distribution in the states of GA, TX, IL, CA, WA, WI, HI, MO, FL, NJ, TN, PA, NY, NC, OH, IN, CT, KY, AZ, NV, IA, | Terminated | November 05, 2014 | 196 | 1278 | Undetermined
Z-2126-2014 | 68789 | No | DPW | Cardiovascular blood flowmeter. | Cardiovascular | Cardiovascular | 510(k) | FlowCOUPLER. An implantable device that is used to detect blood flow in vessels. Used in the end-to-end anastomosis of veins and arteries normally encountered in microsurgical and vascular reconstructive procedures. The FlowCOUPLER System is intended to detect blood flow and confirm vessel patency intra-operatively and post-operatively at the anastomotic site. Baxter Healthcare. |  |  | July 11, 2014 | 2014 | Baxter Healthcare Corporation | Baxter Healthcare is recalling the FlowCOUPLER Device because there is a possibility that some units may not maintain electrical continuity during handling intra-operatively and post-operatively, resulting in a loss of Doppler signal. This may result in the inability to monitor anastomotic patency, with the possibility that blood flow may not be appropriately monitored. | Process control | Synovis Micro Companies Alliance, Inc. (a subsidiary of Baxter International Inc.) sent an Urgent Medical Device Recall letter on July 11, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. The letter informs the customers that the recall is being conducted as a precautionary measure due to the possibility that some units may not maintain electrical continuity during handling intra-operatively and post-operatively, resulting in a loss of Doppler signal. This may result in the inability to monitor anastomotic patency, with the possibility that blood flow may not be appropriately monitored. The recall letter informs the customers of the safety risk and and actions to be taken. Customers with questions about the recall are instructed to contact the Synovis MCA Customer Service unit at (800) 510-3318 or (205) 941-0111. |  | 1315 | 1,315 units | Worldwide Distribution - USA (nationwide) and the countries of United Kingdom, Germany, Sweden, Spain, Israel and Denmark. | Terminated | May 27, 2015 | 320 | 1315 | Undetermined
Z-2128-2014 | 68794 | No | KLE | Resin tooth bonding agent. | Dental | Dental | 510(k) | Henry Schein Natural Elegance SE Bond Self-Etch Light-Cure |  |  | July 16, 2014 | 2014 | Novocol, Inc. | Septodont is recalling Self-Etch bond because the material was incorrectly packaged. | Process design | An email was sent on 7/16/14 to the two distributors requesting the return of all units of Lot C4211-3 in stock. |  | 60 | 60 | Product was sent to two wholsalers/distributors in PA and NY. No product was distributed outside of the US or to government/VA consignees. | Terminated | December 31, 2014 | 168 | 84 | Undetermined
Z-2127-2014 | 68445 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | Small Bone Innovations 3.0/4.0mm AutoFIX" Cannulated Compression Screw System Sets, Implants, and Instruments |  |  | May 19, 2014 | 2014 | Small Bone Innovations | Small Bone Innovations (SBi) is recalling all lots of 3.0/4.0mm AutoFIX" Cannulated Compression Screw System Sets, Implants, and Instruments to update the IFU. | Process control | Small Bone Innovations sent an Urgent Medical Device Recall letter, dated May 19, 2014, to all affected customers. The letter that described the product problem, and the action to be taken by the customer. was sent to sales reps, stocking customers, and distributors. Customers were instructed that a copy of the IFU should be stored with each unpackaged AutoFIX 3.0/4.0mm Compression Screw Sets, Implants, or Instruments in their current inventory. Customers were instructed to return all packaged affected product which would be replaced by the firm with an updated IFU. A Federal Express return label was included to facilitate the return shipment. Customers were instructed to complete, sign, and return with the product the Return Material Authorization (RMA) document.For questions regarding this recall call 215-428-1791. |  | 506 | 506 | Nationwide Distribution including TN, OK, AZ, PA, TX, MO, NY, NC, CA, FL, WA, AL, IN, NH, SD, and OH. | Terminated | October 08, 2014 | 142 | 506 | Undetermined
Z-0135-2015 | 68201 | No | JJX | Quality control material (assayed and unassayed). | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | HbA1c Calibrators LN 4P52-02 consists of: 4P52-02 Hemoglobin A1c 1 x 1.6 mL and 1 x 1.6 mL. A1c Calibrators (lyophilized) contain hemoglobin and glycated hemoglobin from human whole blood. Prior to lyophilization, the calibrator matrix is an MES-buffered solution. Preservative: Ofloxacin. |  |  | April 30, 2014 | 2014 | Abbott Laboratories | HbA1c Calibrators LN 4P52-02, lot 45063UQ12 may contain a value sheet for a different HbA1c calibrator lot. If calibrator lot 45063UQ12 is used in the calibration of the HbA1c assay using values from an incorrect calibrator value sheet then all results, including quality control and patient, may be falsely elevated. | Under Investigation by firm | Product Correction letters dated April 30, 2014 were sent to all customers. The letters included isntructions for customers to: 1) verify that the lot numbr listed on each calibrator bottle agrees with the lot number printed on the value sheet; 2) If HbA1c Calibrators lot 45063UQ12 has already been utilized in your laboratory, verify the concentrations configured for this calibrator lot match the values in the letter; 3) If these do not match, use the values in the letter when calibrating the HbA1c assay using calibrator lot 45063UQ12; and, 4) follow your laboratory protocol regarding the need for reviewing previously reported patient results if incorrect calibrator values were used. |  | 183 | 183 calibrator kits | Worldwide Distribution in the countries of Austria, Canada, Cyprus, Denmark, Finland, France, Germany, Greece, Hong Kong, Ireland, Italy, Kuwait, Malaysia, Norway, Poland, Romania, Russia, Saudi Arabia, Thailand, Turkey, and United Arab Emirates. | Open, Classified |  |  | 183 | Undetermined
Z-0103-2015 | 69292 | No | LOM |  |  | Microbiology | PMA | ADVIA CentaurÂżÂż Systems HBc Total (HBcT) Kit. For In Vitro Diagnostic Use. |  |  | September 15, 2014 | 2014 | Siemens Healthcare Diagnostics, Inc | Increased imprecision and positive bias is observed with patient samples and negative controls and can result in an increase in initial reactive results. | Unknown/Undetermined by firm | Siemens Healthcare issued An Urgent Medical Device Recall (UMDR) sent to all affected Siemens Healthcare Diagnostics customers in the United States on September 15, 2014. An Urgent Field Safety Notice (UFSN) was sent out for distribution to all customers outside the United States on September 15, 2014. These notices inform the customer of the imprecision and positive bias and a potential increase in initial reactive results associated with ADVIA Centaur Systems HBc Total Lot 100064. Customers were instructed to discontinue use of the affected lots and dispose of the product according to local regulations or procedures. The notices should be reviewed with the facility's laboratory or medical director.Each customer will be required to return a completed confirmation fax-back to Siemens. |  | 5434 | 5434 | Worldwide Distribution-USA (nationwide) and the countries of Austria Bahrain Belgium Croatia Cyprus Czech Republic Denmark Estonia Finland France Germany Great Britain Greece Hungary Iraq Ireland Italy Jordan Latvia Lebanon Lithuania Luxembourg Netherlands Norway Poland Portugal Romania Slovakia Spain Sweden Switzerland Turkey Unit.Arab Emir and Yemen. | Terminated | December 28, 2015 | 469 | 5434 | Undetermined
Z-0115-2015 | 66778 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | Symbia E is a multi-purpose SPECT system ideal for hospitals and outpatient centers. |  |  | October 28, 2013 | 2013 | Siemens Medical Solutions USA, Inc. | Reports of loose bolts in the detector support structure of Symbia and E.CAM systems. | Process control | Undated "Customer Safety Advisory Notice" letters were sent to all direct accounts (customers). The letters stated the problem, and included instructions for customers. Customers: 1) can continue to use the devices while waiting for an inspection to be scheduled and performed; and, 2) their local service engineer will inspect the detector pivot connections and make any necessary repairs on site. Customers with any questions can contact their local service representative or Siemens Molecular Imaging at 800-888-7436. |  | 562 | 562 devices | Worldwide Distribution -- USA, including the states of AK, AL, AR, AZ, CA, CO, CT, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, and WY; and, the District of Columbia, Guam, and Puerto Rico; and, the countries of Albania, Algeria, Angola, Argentina, Austria, Australia, Azerbaijan, Bahrain, Bangladesh, Belarus, Belgium, Bolivia, Bosnia and Herzegovina, Botswana, Brazil, Bulgaria, Cameroon, Canada, Chile, China, Columbia, Congo, Costa Rica, Cuba, Croatia, Czech Republic, Denmark, Dominican Republic, Ecuador, Egypt, El Salvador, Estonia, Finland, France, Georgia, Germany, Greece, Honduras, Hong Kong, India, Indonesia, Iran, Iraq, Ireland, Israel, Italy, Japan, Jordan, Kazakhstan, Kuwait, Latvia, Lebanon, Luxembourg, Macedonia, Madagascar, Malaysia, Mexico, Morocco, Myanmar, Nepal, Netherlands, New Zealand, Niger, Nigeria, Norway, Oman, Pakistan, Palestine, Peru, Philippines, Poland, Portugal, Qatar, Romania, Russia, Saudi Arabia, Serbia, Singapore, Slovakia, Slovenia, South Africa, South Korea, Sri Lanka, Sweden, Switzerland, Syria, Taiwan, Thailand, Trinidad and Tobago, Tunisia, Turkey, Uganda, Ukraine, United Arab Emirates, United Kingdom, Uruguay, Uzbekistan, Vietnam, and Yemen. | Open, Classified |  |  | 5424 | Undetermined
Z-0109-2015 | 69418 | No | LPH | Hip joint metal/polymer/metal semi-constrained porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Microseal Augmented (20 degrees) Liner, acetabular liner component made from UHMW polyethylene (ASTM F648). Sizes 36x52x20; and 36x50x20. |  |  | September 23, 2014 | 2014 | Signal Medical Corporation | Incorrect dimension on four liners was detected | Process design | Signal Medical Corporation verbally notified the only consignee via telephone. Please call Signal Medical Corp. at 1-800-246-6324 with any questions. |  | 12 | 12 devices | US Distribution in MO only. | Terminated | January 20, 2015 | 119 | 12 | Undetermined
Z-0113-2015 | 69204 | No | FOZ | Intravascular catheter. | General Hospital | General Hospital | 510(k) | Bionector Needleless connector. Bionector is a multi-purpose catheter accessory; a closed needle-less system permitting blood sampling, intermit ant injection and continuous infusion of fluids or medications to patients in a hospital or homecare situation. |  |  | August 28, 2014 | 2014 | Vygon Corporation | Vygon has uncovered potential problems with membrane elasticity in the 896.038, Bionector Needleless connector and is issuing a voluntary recall of the 896.038, Bionector Needleless connector. | Component design/selection | Vygon sent an Urgent Medical Device Recall letter dated August 28, 2014, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to check their stock for the affected product and immediately cease distribution, and quarantine all affected product immediately. Count and document the affected inventory on the "Recall Acknowledgement and Inventory Return Form. Customers were asked to fax or email the form to Vygon including customer contact information. Vygon customer service will contact the customer and issue a Return Merchandise Authorization number. Customers were instructed to return the product as well as the original signed Recall Acknowledgement and Inventory Return Form using the instructions from Customer Service involving product pick-up. All inventory will be credited at no cost to the customer.For questions regarding this recall call 800-473-5414. |  | 42250 | 42250 | Worldwide Distribution - US including CA and Internationally to CA. | Terminated | March 30, 2015 | 214 | 42250 | Undetermined
Z-0108-2015 | 69499 | No | DYB | Catheter introducer. | Cardiovascular | Cardiovascular | 510(k) | St Jude Medical, 10F Fast-Cath Trio" Hemostasis Introducer, REF # 406308, Sterile EO. |  |  | October 14, 2014 | 2014 | St Jude Medical | St. Jude Medical is conducting a voluntary recall of SMJ Batch # 4641580 of the Fast-Cath Trio" Hemostasis Introducer, Catalog number 406308. The inner diameter of the 10F dilator is too small, such that the guidewire cannot advance through the device which may result in a procedural delay while another dilator of the desired size is obtained. | Process control | St. Jude Medical sent an Urgent Medical Device Recall Notice dated October 14, 2014, to all affected consignees. The letter described the problem and the product affected by the recall. Advised consignees to discontinue us of the device. A St. Jude Medical sales representative were to contact the consignee in the near future to facilitate removal and reconciliation of the recalled device. Customers with questions were instructed to contact their St. Jude Medical sales Representative.For questions regarding this recall call 651-756-2295. |  | 220 | 220 | Nationwide Distribution including FL, IN, ME, MA, NY, NY, OH, RI,TX, and WV. | Terminated | March 25, 2015 | 162 | 220 | Undetermined
Z-0118-2015 | 69236 | No | CAF | Nebulizer. | Anesthesiology | Anesthesiology | 510(k) | The Aeroneb Professional Nebulizer System is a portable medical device for multiple patient use that is intended to aerosolize physician-prescribed solutions for inhalation to patients on and off ventilation or other positive pressure breathing assistance. The Aeroneb Professional Nebulizer System is suitable for use in adult and pediatric patients. |  |  | September 02, 2014 | 2014 | Aerogen Ltd. | Aerogen has received a small number of complaints in relation to the main power adapter manufactured by FRIWO GmbH and used in the Aeroneb Solo and Pro Nebuliser systems. The complaints received identified that the power adapter cover had broken and dislodged from the body of the adapter prior to use. | Nonconforming Material/Component | URGENT FIELD SAFETY NOTICE letters dated September 2, 2014 were sent to all direct accounts (customers). The letters included instructions for customers to: 1) compare their power adapter date codes against those found in the letter as being subject to this recall; 2) until new power adapters are received, customers may continue to use the recalled power adapters if the customer inspects each power adapter prior to use and finds that there is no evidence of damage or separation of the housing and unplug the adapter after each use; and, 3) contact the Aerogen or appropriate Aerogen sales partner immediately so that the power adapter can be replaced free of charge. The letter further requests that the recall notification be forwarded to any parties that may have received the recalled power adapters. Customers with questions can contact Aerogen at 866-423-76436. |  | 2421 | 2,421 power adapters | Worldwide Distribution-USA (nationwide) including the states of CA, GA, IL, MS, NC, and TX, and the countries of Australia, Belgium, Canada, China, Czech Republic, Denmark, Egypt, Finland, France, Germany, Hungary, India, Ireland, Israel, Italy, Kuwait, Libya, Malaysia, Netherlands, Norway, Poland, Qatar, Russia, South Korea, Singapore, Sweden, Switzerland, Thailand, Turkey, United Arab Emirates, United Kingdom, and Vietnam. | Terminated | April 25, 2016 | 601 | 4003 | Undetermined
Z-0123-2015 | 69354 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | The Persona Tibial Articular Surface Inserter is a sterilizable instrument intended for |  |  | October 07, 2014 | 2014 | Zimmer, Inc. | Affected product did not undergo heat treatment hardening step leading to the potential for deformation of the tip. | Process control | Zimmer sent an " URGENT MEDICAL DEVICE RECALL - LOT SPECIFIC " notifications dated September 2014, to all affected customers. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. If after reviewing this notification you have further questions or concerns please call the customer call center at 1-877-946-2761 between 8:00 am and 5:00pm EST. |  | 618 | 618 units | Worldwide Distribution - USA (nationwide) including the states of NY, OR, NC, OH, TX, FL, CO, IN, TN, PA, CA, KS, MO, WI, VA, MN, NJ, UT, MA, GA, MI, OK, AZ, MS, IL and AL., and the countries of Korea, Singapore, Canada and Japan | Terminated | September 23, 2015 | 351 | 618 | Undetermined
Z-0122-2015 | 69321 | No | CGR | Cortisol (hydrocortisone and hydroxycorticosterone) test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | IMMULITE /IMMULITE 1000 Systems COR Cortisol, REF/Catalog Number LKC01, SMN 10381388, IVD. |  |  | September 19, 2014 | 2014 | Siemens Healthcare Diagnostics | Customer complaints were received for positive bias reported on the controls used on the IMMULITE /IMMULITE 1000 on the Cortisol Assay reagent lots IMMULITE /IMMULITE 1000 (LKCO1) Lots 381,382 and 383 . Subsequently, the firm confirmed a positive bias on the IMMULITE/IMMULITE 1000 (LKCO1) Lots 381, 382 and 383. | Nonconforming Material/Component | US Customers were sent an Urgent Medical Device Recall Letter (UMDR2014-09-19, dated 9/19/14) via FedEx on 9/19/14. Outside US Customers were sent an Urgent Field Safety Notice Letter. The letters inform the consignees of the recall of the IMMULITE/IMMULITE 1000 Cortisol (LKCO1) assay, kit lots 381, 382 and 383. Customers were asked to take the following actions: - Please review this letter with your Medical Director. - Discontinue use of and discard the IMMULITE/IMMULITE 1000 Cortisol kit lots listed in the letter. - Review your inventory of these products to determine your laboratorys replacement needs and to provide information to Siemens for reporting to the authorities. --- Customers are requested to complete and return the Effectiveness Check questionnaire attached to the Urgent Field Safety Notice letter within thirty (30) days. |  | 815 | US: 772 units, Canada: 43 units | Worldwide Distribution: US (nationwide) and internationally to Canada. | Terminated | November 18, 2015 | 425 | 815 | Undetermined
Z-0129-2015 | 69449 | No | DBF | Ferritin immunological test system. | Immunology | Immunology | 510(k) | ADVIA Centaur Calibrator C-Used with the ADVIA Centaur Ferritin assay on the ADVIA Centaur CP system only. (In-Vitro Diagnostic) |  |  | October 09, 2012 | 2012 | Siemens Healthcare Diagnostics, Inc | Ferritin Analytical Sensitivity and Calibrator C Realignment with kit lots ending in 22 and 23 used with the ADVIA Centaur Ferritin assay on the ADVIA Centaur CP system | Nonconforming Material/Component | Siemens issued an Urgent Field Safety Notice via e-mail to all US customers affected on October 9, 2012. This notice informs the customer of the Field Correction and the actions that need to be taken. The Urgent Field Safety Notice instructs customers that upon receipt of Calibrator C Lots ending in 25 and higher (Cal C lots CC25LA and higher and Cal C lots CC25HA and higher), discontinue use of the Calibrator C lots ending in 22 and 23 with the ADVIA Centaur Ferritin assay when used on the ADVIA Centaur CP system. A confirmation fax-back form is included to ensure that all customers worldwidehave been notified and have complied with the mandatory action. Further questions please call (508) 668-5000 |  | 2440 | 2440 kits | Worldwide Distribution - Nationwide Distribution and the countries of Argentina, Australia, Brazil, Canada, China, Columbia, Egypt, Hong Kong, India, Israel, Indoenesia, India, Japan, Korea, Mexico, Malaysia, Philippines, Saudi Arabia, Singapore, Taipei, Uruguay, Venezuela, Viet Nam and South Africa. | Terminated | June 18, 2015 | 982 | 10937 | Undetermined
Z-0095-2015 | 68047 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | The GemStar Power Supply is an accessory for the GemStar Infusion Pump. It is a two piece assembled switching power supply. The Power Supply converts alternating current (AC) line input voltage of 120 or 240 V to an output to the device of 3.3 Volts DC. |  |  | February 03, 2014 | 2014 | Hospira Inc. | The GemStar Infusion pump may not receive direct current (DC) power from the power supply. In one instance, smoke was observed and the GemStar pump was operating on battery power while connected to the 3 volt DC power supply. | Process control | An URGENT MEDICAL DEVICE CORRECTION letter dated February 3, 2014 was sent to all direct accounts. The letter included specific instructions for how to handle and use the power supplies so as to minimize the chance of any device failures. There were additional instructions for customers to: 1) Inform potential users in their organization of the product of this notification; 2) Complete the attached reply form and return it to the fax number or e-mail address on the form (even if you do not have any affected product in inventory); 3) If the products were further distributed, their customers should be notified of the recall and asked to contact Stericycle at 888-641-9735 to receive a reply form; and, 4) if a potential user believes that their power supply is damaged for any reason, it should be removed from service and the user should contact Hospira Global Complaints at 800-441-4100 to report the concern. The letter further notes that there is no need to return fully functional power supplies (unless they malfunction). |  | 978 | 978 power supplies | Worldwide Distribution -- USA, including the states of AK, AL, AR, AZ, CA, CO, CT, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, and WY; and, the countries of Austria, Belgium, Brazil, Columbia, France, Germany, Greece, Netherlands, Saudi Arabia, South Africa, Spain, Sweden, Switzerland, Turkey, and United Kingdom. | Terminated | March 07, 2017 | 1128 | 13112 | Undetermined
Z-0106-2015 | 69436 | No | GZL | Depth electrode. | Neurology | Neurology | 510(k) | Box Label: 3 mm Cueva Electrode Kit, Sterile EO, Rx Only, REF. CNE1-3KDINX . |  |  | October 08, 2014 | 2014 | Ad-Tech Medical Instrument Corporation | The recall has been initiated due to concerns that the applicator wand may malfunction potentially causing the inability to monitor the VII cranial nerve. . | Process control | Consignees were sent on 10/8/2014 an AD-TECH "Medical Device Recall" letter dated October 6, 2014. The letter described the Reason for the Voluntary Recall, Risk to Health, Action to be taken by the Customer, Product and Distribution Information and Other Information. Requested consignees to return the "Acknowledgement and Receipt Form" to their Ad-Tech Clinical Specialists Fax 262-634-5668, Telephone 262-634-1555, customersupport@adtechmedical.com. For additional information they can use the same number. |  | 13 | 13 kits (26 electrodes, 26 applicator wands, 26 leadwires). | Worldwide Distribution: US (nationwide) in the states of: CO, FL, LA, NJ, OH, WA, and WI; and internationally to: Finland. | Terminated | April 08, 2015 | 182 | 13 | Undetermined
Z-0134-2015 | 69160 | No | LPL | Soft (hydrophilic) contact lens. | Ophthalmic | Ophthalmic | 510(k) | ProClear Toric contact lenses are made from a material containing 62% water and 38% omafilcon B, a random copolymer of 2-hydroxyethylmethacrylate and methacrylic acid. The lenses have a tint (VAT Blue 6) which is added to make the lens more visible for handling. |  |  | September 17, 2014 | 2014 | Cooper Vision Caribbean Corp. | Routine quality monitoring system has identified that a limitednumber of lots were manufactured with the cylinder axis outside the allowable specification. | Other | The firm sent out a notification letter with the following instructions:Please take the following actions immediately:" Examine your inventory for any affected lot numberso Use the enclosed report to inspect your inventory for affected product that is still in yourWarehouse." Stop any further distribution of the specified lots of Preclear T oric lenses." Quarantine any Preclear Toric lenses with the specified lot numbers." Within the next five (5) business days, please complete the attached Recall Response Form andfax it to 866-484-1048 or scan and email to info@coopervision.com.o If you have indicated that you have product to return, you will be contacted by aCooperVision Representative with instructions for pick-up and removal of affected,quarantined product. Your account will be credited once the product is returned andprocessed.This action is being taken in full cooperation with the Food and Drug Administration (FDA). The recall onlyaffects Proclear T oric lenses with the lot numbers specified on the enclosed list. No other Preclear T oricproducts- or any other CooperVision products - are affected. |  | 3355194 | 13639 lenses/246 product lots | Worldwide Distribution: US (nationwide) and internationally to: Canada, Latin America, EU, Asia Pacific, and Africa/Central Eurasia countries. | Terminated | June 10, 2015 | 266 | 3355194 | Undetermined
Z-0126-2015 | 69347 | No | LYL | Soft (hydrophilic) contact lens care products. | Ophthalmic | Ophthalmic | 510(k) | Equate Moisture Last Multi-Purpose Solution, 2 pack, 12 fl. oz. (355 mL), Distributed by Walmart Stores, Inc. |  |  | September 25, 2014 | 2014 | Bausch &amp; Lomb Inc- Greenville Solutions Plant | Product lot did not meet the specification for its disinfectants at the time of manufacture. | Other | Consignees were notified via letter starting 09/25/2014. |  | 18792 | 18,792 bottles (9,396 twin packs) | US Nationwide Distribution in states of: AL, AR, AZ, CA, CO, DE, FL, GA, IA, IL, IN, KS, LA, MS, NC, NH, NY, OH, OR, PA, TN, TX, UT, VA, and WI. | Terminated | June 12, 2015 | 260 | 18792 | Undetermined
Z-0136-2015 | 69293 | [6 empty] | Animas Vibe Insulin Infusion Pump. |  |  | June 10, 2014 | 2014 | Animas Corporation | Animas determined that the Hebrew language translation in the Vibe Insulin Infusion Pump was incorrect and could lead to potential patient confusion. In addition, one line of Hebrew text was omitted from the display screen. | Employee error | A recall notification letter, dated June 2014, was sent to the recalling firm's sole customer, TriTech Biomed, in Israel, to notify them about the product, problem, and action to be taken. The distributor sent letters to their customers on July 16, 2014. |  | 1028 | 1028 | International Distribution in Israel only. | Terminated | September 02, 2015 | 449 | 1028 | Undetermined
Z-0139-2015 | 69532 | No | LPB |  |  | Cardiovascular | PMA | ThermoCool SmartTouch Navigation Catheters, Catalog No. D132701, D132702, D132703, D132704, D132705, D133601, D133602, D133603. |  |  | October 19, 2014 | 2014 | Biosense Webster, Inc. | The recall was initiated because Biosense Webster is providing additional labeling for the safe and effective use of the ThermoCool SmartTouch Catheter. | Use error | A customer notification letter dated October 19, 2014, will be sent to all customers who purchased the Biosense Webster's ThermoCool SmartTouch Catheter Family. The letter informs the customers of the additional information for the safe and effective use of the ThermoCool SmartTouch Catheter, which will be included in the updated labeling. The letter informs the customers of the problems identified and the actions to be taken. Customers with questions related to the recall letter are instructed to contact their Biosense Webster sales representative or call (866) 473-7823, Monday through Friday from 7am to 8pm EST. |  | 173329 | 173,329 units total (21,812 units in US) | Nationwide Distribution | Terminated | October 06, 2015 | 352 | 173329 | Undetermined
Z-0140-2015 | 69373 | No | NEW | Absorbable polydioxanone surgical suture. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Monoswift PGCL [poly(glycolide-co-caprolactone)] Absorbable Surgical Suture; USP size 3/0. Product is labeled sterile. The product is sealed inside an inner foil pouch with an outside Tyvek pouch. It is an undyed, monofilament synthetic absorbable suture with needle size 3/0. |  |  | September 22, 2014 | 2014 | C P Medical, Inc. | CP Medical Inc. is recalling five lots of Monoswift PGCL [poly(glycolide-co-caprolactone)] Absorbable Surgical Suture for failing to meet the minimum tensile strength as dictated by USP was not met over the length of the product's shelf life. | Nonconforming Material/Component | CP Medical Inc.notified their customers between September 23 through October 3, 2014, of the recall via telephone or e-mail and followed up with a certified letter. Customers were advised to discontinue sale and/or use of the recalled product and return all existing stock to CP Medical. CP Medical also requested their customers to notify their downstream customers of the recall and request they return any unused recalled product, as well as provide information to them regarding distribution of the product. Customers are being asked to fill out and return a questionnaire regarding the recalled product and contact CP Medical Customer Service Representative at 1-800-950-2763 for return authorization and shipping account number. All shipping costs including those from downstream customers will be paid for by CP Medical. Replacement product or credit will be provided. |  | 888 | 74 boxes /12 sutures per box; plus an additional 13 sutures | US Distribution including the states of: AR, AZ, CA, FL, ME, MI, NE, OR, SC, and WA | Terminated | February 17, 2015 | 148 | 888 | Undetermined
Z-0144-2015 | 69265 | No | FQO | Operating tables and accessories and operating chairs and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | IMRIS Operating Room table (ORT200 and ORT300). |  |  | October 01, 2014 | 2014 | IMRIS Inc | One of the three pins that is used to hold the ORT200 table alignment block in place on the floor of the MR Suite went missing during testing at a health care facility. | Use error | Consignees were sent on 10/01/2014 an IMRIS "Urgent Field Safety Notice" letter dated September 25, 2014. The letter described the problem and the product involved in the recall. The letter informed consignees that the pins will be replaced with pins made of non-ferrous material. Customers are to ensure that users are aware of the changes and the existing procedure for table alignment and managing ferrous content in the magnet room. Contact customer support at 1-866-475-0525 for questions. |  | 14 | 14 | Worldwide Distribution -- USA, including the states of CA, FL, NH, NC, TN, and UT; and, the countries of AUSTRALIA, CANADA, and JAPAN. | Terminated | February 20, 2015 | 142 | 14 | Undetermined
Z-0148-2015 | 69372 | No | DIP | Diphenylhydantoin test system. | Toxicology | Toxicology | 510(k) | siemens DimensionÂż Phenytoin FlexÂż Reagent Cartridge |  |  | September 24, 2014 | 2014 | Siemens Healthcare Diagnostics, Inc. | Internal testing has confirmed a low absorbance range between the levels 1 and 5 calibrators with this lot resulting in the potential for a positive or negative bias and imprecision across the PTN assay range. | Under Investigation by firm | Siemens sent an Urgent Medical Device Recall Notification dated September 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Customers were instructed as follows: " Discontinue use and discard any remaining inventory of PTN lot FA5154, and complete the attached form indicating your replacement product needs in the attached Field Correction Effectiveness Form. " Siemens will replace any unused inventory of the affected lot at no charge.We apologize for the inconvenience this situation may have cause. If you have any questions, please contact Siemens Customer Care Center - Technical Solutions at (800) 441-9250, or your local Siemens technical support representative. |  | 1733 | 1733 | Worldwide Distribution - USA (nationwide) including Puerto Rico., and the countries of Australia, Brazil, Bahamas, Canada, Chile, India, Japan and Uruguay | Open, Classified |  |  | 1733 | Undetermined
Z-0149-2015 | 69369 | No | JQW | Pipetting and diluting system for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Freedom EVO-2 100 Base Unit, with Air LiHa Arms |  |  | September 26, 2014 | 2014 | Tecan US, Inc. | Distortion (noise interference) of the liquid level capacitance can result in air aspiration instead of liquid aspiration. This might cause a false signal (air detection) with possible dispense errors resulting in erroneous results. | Nonconforming Material/Component | Tecan sent an Urgent Field Corrective Action letter dated September 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Tecan has instructed field engineers to replace the ZAAP Motion Board in the Freedom EVO instrument. Your instrument was identified as affected and a Tecan Service Representative will contact you to arrange replacement of the board.Please make sure that your internal controls would detect any possible dispense errors that might affect your application when using the instrument before replacement of the board. Alternatively, stop using the instrument until the board is replaced.We apologize for any inconvenience this issue may have caused,and are implementing corrective and preventive actions to prevent recurrence. If you have any further questions, please contact your local Tecan Helpdesk for assistance.For further questions please call (919) 361-5200 ext. 19519 |  | 2 | 2 units | US Distribution to IL. | Terminated | July 10, 2017 | 1018 | 2 | Undetermined
Z-0150-2015 | 69433 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | APTUS Ulna Shortening 2.5 |  |  | October 01, 2014 | 2014 | Medartis, Inc. | A recall of the APTUS Ulna Shortening 2.5 instrument was initiated due to three (3) complaints of breakage of the temporary tension bolt. | Process design | The recall is being extended to end user level of distribution. The sales representatives, customers and a distributor were notified via phone or email on September 19th and notified of the recall via US mail on 10/1/14. US customers, distributors, and sales representatives were instructed to return the affected instrument kits to the importer Medartis, Inc. via transport agency. |  | 44 | 44 | US Nationwide Distribution in the states of WV, IL, CT, AZ, and WA. | Terminated | August 06, 2015 | 309 | 44 | Undetermined
Z-0008-2015 | 69313 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Esprit V1000 and V200 Ventilators, Model No. V1000 and V200; and the 3rd Generation Power Supply Repair Part (RP)Kits |  |  | September 17, 2014 | 2014 | Respironics California Inc | A failure of a specific component may prevent the ventilators from operating on AC power or transitioning back to AC power when operating on battery. Additionally, if a battery is not present or is depleted the ventilators will not operate. This failure will cause the loss of ventilator support, which may result in hypercarbia or hypoxemia. | Unknown/Undetermined by firm | Philips Healthcare sent an Urgent - Field Safety Notice dated September 17, 2014, to all customers who purchased the Esprit V1000 and V200 Ventilators to inform them that Respironics is recalling Esprit V1000 and V200 Ventilators installed with 3rd Generation Power Supplies and also 3rd Generation Power Supply Repair Part Kits. The letter identifies the affected devices, informs the customers of the problems, hazard involved, and the actions to be taken. Customers were reminded to be advised of the warnings provided. A Philips Respironics representative will contact customers to schedule a no-cost replacement for further information or support.Customers with questions were instructed to contact Philips Healthcare at 1-800-722-9377.For questions regarding this recall call 760-918-7300. |  | 2206 | 2,206 units total (665 units in US) | Worldwide Distribution - USA, Argentina, China, Egypt, Ethiopia, Honduras, Hong Kong, India, Indonesia, Iraq, Israel, Italy, Jordan, Libyan Arab Yamahiriya, Malaysia, Mali, Mexico, Myanmar (Burma). Nepal, Nigeria. Pakistan, Peru, Philippines, Serbia, Singapore, South Africa, Sri Lanka, United Republic of Tanzania, Thailand, Vietnam, Zimbabwe, Republic of Korea, and Netherlands. | Open, Classified |  |  | 2206 | Undetermined
Z-0098-2015 | 69319 | No | KDQ | Vacuum-powered body fluid suction apparatus. | General Hospital | General Hospital | 510(K) Exempt | DeRoyal(R) SafeLiner Suction Canister Kit, SEMI-RIGID, 1000cc, REF 71-6518, NON-STERILE, Rx Only, |  |  | August 22, 2014 | 2014 | DeRoyal Industries Inc | These lots of suction canisters and lids have an insufficient seal under low or intermittent suction condition. | Process control | DeRoyal sent an Urgent Voluntary Recall letter dated August 22, 2014, to all affected consignees by overnight letter or by email. DeRoyal was to follow up with non-responding accounts before October 1, 2014. Customers were instructed to check their inventory for the affected product. If they have affected product, customers were instructed to destory both the liner and the lid and dispose of it according to their facility's procedures. Customers were instructed to return the included Notice of Destruction to recalls@deroyal.com even if they no longer have affected product. Customers with questions were instructed to contact their DeRoyal Sales Representative or DeRoyal Customer Service at 1-800-251-9864.For questions regarding this recall call 865-362-2334. |  | 5112 | 5112 units | Nationwide Distribution including MI, OH, IN, FL, WV, VA, MI, NC, and NY | Terminated | April 15, 2015 | 236 | 19012 | Undetermined
Z-2133-2014 | 68400 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | DimensionÂż Integrated Chemistry Systems intended to measure a variety of analytes in human body fluids. |  |  | May 27, 2014 | 2014 | Siemens Healthcare Diagnostics, Inc. | There is a misassembly issue with thephotometer installed on DimensionÂż Integrated Chemistry Systems instruments. Through internal investigation bySiemens Healthcare Diagnostics has confirmed that a limited number of Dimension photometers were misassembled. These photometers are performing within specifications, but may lead to a shortened life of the source lamp. | Employee error | An urgent field safety notice, dated May 2014, was sent to users to inform them of the issue and that a service representative would replace the misassembled photometer. |  | 15 | 15 | FLLAARIAOHOHWIILLAKSMT | Terminated | December 15, 2014 | 202 | 15 | Undetermined
Z-0022-2015 | 69197 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Philips Healthcare Integris Allura 15 -12 Fluoroscopy x ray system |  |  | August 08, 2014 | 2014 | Philips Medical Systems, Inc. | Philips had discovered through a field service testing where a device failed to comply with a Federal standard. | Radiation Control for Health and Safety Act | Philips Healthcare discovered through field service testing, that the device was in violation. Firm service rep. visited site on 7/01/14 and corrected unit.The field service engineer immediately reprogrammed the generator for the X-ray tube installed in the system and re-adapted the x-ray tube to bring the system back into compliance.At this time the Failure to comply has been corrected and the customer has been notified. This correction is to be free of charge to the customer.For further questions please call (978) 687-1501. |  | 1 | one system | US distribution to FL | Terminated | December 03, 2014 | 117 | 1 | Undetermined
Z-0006-2015 | 69024 | No | CIT | Aspartate amino transferase (AST/SGOT) test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | ENVOY 500 AST Reagent Kit, reference Model/Catalog 55255, for Envoy 500. |  |  | August 11, 2014 | 2014 | ELITech Clinical Systems SAS | Some users of ENVOY 500 AST Reagent Kit, reference 55255, for Envoy 500 systems are observing that Quality Control (QC) fails before the product expiry is reached. | Process control | ELITech sent an Urgent Product Recall Notice letter dated August 12, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. To All Users,***Effective immediately, please discontinue using this product.***Actions to be taken:"Laboratories must not use this product and must destroy all remaining inventory of the product."If you have transferred any of this product to another location, please forward a copy of this bulletin to that location."All users receive ENVOY 500 AST REAGENT KIT, reference 55255, from distributors. To be certain everyone who is possibly affected by this is contacted, we are sending this communication to all registered users.As confirmation of receipt of this notice, please return the attached confirmation receipt to Technical Support by fax (401-642-9003) within 48 hours of receipt.On August 20, 2014:, customers were contacted with an interim solution.On September 3, 2014, the firm sent Technical Bulletin 5056 and Distributor Notice D057/Customer Notice C061 to all Envoy users notifying that the recalled Envoy AST Reagent (ref 55255) was being discontinued; the old Envoy AST Reagent (ref. 55250) was being reinstated.On September 23, 2014, the firm sent the URGENT: PRODUCT RECALL NOTICE "UPDATE", dated 19 September 2014, to all Envoy users stating that old (original) Envoy 500 AST Reagent Kit (ref 55250) will replace the recalled Envoy 500 AST Reagent Kit (ref 55255).As confirmation of receipt of this notice, please return the attached confirmation receipt to Technical Support by fax (401-642-9003) within 48 hours of receipt.Customers with questions about this recall, please call the Technical Support at 1-855-354-8324 (1-855-ELITECH). |  | 878 | 878 kits | Nationwide Distribution | Terminated | December 11, 2014 | 122 | 878 | Undetermined
Z-0009-2015 | 69064 | No | EHD | Extraoral source x-ray system. | Dental | Radiology | 510(k) | eXpertDC intraoral x-ray, models: eXpert DC, NA, 75 Inch reach; eXpert DC, NA, 65 Inch reach; eXpert DC, CN, 55 Inch reach; eXpert DC, WW, 75 Inch reach; eXpert DC, WW, 65 Inch reach; eXpert DC, WW, 55 Inch reach; OEM Arm Assembly |  |  | August 12, 2014 | 2014 | Gendex Corp | Two (2) failures of the spring link components within the articulated arm assembly. The spring links broke during normal assembly process for the articulated arm during first cycle of the arm, prior to the 14lb tube head (full load) being installed onto the assembled articulated arm. | Nonconforming Material/Component | Gendex sent a Recall Notification Letter, dated August 12, 2014 and a second Communication Notification letter on September 22, 2014 to all affected customers. The letter identified the affected product, problem, and actions to be taken. For questions contact Gendex Customer Service at 1-800-323-8029. |  | 651 | 651 | Worldwide Distribution - US Nationwide and the countries of Europe, Canada, Australia, China, Russia, Taiwan, Singapore, and Thailand. | Terminated | March 24, 2017 | 955 | 651 | Undetermined
Z-0010-2015 | 69322 | No | JAA | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | GE Healthcare Precision 500D, Classical R&F; System. |  |  | September 05, 2014 | 2014 | GE Healthcare, LLC | GE Healthcare has become aware of a potential safety issue involving missing screws used for mounting the overhead Video Monitor to the tray assembly on a single monitor suspension. This issue affects the Precision 500D X-ray imaging systems. | Employee error | Consignees were sent on 9/5/2014 a GE Healthcare "Urgent Medical Device Letter" GE Ref:10893, dated September 5, 2014.The letter was addressed to Hospital Administrators / Risk Managers, Managers of Radiology / Cardiology &amp; Radiologists / Cardiologists. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction and Contact Information. For questions in USA call 800-437-1171, for other countries contact your local GE Healthcare Service Representative. |  | 204 | 204 (199 US, 5 OUS). | Worldwide Distribution - US Nationwide - AL, AZ, AR, CA, CO, CT, FL, GA, HI, ID, IL, IN, IA, KS,KY, LA, MD, MA, MI, MN, MS, MO, NE, NV, NH, NJ, NM, NY, NC, OH, OR, PA, SC, TN, TX, UT, VA, WA, WV, WI and in the countries CANADA, GERMANY, UNITED KINGDOM. | Terminated | February 20, 2015 | 168 | 204 | Undetermined
Z-0023-2015 | 69249 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Biomedix SELEC-3 I.V. Administration Set. Individually sealed into a TyvekÂż and LDPE/PET pouch. Secondary packaging includes 48 pouched units per box/case. STERILE |  |  | August 14, 2014 | 2014 | Biomedix, Inc. | Customer contacted Biomedix of complaint and returned the failed set along with 86 units of unopened stock (96 units originally shipped). Biomedix performed extensive investigation and found insufficient application of glue. Recall was initiated due to the rate of the failure. | Process control | Biomedix sent an PRODUCT RECALL NOTIFICATION letter dated August 14, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Please contact us with the quantity remaining in inventory or available to be returned to our facility. All units returned to Biomedix will be replaced with the same SELEC-3 item or a similar design at no charge.Please call (1-800-627-2765 ) for further questions. |  | 2928 | 2,928 | US Distribution including the states of OH, NC, and TX | Terminated | November 13, 2014 | 91 | 2928 | Undetermined
Z-0021-2015 | 69280 | No | HTO | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Compress Face Reamer |  |  | September 17, 2014 | 2014 | Biomet, Inc. | Investigation determined that units supplied were missing the 4x21 degree helix angle. | Process control | On September 17, 2014, "URGENT MEDICAL DEVICE RECALL NOTICE" was sent to the consignee Biomet Global Supply Chain Center (GSCC) by email and FedEx. GSCC notified all of the customers within Europe, the Middle East, and Asia that received the affected product. The recall notification included a description of the reason for recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. For questions, call (574) 372-1570. |  | 4 | 4 | International Distribution in the countries of Netherlands, Europe, the Middle East, and Asia. | Terminated | July 14, 2015 | 300 | 4 | Undetermined
Z-0024-2015 | 69149 | No | JIT | Calibrator. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Dimension VistaÂż Mass CKMB Isoenzyme Calibrator (MMB CAL) an in vitro diagnostic product for the calibration of Creatine Kinase MB Isoenzyme (MMB) method on the Dimension VistaÂż System. |  |  | August 28, 2014 | 2014 | Siemens Healthcare Diagnostics, Inc. | Siemens Healthcare Diagnostics has determined that Dimension Vista&reg; Mass CKMB Isoenzyme Calibrator (MMB CAL) shows drift during routine stability monitoring that exceeds our acceptance criteria for drift over the shelf life of this product. Siemens has observed drift from 3% to 12% at MMB concentrations across the assay range. Depending on quality control limits, this drift may not have been d | Process design | An Urgent Medical Device Recall notice, dated August 2014, was provided to all Dimension VistaÂż MMB CAL customers who received the affected lots to notify them of the issue, the potential risk to health, and actions to be taken by customers.Customers were instructed to discard lots 3JD045, 4BD064, 4CD012, and 4DD065, and recalibrate with lot 4FD073. |  | 6562 | 6562 | Worldwided distribution: US (nationwide) and countries of: Australia, Austria, Belgium, Canada, Denmark, Germany, Finland, France, Italy, Japan, Netherlands, New Zealand, Portugal, Saudi Arabia, Slovenia, South Korea, Spain, and Switzerland. | Terminated | March 30, 2015 | 214 | 6562 | Undetermined
Z-0025-2015 | 69085 | No | LJS | Percutaneous, implanted, long-term intravascular catheter. | General Hospital | General Hospital | 510(k) | BioFlo PICC with ENDEXO Technology Maximal Barrier Nursing Kit, 3F-55 cm, UPN Product No. H965750231, REF/Catalog No. 75-023, STERILE --- |  |  | August 25, 2014 | 2014 | Navilyst Medical, Inc | NMI was notified by the manufacturer of the 10 ml BD PosiFlush SF Saline Flush Syringes, that it had initiated a recall on specific lots of the syringes as some syringe unit packages exhibited open seals which may impact package integrity and potentially affect product sterility. Based on the notification, NMI is recalling specific PICC Kits that contain the affected 10 ml BD PosiFlush SF Saline | Nonconforming Material/Component | BD Medical sent an Urgent Voluntary Medical Device Recall letter dated July 22, 2014, to all affected customers via Federal Express on August 25, 2014. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to segregate and quarantine all affected product. Immediately discontinue shipment of the affected product. Complete the Recall Response Card form and fax it back to BD at 1-201-847-6990. If the affected product was further distributed cusotmers should identify their cusotmers and notify them immediately of the product recall. Customers with questions were instructed to contact BD Customer Service at 1-888-237-2762, option 3.For questions regarding this recall call 518-742-4571. |  | 20 | 20 units/pouches | Nationwide Distribution | Open, Classified |  |  | 825 | Undetermined
Z-0061-2015 | 69052 | No | MDM | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Charnley PN Retract &amp; Handle, REF 962004000, Non-Sterile-DuPuy International Ltd. |  |  | August 27, 2014 | 2014 | DePuy Orthopaedics, Inc. | The small extraction peg of the Charnley Pin Retractor and Handle set is breaking from the pin. | Mixed-up of materials/components | The firm, Depuy Synthes, sent an "URGENT INFORMATION - RECALL NOTICE" dated August 27, 2014 to their consignees. The recall notification included a description of the reason for recall, affected product, clinical implications, and instructions for responding to the formal recall notification. The customers were instructed to cease using the affected components immediately, quarantine all affected components in a manner that ensures the affected components are not used; return affected devices immediately to your DePuy; if any affected product has been forwarded to another facility, contact that facility immediately to communicate recall and arrange return; maintain awareness and copy of this notice, and complete and return the Reconciliation forms via fax to 574-371-4939 or emailed to kearle@its.jnj.com. For questions about the device recall information provided, please contact Recall Coordinator, at 574-371-4917 (M-F; 8 a.m. - 5 p.m. EDT) or email Kearle@its.jnj.com.For product-related questions, please contact your local DePuy Orthopaedics, Inc. Sales Consultant. For clinical questions from hospitals/user facilities, please contact DePuy Orthopaedics, Inc. Scientific Information Office at 1-888-554-2482 (M-F; 8 a.m. - 5 p.m. EDT). |  | 65 | 65 units | Worldwide Distribution: US (Nationwide) including states of: AR, CA, CO, CT, IL, MA, NC, OH, PA, TN, TX, UT, VA, WA, WI; and Internationally to: AUSTRALIA, CHILE, UK, INDIA, IRELAND, JAPAN, NEW ZEALAND, RUSSIA, SINGAPORE, and SAUDI ARABIA. | Terminated | November 17, 2015 | 447 | 65 | Undetermined
Z-0018-2015 | 69276 | fixative, formalin-containing | LDY | General purpose reagent. | Pathology | Pathology | 510(K) Exempt | Cardinal Health Prefilled Buffered 10% Formalin. |  |  | September 17, 2014 | 2014 | Richard-Allan Scientific Company | The affected lots could have NBF concentrations that are lower or higher than the desired specifications. | Mixed-up of materials/components | The firm issued an "URGENT MEDICAL DEVICE RECALL" letter to all affected customers. The notification identified the affected products, reason for voluntary recall, risk to health, actions to be taken, and instructions to complete the attached Acknowledgement Form. For questions, contact Sarah Rickert at 269-544-5628 or sarah.rickert@thermofisher.com. |  | 235 | 235 cases | Worldwide Distribution-USA (nationwide) including the states of PA, IL, CA, TX, GA, WA, MA, KY, OH, NY, VA, NC, IN, NJ, MN, UT, HI, FL, KS, TN, MO, MI, MD, LA, NM, DE, CO, WV, OR, WI, AZ, SC, IA, and NE and the country of Canada. | Completed |  |  | 1621 | Undetermined
Z-0059-2015 | 69337 | No | OWQ | Diagnostic intravascular catheter. | Cardiovascular | Cardiovascular | 510(k) | Biosense Webster's Reprocessed SOUNDSTARÂż eco Diagnostic Ultrasound Catheters, Catalog Nos: M-5723-115, Model Number BIO10438577. |  |  | September 24, 2014 | 2014 | STERILMED, INC. | Sterilmed, a division of Johnson &amp; Johnson Medical NV/SA (Sterilmed), has become aware of an issue affecting certain lots of Reprocessed SOUNDSTAR&reg; eco Diagnostic Ultrasound Catheters for which Sterilmed is the reprocessor and Biosense Webster is the distributor. Sterilmed determined that the units of these lots were incorrectly processed and as a result when the affected catheter is conne | Process control | Sterilmed sent an "Field Notification" letter dated September 22, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Actions requested on your part:1. Read the Description of the Problem section below carefully.2.Immediately identify and set aside all units from the affected lots of product listed above in a manner that ensures the affected product will not be used.3. Maintain a copy of this letter with the affected product.4. Review, complete, sign and return the attached Field Removal Certification Form in accordance with the instructions listed on the form.5. Arrange for return of any impacted units from the provided list of affected Reprocessed SOUNDSTARÂż eco Diagnostic Ultrasound Catheters that you may have in your inventory per the instructions on the Field Removal Certification Form.6. Pass on this notice to anyone in your facility that needs to be informed.7. If any of the affected Reprocessed SOUNDSTARÂż eco Diagnostic Ultrasound Catheters have been forwarded to another facility, contact that facility and arrange for the return.8. Maintain awareness of this notice until all affected product with lot numbers listed above has been returned to Biosense Webster.For questions related to this issue and product return please contact your Biosense Webster sales representative. For further questions please call (763) 488-3400. |  | 22 | 22 | US Distribution including the states of WA and CA. | Terminated | August 11, 2015 | 321 | 22 | Undetermined
Z-0092-2015 | 69432 | No | GAD | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | DePuy Synthes Radiolucent Retractor, part number 387.580, lot number 3620715 |  |  | October 02, 2014 | 2014 | Synthes, Inc. | Particle residue on the instrument from adhesive tape which was used to bind the device during transport. There may be a risk of tissue reaction if a patient is exposed to latex particles from the adhesive tape, or a systemic reaction such as anaphylaxis in patients with latex allergy. | Material/Component Contamination | DePuy Synthes sent an Urgent Notice Medical Device Recall letter dated October 2, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Please take the following actions:If you DO have any of the identified devices, please take the following steps:Call DePuy Synthes at 1-800-479-6329 to obtain a Return Authorization (RA) Number.Complete the Verification Section (page 3 of this letter) by checking the appropriate box indicating affected product has been located. Also, please indicate the number of devices found and note the Return Authorization Number. Please include your name, title, address, telephone number and signature in the spaces provided.Send a copy of the completed Verification Section to DePuy Synthes by:Fax: (610) 430-7083 or " Scan/email: Fieldaction@synthes.comNote: If the Verification Section is answered on behalf of more than one facility and/or individual, please clearly indicate the name and address of the facility and/or individual on page 3 of the notification.If you DO NOT have the identified product, please take the following steps:Complete the attached Verification Section (page 3 of this letter) by checking the appropriate box indicating that no affected product has been located. Please include your name, title, address, telephone number and signature in the spaces provided. This return documentation acknowledges your receipt of medical device removal information.Return the documents to DePuy Synthes by Fax: (610) 430-7083 orScan/email: Fieldaction@sythes.comNote: If the Verification Section is answered on behalf of more than one facility and/or individual, please clearly indicate the name and address of the facility and/or individual on page 3 of the notification.If you have any questions, please call 610-719-5450, or contact your DePuy Synthes sales consultant. Thank you for your attention to this issue. |  | 11 | 11 | US Distribution including the states of PA, NC, TN, NY, WI, WA and MI. | Terminated | December 18, 2015 | 442 | 11 | Undetermined
Z-0057-2015 | 69360 | No | HWQ | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Green Mamba Suture Passer. |  |  | September 26, 2014 | 2014 | Biomet, Inc. | An investigation identified that high level friction may exist between the needle and Mamba suture passer instrument, causing the needle to break during use. | Nonconforming Material/Component | On 9/26/2014, "URGENT MEDICAL DEVICE RECALL NOTICE" notifications were sent via courier to the affected distributors with instructions for returning the affected product. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, contact information, and instructions for responding to the formal recall notification. |  | 12 | 12 units | Nationwide Distribution-including the states of FL, NY, TX, VA, NV, MI, CA, IN, SD, GA, KY, NC, and AR. | Terminated | July 14, 2015 | 291 | 191 | Undetermined
Z-0052-2015 | 69409 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | ZOLL Pro-padz Sterile Multi-Function Electrodes (MFE) |  |  | September 16, 2014 | 2014 | Bio-Detek, Inc. | Product sterility is not assured | Material/Component Contamination | Bio-Detek notiifed customers of the expanded recall on September 16, 2014 for Lot 1914 and September 22, 2014 for Lots 1614, 1814, 2114 via UPS overnight delivery. Customers were requested to return the suspect product(s) in their possession. A Device Correction Response Return Form was provided to each Customer. For questions call 1-800-348-9011 or 1-978-421-9460. |  | 237 | 237 cases (6/case) | Worldwide Distribution - US Nationwide and the countries of: Australia, Canada, Belgium, France, and Germany. | Completed |  |  | 489 | Undetermined
Z-0065-2015 | 69065 | No | HSB | Intramedullary fixation rod. | Orthopedic | Orthopedic | 510(k) | Stryker T2 Tibia System Nail Insertion Sleeve, Elastic, Sterile, a) REF 1806-1406S, SPI 08-11 and b) REF 1806-1407S, SPI 08-13. |  |  | July 24, 2014 | 2014 | Stryker Howmedica Osteonics Corp. | Stryker Orthopaedics received a report indicating some of the Nail Insertion Sleeves had a different inner diameter than described on the package label. | Mixed-up of materials/components | Stryker Orthopaedics sent an "Urgent Medical Device Removal" letter dated July 23, 2014 and Business Reply Form to their Branches/Agencies via e-mail on July 23, 2014 and their Hospital Risk Management was sent Notification Letters and Product Accountability Forms via FedEx (with return receipt) by July 24, 2014. The recall letter informed the customer of the problem. The customer was asked to return all affected products available to Stryker Osteosynthesis, c/o Paul Jahnke, Stryker Orthopaedics, 325 Corporate Drive, Mahwah, NJ 07430, REF: PFA #2014-101 or Contact Stryker's customer service and refer to PFA #2014-101 for returning the product to Stryker. Customers were asked to complete and sign the enclosed Business Reply Form and fax a copy to: 1 (855) 251-3635 or email a copy to Aminah Crawford, Recall Coordinator at aminah.crawford@stryker.com. |  | 484 | 484 units | US Nationwide Distribution in the states of CA, FL, IN, MD, NC, NH, NY, OH, PA and WI. | Terminated | June 08, 2017 | 1050 | 484 | Undetermined
Z-0069-2015 | 69179 | No | MWI | Cardiac monitor (including cardiotachometer and rate alarm). | Cardiovascular | Cardiovascular | 510(k) | Philips Expression MR200 MRI Patient Monitoring System |  |  | September 05, 2014 | 2014 | Invivo Corporation | Failure to produce the Non-Invasive Blood Pressure (NIBP) Measurement. | Process control | Philips Medical Systems sent an Urgent - Medical Device Correction letter dated July 9, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Our records indicate that you have an affected device. The following page provides additional instructions and actions that will be taken to address this problem.If you need any further information or support concerning this issue, please contact your local Philips representative: 877-468-4861 option 1, then option 2. |  | 24 | 24 units. | Worldwide Distribution - US including the states of CA, GA, MI, NJ, NV, and TX., and the countries of: Austria, Finland, Germany, Poland, Singapore, Sweden, Switzerland and United Kingdom. | Open, Classified |  |  | 24 | Undetermined
Z-0075-2015 | 69340 | No | HWT | Template for clinical use. | Orthopedic | Orthopedic | 510(K) Exempt | Discovery Elbow Humeral Trial |  |  | September 24, 2014 | 2014 | Biomet, Inc. | An investigation identified that the product was incorrectly manufactured as a left humeral trial but is etched as a right humeral trial. | Process control | Biomet sent an "URGENT MEDICAL DEVICE RECALL NOTICE" letters dated September 24, 2014, via FedEx overnight to affected customers. The notification included a description of the reason for recall, instructions for responding to the formal recall notification, and affected product, actions to be taken. For questions call (574) 372-1570. |  | 6 | 6 | Worldwide Distribution - US Nationwide in the state of SD, OH, CT and the country of: The Netherlands and India. | Terminated | July 14, 2015 | 293 | 6 | Undetermined
Z-0081-2015 | 69228 | No | JIL | Urinary glucose (nonquantitative) test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | External power supply adaptors that accompany Clinitek Status+ |  |  | September 19, 2014 | 2014 | Siemens Healthcare Diagnostics, Inc | The external power supply for the Clintek Status analyzer, which is provided separately, is damaged. It can result in an electric shock to the User. | Process control | Siemens sent an Urgent Field Safety Notice dated September 2014, to all by e-mail to all affected Siemens Healthcare Diagnostics Regional offices on September 19, 2014 both in the United States and elsewhere for communication with affected customers. This notice informs customers of the issue with the power supply and the potentially affected products that may contain a damaged power supply. Consignees are asked to complete a Field Correction Effectiveness Check form and return in to Siemens Healthcare Diagnostics via Fax to (312) 275-7795. All affected domestic customers were sent a hard copy of the Urgent Field Safety Notice via Federal Express on 9/23/14. All affected customers outside the US will be provided a copy of an Urgent Field Safety Notice via hard copy, e-mail, and/or fax as determined by each countries local regulations and procedures.Customers with questions were instructed to contact their local Siemens technical support representative.For questions regarding this recall call 877-229-3711. |  | 4200 | approx. 4200 | Worldwide Distribution - USA (nationwide) including Puerto Rico and Internationally Canada, Mexico, Argentina, Australia, Austria, Bangladesh, Belgium, Brazil, Bulgaria, Canada, Chile, Croatia, Czech Republic, Denmark, Egypt, Estonia, Finland, France, Georgia, Germany, Great Britain, Greece, Hungary, India, Indonesia, Ireland, Israel, Italy, Japan, Kazakhstan, Kuwait, Latvia, Malaysia, Mexico, Morocco, Netherlands, New Zealand, Norway, Pakistan, Poland, Portugal, Qatar, Republic Korea, Romania, Russian Fed., Saudi Arabia, Senegal, Singapore, Slovakia, Slovenia, South Africa, Sweden, Switzerland, Tadjikistan, Taiwan, Turkey, Turkmenistan, Unit.Arab Emir., United Kingdom, Vietnam. | Terminated | July 20, 2015 | 304 | 4200 | Undetermined
Z-0080-2015 | 69253 | Yes | KDJ | Peritoneal dialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | 15 L Cycler Drainage Bag |  |  | October 17, 2013 | 2013 | Baxter Healthcare Corp | The large tube clamp used to close the drain tube on the 15 L Cycler Drainage Bag may not fully close the drain tube and may result in leakage. | Component change control | Baxter sent an Urgent Product Recall letter dated May 30, 2014 to all affected customers. The letter identified the affected product, problem and actions to be taken. The recalling firm is instructing their customers to remove all affected products from their facility and quarantine the product subject to the recall. Customers are instructed to return the affected product to Baxter and receive a credit or replacement. For questions contact the Center for One Baxter at 1-800-422-9837. |  | 3085080 | 3,085,080 | Worldwide Distribution - US Nationwide (including Puerto Rico) and countries of, Canada, Australia, New Zealand, Thailand, Malaysia, Singapore, Korea, Japan, Philippines, Columbia, Ecuador, Austria, Bahran, Belgium, Bosnia &amp; Herzegovina, Bulgaria, Czech Republic, Denmark, Estonia, Finland, France, Georgia, Germany, Great Britain, Greece, Hungary, Ireland, Isreal, Italy, Jordan, Kazakhstan, Latvia, Luxembourg, Malta, Norway, Oman, Poland, Portugal, Qatar, Russia, Saudi Arabia, Slovakia, South Africa, Spain (including Canary Islands), Sweden, Switzerland, The Netherlands, Turkey and United Arad Emirates. | Terminated | January 04, 2016 | 809 | 3085080 | Undetermined
Z-0086-2015 | 69371 | No | MUE | Full-field digital mammography system. | Radiology | Radiology | 510(k) | Siemens Mammomat Inspiration mammography systems. |  |  | September 18, 2014 | 2014 | Siemens Medical Solutions USA, Inc | The stereo biopsy devices for Mammomat Inspiration mammography systems might have integrated a safety switch which causes a failure of the functionality. The pin implemented in the safety switch may not put enough pressure on the safety circuit to prohibit movement. The needle positioning device may move even with the safety switch being set sideward. If this occurs with the needle being already i | Process control | A safety advisory notice, dated September 18, 2014, was sent to end users that identified the product, problem, and actions to be taken. |  | 3 | 3 | US Distribution in states of: NY, VA, and MA. | Terminated | February 24, 2015 | 159 | 3 | Undetermined
Z-0088-2015 | 69291 | Yes | LWP |  |  | Cardiovascular | PMA | Boston Scientific TELIGEN, Implantable Cardiac Defibrillator models E102, E110, the following models are not available in the US: models E103, E111, F102, F103, F110, F111. Sterilized using ethylene oxide. |  |  | September 17, 2014 | 2014 | Boston Scientific CRM Corp | Boston Scientific identified a subset of COGNISTM CRT-Ds and TELIGENTM ICDs that had experienced an increased rate of premature battery depletion due to compromised performance of a low voltage (LV) capacitor. Boston Scientific updated their August 2013 Physician communication to include additional devices that may experience premature battery depletion due to compromised performance of a low volt | Component design/selection | Boston Scientific sent an Important Medical Device Information letter to physicians beginning September 17, 2014. The letter provided a summary of the issue, described clinical implications and expected device performance, and provided patient monitoring recommendations. Boston Scientific is introducing upgraded program software (Model 2868, version 3.04) that enhances the effectiveness of the Safety Architecture tools. Questions should be addressed to Boston Scientific Representative or Technical Services 1-800-227-3422 or tech.services@bsci.com. |  | 31144 | 31,144 devices total | Worldwide Distribution - US Nationwide including American Samoa, Guam, Puerto Rico, and U.S. - Virgin Islands. OUS countries to include: Argentina, Aruba, Australia, Austria, Bahamas, Barbados, Belgium, Bermuda, Brazil, Canada, Cayman Islands, Chile, Colombia, Costa Rica, Cyprus, Czech Republic, Denmark, Dominican Republic, Ecuador, Egypt, El Salvador, Finland, France, French Polynesia, Germany, Greece, Guadeloupe, Haiti, Honduras, Hong Kong, Hungary, India, Indonesia, Iran, Ireland, Israel, Italy, Jamaica, Japan, Jordan, Korea, Kuwait, Latvia, Lebanon, Luxembourg, Malaysia, Mexico, Monaco, Mozambique, Netherlands, Netherlands Antilles, New Caledonia, New Zealand, North Yemen, Norway, Panama, Peru, Philippines, Poland, Portugal, Qatar, Reunion, Russia, Saudi Arabia, Singapore, Slovakia, Slovenia, South Africa, South Korea, Spain, Sri Lanka, Sudan, Sweden, Switzerland, Syrian Arab Republic, Taiwan R O C, Thailand, Trinidad and Tobago, Tunisia, Turkey, United Arab Emirates, United Kingdom, Uruguay, Venezuela. | Terminated | February 20, 2015 | 156 | 31144 | Undetermined
Z-0155-2015 | 69412 | bag, urine collection, leg, for external use, non-sterile | NNW | Urine collector and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | Coloplast Conveen Urine Collection Leg Bag, Product Numbers 5170, 5171, and 5174. |  |  | September 29, 2014 | 2014 | Coloplast Manufacturing US, LLC | Coloplast is recalling the Conveen Contour Leg Bag due to potential leakage from the bag seam when filled. | Process control | Consignees were sent a Coloplast "Recall Voluntary Notification " letter on 9-30-2014. The letter described the problem and the product affected by the recall. Requested consignees to complete the attached Acknowledgement Form and return to Coloplast or contact Coloplast Customer Service at 800-533-0464 to communicate the status of any product in their possession. If the product was distributed, they advised consignees to contact customers immediately and advise them of the recall and have them return the product to them. For questions they can contact Coloplast at Customer Service at 800-533-0464. |  | 217540 | 217,540 | Nationwide Distribution including Puerto Rico and the states of AL, AZ, CA, CO, FL, GA, ID, IL, IN, IA, KS, KY, LA, MD, MA, MI, MN, MS,MO,NE,NH, NJ, NY, NC, ND, OH, OK, OR, PA, RI, SC, TN, TX, UT, VA, WA, and WV. | Terminated | March 10, 2015 | 162 | 217540 | Undetermined
Z-0162-2015 | 69452 | No | KRA | Continuous flush catheter. | Cardiovascular | Cardiovascular | 510(k) | Merit Maestro Microcatheter, 2.4F 130cm Straight, REF28MC24130ST, Rx Only, Non-pyrogenic Sterile if package is unopened or undamaged. |  |  | September 26, 2014 | 2014 | Merit Medical Systems, Inc. | These microcatheters were packaged with a 2.8 F distal tip, rather than the labeled 2.4 F. | Mixed-up of materials/components | Consignees were notified via letter on September 26, 2014. Written notification included product identification information, instruction to immediately quarantine any devices and discontinue use, ensure all personnel to whom devices were distributed are made aware of this field action, and instructions to contact their Merit representative. An email notification was sent to US Merit sales representatives with accounts that had received affected product. The affected sales representatives were provided a Customer Response Form applicable to their account, a representative copy of communications sent directly to their affected accounts, and a copy of the Frequently Asked Questions. Merit sales representatives were instructed to visit the account and assist with the return of the affected units back to the Merit Field Assurance Department. |  | 56 | 56 units | Nationwide and to India. No govt/military/VA consignees. | Terminated | November 19, 2014 | 54 | 56 | Undetermined
Z-0163-2015 | 69462 | No | FFL | Ureteral stone dislodger. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | Bard Dimension Stone Basket |  |  | October 07, 2014 | 2014 | C.R. Bard, Inc. | The product is non-sterile. | Release of Material/Component prior to receiving test results | Bard sent an Urgent: Medical Device Recall letter to each customer on 10/7/2014 with proof of delivery notice. Wholesalers were asked to conduct a sub-recall. Letter identify the affected product, problem and actions to be taken. Customers were instructed to examine their inventory and quarantine product subject to recall, do not use or further distribute any of the affected product, remove and return affected product to Bard Medical Division, and if the product have been further distributed identify your customers and notify them at once of the product recall. Customers were instructed to complete the enclosed Recall &amp; Effectiveness Check Form and fax to 1-770-784-6469. For questions call 1-770-784-6471. |  | 86 | 86 units | Worldwide Distribution in the states of AL, FL, NJ, NY, NC, PA, SC, TX, WI, and countries of Germany, Italy, Kuwait, and Luxembourg. | Terminated | June 12, 2015 | 248 | 86 | Undetermined
Z-0493-2015 | 69683 | No | FRO |  |  | General &amp; Plastic Surgery | 510(k) | GuardIVa Antimicrobial Hemostatic IV Dressing, REF FP-23-AD008, Manufacturer: Bard Access Systems, Inc. USA. Product Code FP23AD006. |  |  | November 03, 2014 | 2014 | Bard Access Systems | Bard Access Systems is recalling GuardIVa (Ref. No. FP23AD006) from lot ASXHT003 because they were non-sterile unit samples which potentially could be used clinically. | Mixed-up of materials/components | A customer letter was sent on 11/3/14 requesting immediate examination of inventory and to quarantine the products for return. |  | 2532 | 2532 units | US Distribution to state of: TX, OK, LA, and AR. | Terminated | May 12, 2015 | 190 | 2532 | Undetermined
Z-0416-2015 | 68780 | walker, mechanical | ITJ | Mechanical walker. | Physical Medicine | Physical Medicine | 510(K) Exempt | MediChoice Dual-Release Folding Walker with 5" wheel |  |  | June 25, 2014 | 2014 | Owens &amp; Minor Distribution, Inc. | Product wheels may be too loose, too tight, or fall off. | Employee error | Customers were notified via letter and email by June 27, 2014. to take the following actions: 1. Designate a quarantine location in the DC for the recalled lots and serial numbers. 2. Label this area very clearly as "RECALLED PRODUCT IN QUARANTINE". 3. Clearly segregate the area to prevent the product from being moved or handled and to further designate that the product is in quarantine. 4. Record each case lot# or walker serial number if the product is out of the case on the recall inventory report.5. Label each unit in the quarantine area with a label that clearly states RECALLED PRODUCT IN QUARANTINE. 6. Conduct a daily inventory verification check using to ensure the inventory still remains in quarantine. In addition, please forward a copy of this letter to hospital personnel in a position to notify patients who may have received and been discharged with a walker from the lots and serial numbers identified above. Hospital Customers: Contact your local DC Customer Service Representative who will assist in returning and replacing the recalled walker. Patients Who May Have Received a Recalled Walker from a Hospital: Contact Owens &amp; Minor Mainstreet Customer Service at 1-800-818-6664 who will assist you in returning and replacing a walker. |  | 4856 | 1214 cases; 4 units/case | US Distribution including the states of CA, CO, CT, GA, IL, MA, MD, ME, MI, NE, NH, NJ, NY, OH, OR, PA, SC, TX, UT, VA, VT, WA and WY. | Terminated | January 06, 2017 | 926 | 4856 | Undetermined
Z-0420-2015 | 69434 | No | LCP | Glycosylated hemoglobin assay. | Hematology | Hematology | 510(k) | Siemens Dimension HB1C Flex reagent cartridge; Dimension Hemoglobin A1c (HB1C) Flex reagent cartridge on the Dimension clinical chemistry system is an in vitro diagnostic assay for the quantitative determination of hemoglobin A1c (HbA1c) in human anticoagulated whole blood. Measurements of hemoglobin A1c are effective in monitoring long-term glucose control in individuals with diabetes mellitus. |  |  | September 30, 2014 | 2014 | Siemens Healthcare Diagnostics, Inc. | Certain lots of the reagent cartridges show confirmed positive bias averaging 0.4% Hemoglobin A1c units and occasionally up to 1.0% HbA1c units for patient samples when compared to the National Glycohemoglobin Standardization Program (NGSP). QC samples may exhibit a similar bias. | Nonconforming Material/Component | Siemens sent an Urgent Medical Device Correction letter dated September 2014 was provided to all Dimension VistaÂż HbA1c and DimensionÂż HB1C customers who received the affected lots to notify them of the issue, the potential risk to health, and actions to be taken by customers. Customers were instructed to discontinue use of and discard the affected devices.Customers need to review the letter with their Medical Director. Siemens will replace any unused inventoryof the affected lots at no charge. Customers were asked to complete the attached form and fax it to 302-631-7795. Customers should retain a copy of the letter with their laboratory records, and forward the letter to those who may have received this product. Customers with questions should contact their Siemens Customer Care Center - Technical Solutions or their local Siemens technical support representative.For questions regarding this recall call 800-441-9250. |  | 53046 | 53, 046 | Worldwide Distribution: US (nationwide) and internationally to: Canada. | Terminated | August 10, 2015 | 314 | 53046 | Undetermined
Z-0418-2015 | 69632 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Zimmer Patient Specific Instrument PERSONA CR Pin Guides (Femur and Tibia) is intended to be used as a surgical instrument to assist in the positioning of Total Knee Replacement components intraoperatively and in guiding the marking of bone before cutting provided that anatomic landmarks necessary for alignment and positioning of the implant are identifiable on patient imaging scans. |  |  | October 16, 2014 | 2014 | Materialise USA LLC | The box of guides for a specific case arrived intact but contained two (2) femoral guides instead of one (1) femoral guide and one (1) tibial guide. The device did not meet specification: One tibia guide missing and one femur guide contained an edge which was +0.04 mm over specification. | Nonconforming Material/Component | Materialise USA was notified on October 16, 2014 by the Field Representative of the distributor (Zimmer) that the box of guides for case JTOMA-JL-0636-L arrived intact but contained two (2) femoral guides instead of one (1) femoral guide and one (1) tibial guide. A new set of tibial and femoral guides were manufactured and sent to Zimmer (headquarters distribution center) on October 17, 2014 before the surgery date on October 20, 2014.For questions regarding this recall call 734-662-5057. |  | 1 | 1 | Nationwide Distribution to MN only | Terminated | February 17, 2015 | 124 | 1 | Undetermined
Z-0421-2015 | 69505 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | GemStar infusion pump, list numbers 13000 and 13100, GemStar Infusion Pump (pain management) list number 13150. |  |  | March 19, 2013 | 2013 | Hospira Inc. | Some patients are obtaining access to the dosage reset codes required to change the settings on GemStar infusion pumps and are subsequently changing the infusion rate set by their doctors. Patients are obtaining access to these codes via website sponsors by downloading the GemStar User/Service Manual featured on the sponsors' websites. The GemStar User/Service Manual should be removed from these w | Other | Hospira sent a letters dated March 19, 2013 to Internet website sponsors to request removal of the GemStar User/Service Manual from their websites as soon as possible. The Internet website sponsors were requested to return a confirmation form after removing the GemStar User Service Manual from their site. For questions call 224-212-2891. |  |  | Not Applicable | Worldwide Distribution - US: Nationwide and the countries of Australia, Austria, Bahrain, Belgium, Croatia, Egypt, Finland, France, Germany, Gibraltar, Greece, Hungary, Icelanofd, Ireland, Israel, Italy, Jordan, Kuwait, Lebanon, Luxembourg, Malta, Netherlands, New Zealand, Oman, Portugal, Qatar, Romania, Saudi Arabia, Slovenia, South Africa, South Korea, Spain, Sweden, Switzerland, Taiwan, Turkey, United Arab Emirates, United Kingdom | Completed |  |  |  | Undetermined
Z-0425-2015 | 69625 | No | LBH | Cavity varnish. | Dental | Dental | 510(k) | Relief ACP Oral Care Gel intended to relieve discomfort from dentin sensitivity. Sold separately in single, single syringe, 4 packs, 6 packs and in teeth whitening kits |  |  | October 24, 2014 | 2014 | Discus Dental LLC | Product sold after April 1, 2014 were inadvertently filled with hydrogen peroxide teeth whitening gel (6% concentration as dispensed) on one side of the dual barrel syringe instead of the intended component (phosphate gel). Consequently, use of this product by patients may result in a neutral effect or increased sensitivity. | Mixed-up of materials/components | Cease distribution of the product immediately and contact patients who may have received the product from April 1, 2014. Patients who are still in possession of Relief ACP Oral Care Gel should return affected product to dentinal office.Discus Dental is offering a non-cost replacement for all affected product, including free shipping. To remedy the situation take the following steps:1. Contact the patients who have received affected product after April 1, 2014 either by letter or telephone, and instruct them to cease use of the product immediately and return any remaining product.2. Please use Attachment A: Product Shipped, Acknowledgement and Response Form to document the quantity of affected lots of product being returned to Discus Dental LLC. Go to www.philips.com/discus-reliefgel and enter the requested information with contact details.3. Please contact Discus Dental customer service at CustomerServicePhone if you have any questions and to arrange for replacement products and prepaid return shipping label. |  | 34904 | 20,251 distributed in U.S. on nationwide level and 14,653 internationally | Worldwide distribution including Nationwide (U.S.) and the countries of New Zealand, Australia, Canada, Germany, Switzerland, Italy, Austria, Finland, South Africa, United Arab Emirates, Korea, Turkey, Egypt, Jordan, Thailand, Lebanon, Kuwait, Tunisia, India, and Dominican Republic. | Terminated | April 14, 2015 | 172 | 34904 | Undetermined
Z-0449-2015 | 69629 | stains, microbiologic | JTS | Dye and chemical solution stains. | Hematology | Pathology | 510(K) Exempt | Remel TB Potassium Permanganate (.5% Aqueous) 40192, packaged in 250 ml bottles, 5/pkg. |  |  | October 29, 2014 | 2014 | Remel Inc | The product may appear cloudy or discolored and may not perform correctly in presumptive diagnosis of mycobacterial disease. | Process control | Thermo Fisher Scientific sent a Medical Device Recall letter dated October 29, 2014, to all affected customers. The letter notified customers that Remel, Inc., a part of Thermo Fisher Scientific, was voluntarily recalling the affected product. Customers were instructed to notify any personnel who need to be notified of the potential for false negative results. Customers should review results reported using this lot and consider retesting and/or seek appropriate expert advice. Return the attached Medical Device Recall Return Response form acknowledging receipt of the notice and disposal of the affected product.Customers with questions were instructed to contact the Technical Services Department at 800-255-6730 (US) or 913-888-0939 (International).For questions regarding this recall call 913-888-0939. |  | 395 | 79/5/250- packages/bottles/ml | Worldwide Distribution to USA including AL, CA, CO, IA, MO, NC, NY, NV, OH, PA, SC, SD, TN, TX, VA, WI and Internationally to Canada. | Terminated | June 01, 2015 | 215 | 395 | Undetermined
Z-0226-2015 | 69603 | No | NKB | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | Synthes Titanium Polyaxial Reduction Head for Titanium Matrix Spine Screws: The Titanium Polyaxial Reduction Head is a part of the Matrix Spine System. |  |  | October 23, 2014 | 2014 | Synthes (USA) Products LLC | Two nonconforming parts of the Titanium Polyaxial Reduction Head for Titanium Matrix Spine Screws may have been inadvertently packaged and shipped. It is believed that the inner diameter of the Reduction Head may not have been machined to the correct specification. | Process control | DePuy Synthes sent an Urgent Notice: Medical Device Recall letter, dated October 23, 2014, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to check inventory and remove affected stock, call Synthes to obtain a Return Authorization for the device, and to return the completed verification section of the response form by fax to (610) 430-7083 or Scan/email: Fieldaction@synthes.com. Customers with questions were instructed to call 610-719-5450. |  | 24 | 24 | Distributed to MI, MT, and CO. | Terminated | August 10, 2015 | 291 | 24 | Undetermined
Z-0490-2015 | 69786 | Yes | KDI | High permeability hemodialysis system. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Polyflux Revaclear Max Dialyzer, Model Number 110634. |  |  | November 05, 2014 | 2014 | Gambro Renal Products Inc | Gambro is initiating a field action for Polyflux Revaclear Dialyzer and Polyflux Revaclear Max Dialyzer due to potential tears in the individual packaging that may compromise the sterile barrier. | Process control | Recall letter was sent via UPS delivery next day to US customers on November 5th, 2014. Product that is returned will be quarantined at the distribution centers. |  | 20000000 | 20 million dialyzers | Worldwide Distribution: US (nationwide) including PR, and countries of: CA, BB, BM, BS, DO and LC. | Terminated | May 20, 2015 | 196 | 20000000 | Undetermined
Z-0468-2015 | 69631 | No | IYB | Diagnostic x-ray tube mount. | Radiology | Radiology | 510(K) Exempt | The J700 Floor Mounted Tube Stand (FMTS) is intended to support and position the diagnostic x-ray tube housing assembly for a medical radiographic procedure. The product consists of a vertical column (with counterbalanced vertical carriage, counterweight and pulley system), horizontal floor car and a floor mounted track to ride on. |  |  | October 28, 2014 | 2014 | Summit Industries Inc. | The welds on J700 tube stands may be insufficient resulting in tube stand breakage, damage to the X-Ray machine, and/or patient or healthcare provider injury. | Process control | Summit Industries sent an Urgent Medical Device Recall letter dated October 27, 2014, to all direct accounts (distributors). The letters included instructions for distributors to: 1) make every effort to locate the affected devices; 2) replace and return the defective parts; 3) contact the end user for any defective device(s) that were installed and determine if the affected device(s) is/are in service; 4) if the device(s) is/are still in service, deliver a copy of the "End User Letter", immobilize the tube arm, and place the REMOVE FROM SERVICE sticker on the device; 5) complete and return the attached questionnaire for a no-charge purchase order for the necessary replacement parts; and, 5) inform the end user that new replacement pats will be installed as soon as possible. Distributors with questions can contact Summit Industries at 800-729-9729 (Monday - Friday, 7:00 AM - 3:45 PM CST). The Urgent Medical Device Recall letter dated October 27, 2014, to end users advised them of the recall and the steps that Summit Industries and their distributor would be taking to correct the issue. |  | 32 | 32 | US Distribution including the states of: AZ, CO, IL, KY, LA, NE, NY, OH, PA, TN, TX, VA, WA and WV. | Terminated | May 21, 2015 | 205 | 32 | Undetermined
Z-0474-2015 | 69289 | No | CDP | Total triiodothyronine test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | ARCHITECT Total T3 Reagent Kit (7K64) consists of: |  |  | September 11, 2014 | 2014 | Abbott Ireland Diagnostics Division | 17% of ARCHITECT Total T3 Reagent Kit, LN 7K64-20/-25 reagent lots 38901UI00 and 38901UI01 may exhibit lower Relative Light Units (RLUs) than expected, which in turn may result in controls out of range or patient results higher than expected. | Employee error | Abbott sent an Product Recall letter dated September 11, 2014, to all direct accounts (customers). The letters included instructions for customers to: 1) Discontinue use of and destroy any remaining inventory according to your laboratory procedures; 2) follow your laboratory protocol regarding the need for review of previously reported patient results; 3) replacement material is available; for ordering, please order through your normal ordering process; and, 4) if you have forwarded any of the affected lots to another laboratory, please provide them with a copy of this communication. Customers or healthcare providers with questions can contact Customer Service at 1-877-422-2688 (24 hours a day, 7 days a week). |  | 7415 | 7,415 kits | Worldwide Distribution - USA (nationwide) including Puerto Rico and the states of AL , AR , AZ , CA , CT , FL , GA , IL , IN , KS , LA , MA , MD , MI , MN , MO , MS , MT , NC , NJ , NV , NY , OH , OK , PA , SC , SD , TN , TX , UT , VA , WA , WI and WV., and the countries of : Andorra , Angola , Argentina , Armenia , Austria , Azerbaijan , Bangladesh , Barbados , Belarus , Belgium , Bosnia &amp; Herzegovina , Brazil , Burkina Faso , Canada , Cayman Islands , Chile , China , Colombia , Costa Rica , Croatia , Czech Republic , Ecuador , Egypt , El Salvador , Ethopia , Georgia , Germany , Greece , Greenland , Guatemala , Honduras , Hong Kong , India , Indonesia , Iran , Ireland , Israel , Italy , Jamaica , Japan , Kazakhstan , Kenya , Korea , Kuwait , Latvia , Lebanon , Libya , Malaysia , Mexico , Moldova , Montenegro , Oman , Pakistan , Gaza &amp; Jericho , Panama , Paraguay , Peru , Philippines , Poland , Portugal , Qatar , Romania , Russia , Saudi Arabia , Serbia , Singapore , Slovakia , Spain , Sri Lanka. , Switzerland , Taiwan , Tajikistan , Thailand , The Netherlands , Trinidad &amp; Tobago , Turkey , Ukraine ,United Arab Emirates , United Kingdom , Venezuela , Vietnam and Yemen. | Open, Classified |  |  | 7415 | Undetermined
Z-0166-2015 | 69408 | No | MEA | Infusion pump. | General Hospital | General Hospital | 510(k) | CADD-Solis Medication Safety Software, Administrator CD, Version 3.1, REF 21-2194-0301-01. |  |  | September 29, 2014 | 2014 | Smiths Medical ASD, Inc. | Smiths Medical has identified an issue with a single batch (Lot Number 2752712) of CADD&quot;-Solis Medication Safety Software Administrator 3.1 CDs. CADD&quot;-Solis Medication Safety Software Administrator 3.1 CDs were sent with the CADD&quot;-Solis Medication Safety Software Point of Care 3.1 software loaded on them. | Process control | Consignees were hand delivered by Smiths Medical Sales representatives on 9/29/2014 a Smiths Medical "Urgent Field Safety Notice" dated 25-September-14. The letter described the problem and the product being recalled. The letter also provided the Advice on Action to be Taken by the User, Transmission of this Urgent Medical Device Recall Notice and requested consignees to complete and return the Confirmation Form by fax to 1-800-237-8033 or via email to recall.response@smiths-medical.com. For questions customers can contact Smiths Medical's Customer Service Department at 1-800-258-5361. |  | 24 | 24 | US Nationwide Distribution in the state of AK, CA, CO, CT, GA, MI, MN, NH, NC, SC, TX, VA, WI. | Terminated | February 25, 2015 | 149 | 24 | Undetermined
Z-0477-2015 | 69735 | No | PAV | Intravascular occluding catheter. | Neurology | Neurology | HDE - Humanitarian Device Exemption | 7F NeuroFlo Catheter, Model 1027 |  |  | April 01, 2014 | 2014 | ZOLL Circulation, Inc. | After aquisition of assets of CoAxia, Zoll discovered that one lot of NeuroFlo Catheters was manufactured by another firm and an HDE supplement listing as the owner of the product was not filed. | Other | Hospital was notified in April 2014 of the recall by UPS. |  | 1 | 1 device | One location in NC | Terminated | November 26, 2014 | 239 | 1 | Undetermined
Z-0206-2015 | 69568 | protector, skin pressure | FMP | Skin pressure protectors. | General Hospital | General Hospital | 510(K) Exempt | Gel-E Donut : 92025-A (Extra Small), 92025-B (Small), |  |  | October 10, 2014 | 2014 | Philips Medical Systems, Inc. | Mold on gel filled Gel-E and Squishon products | Material/Component Contamination | Philips Healthcare/Children's Medical Ventures issued a Medical Device Recall Letter dated Oct 10, 2014. Customers were asked to discontinue use and dispose of all product in their facility, even if mold is not visible. Following these actions, customers will be asked to return a reply form indicating that these actions have been completed. Customers will be given credit for all disposed products. Further, the Gel-E Donut and Squishon 2 products will be withdrawn from the market. Customers with questions were instructed to contact their Philips/Children's Medical Ventures representative.For questions regarding this recall call 978-687-1501. |  | 336695 | 336,695 units | Worldwide Distribution - USA (nationwide) and Internationally to: Canada, Australia, Austria, Belgium, France, Germany, Iceland, Ireland, Italy, Japan, Kuwait, Netherlands, New Zealand, Norway, Portugal, Reunion, Romania, Saudi Arabia, South Africa, Spain, Sweden, Switzerland, Thailand, and the United Kingdom. | Terminated | November 29, 2016 | 781 | 336695 | Undetermined
Z-0467-2015 | 69621 | No | JKR | Calcitonin test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Calcitonin ELISA Kit, Code No. 7024, 7024BU. |  |  | September 18, 2014 | 2014 | Biomerica | A potential issue with false high values. | Material/Component Contamination | An Urgent Field Safety Notice dated 9/18/14 was sent to all their customers who purchased the Calcitonin ELISA Kit. The letter informed the customers of the problems identified and the actions to be taken. Customers were to examine their inventory immediately and quarantine any affected product. Also, customers were asked to notify any customers that the product was further distributed to. Customers with questions were instructed to contact their local Customer Service or Sales Representative regarding the recall notification. Customers were instructed to complete and return the enclosed response form by either fax or email as soon as possible and not later than 10/1/14. |  | 1062 | 1,062 units | Worldwide Distribution -- USA, including the states of NH, CA, NJ, FL; and, the countries of Russia, Ukraine, Switzerland, Spain, Italy, Serbia, China, France, Pakistan, Austria, United Kingdom, Egypt, and India. | Terminated | March 28, 2015 | 191 | 1062 | Undetermined
Z-0482-2015 | 69669 | No | NJE |  |  | Neurology | HDE - Humanitarian Device Exemption | Gateway OTW 2.75mm x 9mm; |  |  | October 22, 2014 | 2014 | Stryker Neurovascular | Units of GDC 360 coils, Target coil and Gateway PTA Balloon Catheter were shipped to US customers beyond the expiration date. | Process control | Stryker Neurovascular sent an Urgent Medical Device Recall Notification dated October 22, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Letters request that users take the following actions:1. Immediately locate the subject devices and quarantine them.2. Distribute this notice to all affected Departments in your facility.3. Inform Stryker Neurovascular if any of the subject devices have been distributed to other organizations.a. Please provide contact details so that Stryker Neurovascular can inform those recipients appropriately.4. Complete and sign the attached Customer Response Form and:a. Fax it to: 1-866-876-4355 orb. E-mail a copy to: NeuroComplaints@stryker.comc. Please complete this form even if you do not have any subject devices to return to prevent the need of follow up notices.5. Keep a copy of the completed, signed Customer Response Form for yourrecords.Stryker Customer Service reps will contact customers to arrange for replacement.Any returned product will be quarantined and destroyed. |  | 1 | 1 unit | US Distribution to the states of: VA, NH, OH and TX. | Terminated | February 12, 2015 | 113 | 1 | Undetermined
Z-0167-2015 | 69345 | No | HQG | Prescription spectacle lens. | Ophthalmic | Ophthalmic | 510(K) Exempt | Prescription eyeglass safety lenses. |  |  | September 12, 2014 | 2014 | Eyemart Express Ltd | Prescription eyeglass safety lenses did not meet specifications. | Employee error | Eyemart Express sent an Important Recall Notice dated September 12, 2014, to all affected customers. The notice identified the product, the problem, and the action to be taken by the customer. Customers were instructed to call Eyemart Express Customer Service at 1-888-DRBARNES to discuss their lens replacement options of if they have any questions. |  | 76 | 76 pairs of eyeglasses | Nationwide Distribution including AK, AL, AR, CA, LA, MD, MO, MS, MT, NC, ND, NM, OH, OK, TN, SC, TX, VA, WI, and WV. | Terminated | October 09, 2015 | 392 | 76 | Undetermined
Z-0186-2015 | 69570 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | Discovery NM/CT 670 dual detector free-geometry integrated nuclear imaging system. Models H3100NA, H3100PP, H3100RA, H2401LS, H2401LZ, H2402LS. |  |  | October 14, 2014 | 2014 | GE Healthcare, LLC | GE is issuing this recall due to a potential safety issue related to a portion of the system falling onto the patient during a scan due to fasteners being loose that secured the camera to the gantry. GE is updating the Preventative Maintenance procedure and schedule, and added an additional Preventative Maintenance check for loose fasteners on the Nuclear Medicaine systems. | Nonconforming Material/Component | The firm, GE Healthcare, sent an "URGENT MEDICAL DEVICE CORRECTION" letter, GEHC Ref# 40860, dated October 14, 2014 to its consignees. The letter described the Safety Issue, safety Instructions, Affected Product Details, Product Correction, actions to be taken and Contact Information. For questions, consignees in the United States, can contact 1-800-437-1171, for other countries they can contact their GE Healthcare Service Representative. |  | 255 | 255 | Worldwide Distribution: US (nationwide) including Guam, Puerto Rico and DC; and countries of: ALGERIA, ARGENTINA, AUSTRALIA, AUSTRIA, BELGIUM, BELRARUS, BRAZIL, BULGARIA, CANADA, CHILE, CHINA, COLOMBIA, CROATIA, CYPRUS, CZECH REPUBLIC, DENMARK, DOMINICAN REPUBLIC, ECUADOR, EGYPT, ESTONIA, FINLAND, FRANCE, GEORGIA, GERMANY, HUNGARY, GREECE, GUADELOUPE, HONG KONG, INDIA, INDONESIA, IRAN, IRELAND, ISRAEL, ITALY, JAPAN, REPUBLIC OF KOREA, KUWAIT, LATVIA, LEBANON, LITHUANIA, MALAYSIA, MALTA,MARTINIQUE, MEXICO, MOROCCO, NAMIBIA, NETHERLANDS, NEW ZEALAND, NORWAY, PAKISTAN, PANAMA, PERU, PHILIPPINES, POLAND, PORTUGAL QATAR, ROMANIA, RUSSIA, SAUDI ARABIA, SINGAPORE, SLOVAKIA, SOUTH AFRICA, SPAIN, SWEDEN,SWITZERLAND, TAIWAN, THAILAND, TUNISIA, TURKEY, UNITED KINGDOM, UKRAINE, URUGUAY, VENEZUELA, VIET NAAM. | Terminated | February 27, 2015 | 136 | 3410 | Undetermined
Z-0466-2015 | 69566 | No | LIB | Microorganism differentiation and identification device. | Microbiology | Microbiology | 510(K) Exempt | Oxoid Legionella Latex Test, DR0800M, containing 50 tests per box. The box contains a variety of components, one of which is Oxoid L. pneumophila Latex Reagent, DR0802, Lot 1301912. The responsible firm on the label is Oxoid Ltd. |  |  | October 23, 2014 | 2014 | Remel Inc | A reagent within the test may return false negative results. | Other | ThermoFisher sent a Medical Device Recall letter dated October 23, 2014. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to notify any personnel who need to be aware of the potential for false negative results for the affected product. Customers should review their inventory for the affected products and discard them. Return the attached Medical Device Recall Return Response acknowledging receipt of the notice and disposal of the affected product. If product was further distributed customers should contact those entities, advise them of the situation and provide them with a copy of this letter. Customers with questions were instructed to contact their Technical Services Department at 800-255-6730 (US) or 913-888-0939 (International). |  | 2950 | 59/50 test boxes | Nationwide Distribution including AZ, CA, DC, FL, IL, ME, NJ, NY, PA, VA, and VT. | Terminated | January 22, 2015 | 91 | 3050 | Undetermined
Z-0492-2015 | 69787 | No | GAP | Natural nonabsorbable silk surgical suture. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Ethicon PERMA-HAND Silk siliconized black braided 10 strands per packet Ethicon LLC- Made in USA. |  |  | October 17, 2014 | 2014 | Ethicon, Inc. | Ethicon is unable to ensure that PERMA-HAND silk suture (siliconized black braid) meets the expected sterility assurance level for sterile products. | Other | Ethicon sent an Urgent Medical Device Recall Letter/Business Reply Form dated 10/17/2014 via UPS next day mail. |  | 2508 | 2,508 each | Nationwide Distribution | Terminated | August 14, 2015 | 301 | 2508 | Undetermined
Z-0419-2015 | 69584 | No | KNH | Contraceptive tubal occlusion device (TOD) and introducer. | Obstetrics/Gynecology | Obstetrics/Gynecology | PMA | Gyrus ACMI Disposable Falope-Ring Band Applicator Kits. REF 006889-901 and REF 005280-901 (with 8 mm Disposable Trocan). Sterile EO. Single Use Only. The disposable Falope-Ring Band Applicator is a sterile, two-band, dual-incision instrument. The two-band applicator enables the physician to occlude both Fallopian tubes without removing the applicator from the abdominal cavity. The Falope-Ring Band and Applicator Kits are indicated for use for female sterilization (permanent contraception). |  |  | October 21, 2014 | 2014 | Gyrus Acmi, Incorporated | All packages of Falope Ring Band Applicator kits are being recalled due to compromises in sterile packaging that could be associated with bacterial contamination that might lead to patient infection. | Process control | Olympus sent an Urgent Medical Device Recall letter dated October 21, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Olympus requires you to take the following action: I. Immediately cease any further use of any affected product you have, remove it from your inventory and quarantine it until it is shipped back to us. 2. Call your Olympus customer service representative at 1-888-524-7266 to obtain a Returned Goods Authorization so that you may return the product with no charge to you. Olympus will issue a credit t or replacement to your facility y for any returned product. 3. Please note on the enclosed questionnaire that you have received this information. 4. Fax the completed questionnaire to 484-896-7128 regardless of whether you have any affected inventory at your facility. ln addition, if you may have further distributed this product, please identify your customers, notify them at once of this product recall, and appropriately document your notification process. Your notification to your customers may be enhanced by including a copy of this recall notification letter. |  | 74792 | 74,792 units | Worldwide distribution. US Nationwide, Canada, Vietnam, and Europe | Completed |  |  | 74792 | Undetermined
Z-0227-2015 | 69514 | No | GEX | Laser surgical instrument for use in general and plastic surgery and in dermatology. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | SideFire Directional Laser Probe, REF: NBP101-01, NBP201-01, NBP301-01, NBP401-01, NBP501-01. These laser probes are part of the NeuroBlate System, a unique collection of MRI-compatible laser devices and accessories that create an MRI guided delivery of precision thermal therapy in the brain. The NeuroBlate System components include Gas-cooled Laser Delivery Probes to deliver controlled energy to a target zone. |  |  | October 03, 2014 | 2014 | Monteris Medical Corp | Monteris Medical received a complaint of a separation of the distal sapphire tip from the laser delivery probe. | Process control | Monteris Medical sent an Urgent Advisory Notice of Product Recall letter dated October 3, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Action to be taken by user:Our distribution records document that you were shipped products from the affected lots. Please check your inventory to confirm whether you have any stock affected by this notice. An acknowledgment form is included to assist you in the assessment of your inventory. Your Monteris Medical Sales Representative will be glad to assist you with this assessment. Once your inventory has been verified, please fax the form back to Monteris Medical at 763-253-4728. |  | 234 | 234 devices | US Distribution including the states of CA, CT, FL, GA, KS, MN, MO, NC, OH, PA, TX, and VA | Terminated | December 08, 2014 | 66 | 367 | Undetermined
Z-0225-2015 | 69693 | No | MAX | Intervertebral body fusion device. | Orthopedic | Orthopedic | 510(k) | The MectaLIF Oblique Handle, model number 03.22.10.0262, is used to insert the MectaLIF oblique spinal cage implants. |  |  | October 21, 2014 | 2014 | Medacta Usa | The MectaLIF Oblique Handle has the laser marking &quot;MEDIAL&quot; on the incorrect side and could result in incorrect insertion of the device during surgery. | Employee error | Medacta sent an URGENT MEDICAL DEVICE RECALL letter dated October 21, 2014 were via e-mail to all sales representatives and distributors. The letters included instructions for the sales represetantives and distributors to: 1) check their inventory for the recalled devices; 2) immedaitely sequester the recalled devices and do not use them during surgery; 3) if the devices have been further distributed, inform them of the recall; and, 4) replacement devices will be provided. Customers with questions can contact Medacta USA at 805-910-6511. |  | 32 | 32 oblique handles | Worldwide Distribution - US including CA, CO, ID and Internationally to Australia and Germany. | Terminated | March 14, 2016 | 510 | 32 | Undetermined
Z-0223-2015 | 69508 | No | LZS |  |  | Ophthalmic | PMA | Nidek Excimer Laser Corneal Surgery System, model EC-5000 CXIII. |  |  | August 06, 2012 | 2012 | Nidek Inc | Contact failure in F1 and F4 fuse holders was detected on the EC-5000 CXIII Excimer Laser Corneal Surgery System that could cause a serious adverse event for the patient. | Component design/selection | Nidek Engineers visited each site to perform correction - fuse holders have been inspected and replaced if necessary.An Engineering Change Order was issued on June 1, 2012, for inspection and/or replacement of affected fuse holders by Nidek Service Engineers.For further questions please call (510) 353-7785. |  | 14 | 14 units | US Distribution including the states of CO, NY, MO, LA, ME, CA, MI, AZ, CT and FL. | Terminated | November 17, 2014 | 833 | 14 | Undetermined
Z-0214-2015 | 69282 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | ADVIA CentaurÂż and ADVIA CentaurÂż XP and ADVIA CentaurÂż CP Cleaning Solution (For in vitro diagnostic use with the ADVIAÂż Centaur systems) |  |  | September 12, 2014 | 2014 | Siemens Healthcare Diagnostics, Inc | Stability of the ADVIA Centaur Cleaning Solution is not assured | Material/Component Contamination | Siemens issued an Urgent Medical Device Recall (UMDR) to customers in the US on September 12, 2014. An Urgent Field Safety Notice (UFSN) was sent out for distribution to all customers outside the United States on September 12, 2014. These notices inform the customer of the degradation of hypochlorite inthese lots of cleaning solution. Customers were instructed to discontinue use of the affected lots and dispose of the product according to local regulations or procedures. The notices should be reviewed with the facility's laboratory or medical director. Fax-back forms will be used to confirm that the customers received the UMDR. For those customers that do not return a fax-back form, a Siemens representative will follow up with those customers to ensure they received the Urgent Medical Device RecallFor questions regarding this recall call 508-668-5000. |  | 4645 | 4645 (12 pk) | Worldwide Distribution - USA (nationwide) and Internationally to Canada. Grerat Britain, Saudi Arabia, Nepal, and Viet Nam. | Terminated | July 17, 2015 | 308 | 4645 | Undetermined
Z-0037-2015 | 68777 | No | KNH | Contraceptive tubal occlusion device (TOD) and introducer. | Obstetrics/Gynecology | Obstetrics/Gynecology | PMA | The Hulka Clip is a sterile packaged medical device that carries a two year expiration date. It is packaged two clips per procedural tray. Theses procedural trays are then sealed in a Tyvek/Mylar pouch and packaged in a cardboard shipping carton with ten pouches in each shipping carton. |  |  | July 09, 2014 | 2014 | Richard Wolf Medical Instruments Corp. | Richard Wolf Medical Instrument Corporation is recalling Hulka Clip Tubal Occluding Device lot 51000232. The devices are being recalled as a part of an investigation into a biocompatibility test discrepancy where the recalled lot (lot 51000232) passed final release testing, but the subsequent lot failed biocompatibility testing on a component that was also used to manufacture lot 51000232. | Nonconforming Material/Component | "Urgent: Medical Device Recall" letters dated July 9, 2014 were sent to all direct accounts (customers). The letters included instructions for customers to: 1) inspect their inventories for the recalled products; 2) contact Richard Wolf Medical Instruments at 800-323-9653 x 225 or 380 to notify Richard Wolf Medical Instruments if you have any devices in inventory; 3) review your records and report the number, if any, of Hulka Clips from lot 51000232 that were implanted in your patients on the Recall Response Form; 4) complete and return the Recall Response Form attached to the letter; and, 5) a RMA number will be provided to customers returning the recalled products and those customers will receive a credit. All returned product should be sent to the attention of Ms. Lisa Williams, Regulatory Specialist. Customers with questions can contact Richard Wolf Medical Instruments at 800-323-9653 x 225 or 380. |  | 73 | 73 cartons (1,460 clips) | Worldwide distribution: US (Nationwide) including states of: (AK, AL, CO, CT, FL, GA, IA, IN, KY, LA, MI, MO, MS, NC, NY, OH, OR, PA, SC, TN, TX, VA, WA, WV; and country of: Germany. | Terminated | April 24, 2017 | 1020 | 73 | Undetermined
Z-0133-2015 | 69395 | No | DXC | Vascular clamp. | Cardiovascular | Cardiovascular | 510(k) | SAM Junctional Tourniquet (SJT) Auxiliary, Model SJT 102 and SJT 112. |  |  | August 14, 2014 | 2014 | Sam Medical Products | SAM Junctional Tourniquet (SJT) Auxiliary strap, model SJT 102 and SJT 112, is recalled because the snap hook on auxiliary strap broke during training demonstration. | Component design/selection | SAM Medical Products issued the press release on August 29, 2014 and the link is http://www.ireachcontent.com/news-releases/sam-medical-products-implements-voluntary-recall-of-its-accessory-axilla-strap-for-the-sam-junctional-tourniquet-273239381.html. SAM Medical Products sent the "URGENT: MEDICAL DEVICE RECALL" letter, dated September 10, 2014, to their consignees.The 5 additional lot numbers (7/1/14, 7/2/14, 8/4/14, 8/7/14, and 8/11/14) that were erroneously placed on the recall notification letter and they are NOT part of the recall.SAM Medical Distributors will need to notify their customers via this letter and the included FAX Back document of the actions being taken as well as ship the replacement straps to individual customers.Distributors will document and report to SAM Medical that all customers have been notified and replacement straps delivered to these customers. Distributors will notify SAM Medical that the replaced straps have been destroyed. This information will be documented on the included FAX Back document and then returned to SAM Medical (via FAX, e-mail, or mail, perinstructions on the FAX Back document).Customers will receive one complete Auxiliary Accessory Strap at no charge with a reinforced metal clip as a replacement. Once received, the existing strap should be removed from the SJT and replaced with the new strap. Customers/End Users need to complete and return the FAX Back document to the distributor from whom the FAX Back was sent.Customer can contact SAM Medical Customer Service at:-Email customerservice@sammedical.com-Phone (503) 639-5474, Monday through Friday, 7:30-4:30, Pacific Time-Toll free l-800-818-4726-Website: vww.sammedical.com |  | 3633 | total 3633 units (3342 units in US and 291 units Outside US) | Worldwide Distribution: US (nationwide) in states of Ohio, Oregon, New Mexico, North Carolina, Texas and to the following countries: Australia, Canada, Hong Kong, Iceland, Japan, Lebanon, Singapore, Thailand, and United Kingdom. | Terminated | July 21, 2015 | 341 | 3633 | Undetermined
Z-0125-2015 | 69326 | No | MIH |  |  | Cardiovascular | PMA | Ovation Prime Abdominal Stent Graft System - 29mm Aortic Body stent grafts. |  |  | September 20, 2014 | 2014 | Trivascular, Inc | TriVascular has identified select lots of Ovation Prime 29mm Aortic Body stent grafts that may be associated with rapid emptying of the fill polymer syringe when it is connected to the Autoinjector during the fill polymer injection step of the primary implant procedure. | Under Investigation by firm | Urgent Field Safety Notice letter was sent to all consignees on 9/19/14 via trackable mail system. Users are notified that TriVascular has identified select lots of Ovation Prime 29mm Aortic Body stent grafts that may be associated with rapid emptying of the fill polymer syringe when it is connected to the Autoinjector during the fill polymer injection step of the primary implant procedure. |  |  | 0 Devices | International Distribution Only -- Canada, Turkey, Austria, Ireland, Belgium, Italy, Switzerland, Spain, France, and Cyprus. | Terminated | December 04, 2014 | 75 | 0 | Undetermined
Z-0172-2015 | 69382 | No | HWE | Surgical instrument motors and accessories/attachments. | General &amp; Plastic Surgery | Orthopedic | 510(K) Exempt | Stryker SmartLife Large Aseptic Housing |  |  | October 08, 2014 | 2014 | Stryker Instruments Div. of Stryker Corporation | The SmartLife Aseptic Housings are being recalled in that the top section of the housing may separate from the bottom section due to insufficient welding. | Process control | On October 8, 2014, Stryker Instruments issued "URGENT MEDICAL DEVICE RECALL NOTIFICATION" to their consignees. Notification included affected product description, reason for recall, risk to health, and actions to be taken by the customer/user. The customer/user were instructed to immediately review this recall notification, check all stock areas; discontinue use of any product which shows sign of cracking or damage at the weld; and complete and return the enclosed Business Reply Form (BRF) via fax to Stryker Instruments Regulatory Department, 866-521-2762 or scan and email a copy to kara.spath@stryker.com. Upon receipt of completed BRF, a replacement product will be shipped and the customer/users, upon receipt of replacement product, are to remove and return all recalled product from use and replace them with new product. For questions regarding this recall, please contact Stryker Instruments: at 269-389-4518 or email: kara.spath@stryker.com. |  | 3138 | 3,138 in total | Worldwide Distribution: US (nationwide) WI, MD, NY, NM, AZ, GA, CA, MN, DE, PA, TX, MT, MI, KS, KY, WY, IL; and countries of: Australia, Canada, Sweden, Netherlands, France Switzerland, India, Japan, Poland, South Africa, England, and Hong Kong. | Terminated | May 31, 2016 | 601 | 3138 | Undetermined
Z-0178-2015 | 69535 | No | DTL | Cardiopulmonary bypass adaptor, stopcock, manifold, or fitting. | Cardiovascular | Cardiovascular | 510(K) Exempt | Cardiopulmonary bypass connector; P/N: 321334-000 DESC: CONNECTOR, 3/8 X 3/8 X 1/2 CLEAR PAY LOT: 295424 STERILE Expiration May 2019 Manufacturing Date May 2014 |  |  | September 18, 2014 | 2014 | Dokitz NovoSci | The single leg of the PAY connector has been observed to become detached. | Process control | The recalling firm notified the affected customers via letter and requested the affected lot be returned to NovoSci. |  | 240 | 240 units | MO, TX, PA, IL, MD, UT | Terminated | May 15, 2015 | 239 | 240 | Undetermined
Z-0179-2015 | 68079 | No | NJT | Immunohistochemistry reagents and kits. | Hematology | Pathology | 510(K) Exempt | Thyroglobulin (1D4) is a mouse anti-human monoclonal antibody produced as a tissue culture supernatant, and supplied in Tris buffered saline with carrier protein, containing 0.35% ProClin 950 as a preservative. Total volume = 7 mL. |  |  | April 18, 2014 | 2014 | Leica Microsystems, Inc. | Bond Ready to Use Primary Antibody Thyroglobulin does not function as intended up to the expiry date on the product labelling. This was detected through an on-going stability program. There is a link between the age of the product and staining intensity. | Unknown/Undetermined by firm | Leica Biosystems sent a Medical Device Recall Notification letter dated April 18, 2014, to all affected customers. The letter included instructions for customers to: 1) assess any results previously obtained by using the recalled products; 2) identify and destroy any unused or partly used recalled products or confirm that the recalled products are no longer in stock; and, 3) complete and return the attached Medical Device Recall Notification Acknowledgement Form. Customers with questions were instructed to contact Leica Microsystems by e-mail at LMGRA@leica-microsystems.com.For questions regarding this recall call 847-405-5413. |  | 180 | 180 units | Worldwide Distribution - USA including AR, AZ, CA, CO, CT, FL, GA, IL, KY, LA, MA, MD, MI, MN, MO, MT, MY, NC, NH, NJ, NY, OH, ON, OR, PA, TN, TX, UT, VA, WA, WI and the District of Columbia, Internationally to Belgium, Bulgaria, France, Italy, Pakistan, Philippines, Russia, Saudi Arabia, Spain, Turkey, and the United Kingdom. | Terminated | August 14, 2017 | 1214 | 180 | Undetermined
Z-0175-2015 | 69437 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Access 2 Immunoassay Systems, Catalog No. 81600N, 386220, Instructions for Use Part No. B14253A. |  |  | October 06, 2014 | 2014 | Beckman Coulter Inc. | Beckman Coulter is recalling the Small Form Factor Console PC (SFF PC) (B23083) for use with the Access 2 Immunoassay systems because it may experience a "MFC Exception" error during normal operation of the Access 2 Immunoassay Systems. | Vendor change control | An Urgent Medical Device Recall Letter dated 10/6/14 was sent to all customers who purchased the Small Form Factor Console PC for use with the Access 2 Immunoassay Systems Family. The letter informs the customers of the problems identified and the actions to be taken. Customers are instructed to complete and return the enclosed Response Form within 10 days. Customers with questions are instructed to contact Customer Support Center at http://www.beckmancoulter.com/customersupport/support or call (800) 854-3633 in the US and Canada. Outside the US and Canada, customers are instructed to contact their local Beckman Coulter representative. |  | 1519 | 1,519 units total (403 units total) | Worldwide Distribution - US Nationwide in the countries of Algeria, Argentina, Australia, Austria, Bangladesh, Belarus, Bosnia-Herzegovina, Botswana, Brunei Darussalam, Bulgaria, Canada, China, Colombia, Costa Rica, Croatia, Czech Republic, Ecuador, Egypt, El Salvador, Ethiopia, France, Germany, Gibraltar, Greece, Guam, Hong Kong, Hungary, India, Indonesia, Ireland, Israel, Italy, Japan, Kazakhstan, Kenya, Korea, Lebanon, Libyan Arab Jamahiriya, Malaysia, Mexico, Morocco, Namibia, Pakistan, Palestinian Territory, Panama, Peru, Philippines, Poland, Portugal, Romania, Russian Federation, Saudi Arabia, Serbia, Slovakia, South Africa, Spain, Sri Lanka, Sweden, Switzerland, Taiwan, Thailand, Tunisia, Turkey, United Kingdom, United States including Puerto Rico, Venezuela, and Vietnam. | Terminated | June 29, 2015 | 266 | 1519 | Undetermined
Z-0192-2015 | 69474 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | OECÂż UroView 2800. |  |  | October 03, 2014 | 2014 | GE OEC Medical Systems, Inc | GE Healthcare initiated a field correction because prior revision of components were potentially installed on certain OEC 9800, FlexiView 8800, OEC MiniView 6800 and OEC UroView 2800 systems. | Component design/selection | On October 3. 2014. GE Healthcare Surgery began distributing an Urgent Medical Device Correction letter tocustomers to inform them of the issue and provide instructions for immediate mitigation of the issue. The letterwas sent to 3 titles at each facility affected: 1.1 Facility Administrator. 2.1 Director/Manager of Radiology andRadiologists, 3.) Radiology Department. Firm will track allservice requests for the field action to completion. These service requests will document the inspection and component replacement (if applicable) on the affected units. Exceptions to completion of the service request will be recorded. Records of field action completion will be retained per our processes and procedures. No product is being returned. |  | 2 | 2 | Worldwide Distribution. | Terminated | March 23, 2015 | 171 | 344 | Undetermined
Z-0170-2015 | 69540 | No | LON | Fully automated short-term incubation cycle antimicrobial susceptibility system. | Microbiology | Microbiology | 510(k) | Rapid Neg BP Combo Panel Type 3; Catalog number B 1017-117. |  |  | October 17, 2014 | 2014 | Siemens Healthcare Diagnostics, Inc. | An increase in false positive susceptible results on Rapid Neg BP Combo Panel Type 3 and Rapid Neg Urine Combo Panel Type 1 used with MicroScan Microbiology Systems. | Nonconforming Material/Component | Customer Recall notification letters were sent on October 17, 2014. Letters recommend that any remaining inventory be discarded and that consideration be given to reviewing previous test result, conduct patient follow up and/or repeat testing if isolates are still available.A second letter clarifying the recall was issued November 7, 2014 which provided additional information as requested by CDRH. |  | 1381 | 1381 | Distributed US (nationwide) including the states of IA, CA, NY, WI, MI, SD, NC, FL, HI, OH, WA, TX, WA, IN, NM, MS, KY, IL, UT, ND, and WY and the country of Canada. | Terminated | June 29, 2015 | 255 | 3620 | Undetermined
Z-0198-2015 | 69397 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Persona Stemmed 5 Degree Cemented Tibia |  |  | October 09, 2014 | 2014 | Zimmer, Inc. | Cleaning process validation failure. | Nonconforming Material/Component | Zimmer sent an Urgent Medical Device Recall letter dated October 9, 2014 to affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to: 1. Review the notification and ensure affected personnel are aware of the contents. 2. Assist your Zimmer sales representative with the quarantine of any affected product. 3. Your Zimmer sales representative will remove the recalled product from your facility. 4. If after reviewing this notification you have further questions or concerns please call the customer call center at 1-800-348-2759. Hours of operation are Monday through Friday, 8 a.m. through 8 p.m. EST. |  | 138 | 138 units | US Nationwide Distribution and countries of Austria ,Belgium, Switzerland, Germany, France, United Kingdom, Italy and South Africa. | Terminated | June 24, 2015 | 258 | 138 | Undetermined
Z-0197-2015 | 69596 | No | KWS | Shoulder joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Integra Titan Reverse Shoulder System right and left head cutting templates, 142.5 degrees. |  |  | October 24, 2014 | 2014 | Integra LifeSciences Corp. | A single lot of left and a single lot of right Reverse Shoulder System cutting templates were manufactured incorrectly. Specifically, the threaded handle has been welded backwards on the template resulting in the cutting angle being the reverse of what it should be. | Nonconforming Material/Component | The recalling firm will notify their consignees via letter delivered by traceable courier service. The consignee will be asked to complete and return a form to the recalling firm stating if they have affected product or not. |  | 12 | 12 units | Nationwide Distribution including the states of TX, CA, OH, MS, VA, IL, TN, GA, CA, FL, and AZ. | Terminated | July 21, 2015 | 270 | 12 | Undetermined
Z-0200-2015 | 69521 | No | JSK | Supplement for culture media. | Microbiology | Microbiology | 510(K) Exempt | Oxoid Legionella BCYE Growth Supplement SR0110A, packaged in 100 ml vials, 10/pkg. |  |  | October 17, 2014 | 2014 | Remel Inc | Recalled product may not provide adequate recovery of Legionella pneumophila when used as directed and may result in false negative identification of Legionella pneumophila. | Under Investigation by firm | Thermo Fisher sent an Urgent Medical Device Recall letter dated October 17, 2014, to all affected customers. |  | 160 | 16/10/100 ml, vial, pkgs | Distributed in the states of ND, VT and TX. | Open, Classified |  |  | 160 | Undetermined
Z-0199-2015 | 69271 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | ATTUNE Intuition Distal Femoral Jig; sold non-sterile; product is wrapped in bubble wrap and then sealed inside a polyethylene bag |  |  | October 15, 2014 | 2014 | DePuy Orthopaedics, Inc. | Issuing a device correction because if the pin bushing is over loaded or inverted, there is the potential for partial or full displacement and fracture of the plastic around the pin bushing. | Process control | DePuy sent an "URGENT DEVICE CORRECTION" notifications dated October 15, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. For questions, please call 574-371-4917. |  | 3082 | 3,082 | Worldwide Distribution - USA (nationwide) except AK, including Puerto Rico and the countries of Canada, Australia, Austria, Belgium, China, France, Germany, India, Ireland, Israel, Italy, Japan, Malaysia, Netherlands, New Zealand, Singapore, South Africa, South Korea, Spain, Switzerland, Thailand, UAE, United Kingdom and Chile. | Terminated | November 19, 2015 | 400 | 3082 | Undetermined
Z-0202-2015 | 69493 | No | HSB | Intramedullary fixation rod. | Orthopedic | Orthopedic | 510(k) | GMK Intramedullary Extension Rod 100 mm is a rod to obtain correct leg alignment during total knee replacement. |  |  | September 22, 2014 | 2014 | Medacta Usa | Medacta has registered in its database 3 cases of breakage of the GMK Intramedullary ExtensionRod 100 mm, code 02.07.10.0162. | Nonconforming Material/Component | Medacta sent an URGENT MEDICAL DEVICE letter dated September 22, 2014 to the sole US consignee. The letter included instructions for the consignee to: 1) check its inventory of the device; 2) quarantine and do not use the recalled device; and, 3) return the recalled device upon receipt of a RMA # from the Medacta QA Department. For questions regarding this recall call 805-910-6511. |  | 1 | 1 rod (US); 145 rods (global) | Worldwide Distribution - US including OH and Internationally to Australia, Austria, Belgium, France, Germany, Greece, Italy, Japan, Portugal, South Africa, Spain, and Switzerland. | Terminated | May 07, 2015 | 227 | 1 | Undetermined
Z-0210-2015 | 69668 | No | FMT | Infant radiant warmer. | General Hospital | General Hospital | 510(k) | Fisher and Paykel Healthcare (FPH) IW900-Series Infant Warmer |  |  | November 06, 2014 | 2014 | Fisher &amp; Paykel Healthcare, Ltd. | Breakage of the nut that secures the heater head in place, which can cause the heater head to become partially detached and swing towards the bassinet. There is the potential that a heater head could contact a patient situated on the bassinet and cause a serious injury. | Manufacturing material removal | Fisher &amp; Paykel Healthcare sent an Urgent Medical Device Recall letter to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to:Step 1: Identify any affected Infant Warmers in your inventory by checking the REF and SERIAL number on the device label (see examples in Figures 1 and 2).Step 2: Identify any affected spare parts in your inventory by checking the REF and LOT number on the parts (see example in Figure 3).Step 3: Check your Infant Warmer service records to determine if affected spare parts were installed on any other Infant Warmers. If you are unable to identify specific Infant Warmers that may have had affected spare parts installed, then please contact FPH at the contact details below.Step 4: Replacement of Affected Nut. - Remove the affected Infant Warmer from use.- Switch the warmer off and disconnect it from the mains power supply.- Support the heater head during disassembly to prevent it from falling.- Follow Steps A to F in Figure 4 below to replace the nut.- Ensure the pivot washer remains in place when installing the new nut.- Ensure that the heater head is secure, can rotate smoothly, and the detent locates the head in the central position.- Reconnect the mains power supply, switch the warmer on, and ensure there are no error codes or alarms before returning to use.Step 5: Complete the Urgent Medical Device Recall Response Form and return it to your FPH Representative |  | 661 | Nationwide (U.S.): 661 Infant Warmers and 180 Spare Parts, Internationally: 2,451 Infant Warmers and 280 Spare Parts | Worldwide Distribution - USA (nationwide) and Internationally to Albania, Algeria, Australia, Austria, Bahrain, Bangladesh, Bekgium, Bermuda, Chile, China, Croatia, Dominican Republic, Estonia, Finland, France, Germany, Great Britain, Greece, Hong Kong, Iceland, India, Indonesia, Iran, Ireland, Israel, Jordan, Kenya, Kuwait, Latvia, Lebanon, Lithuania, Malaysia, Mauritius, Mexico. Morocco, New Zealand, Nigeria, Northern Ireland, Norway, Pakistan, Palestinian Territory, Panama, Peru, Philippines, Poland, Portugal, Qatar, Romania, Russia, Saudi Arabia, Serbia, singapore, Slovenia, South Africa, Spain, Sri Lanka, Sultanate of Oman, Syria, Taiwan, Thailand, The Netherlands, Tunisia, Turkey, United Arab Emirates, Uruguay, Venezuela, Vietnam, and West Bank. | Terminated | April 27, 2015 | 172 | 661 | Undetermined
Z-0213-2015 | 69504 | No | DQA | Oximeter. | Cardiovascular | Anesthesiology | 510(k) | Masimo SET uSpO2 Cable |  |  | September 30, 2014 | 2014 | Masimo Corporation | Firm has identified a small number of Oximetry Cables with crossed-internal wires. The effect on performance when these wires are crossed is that when a low SpO2 value is measured, a high SpO2 value would be displayed on the monitoring device. Additionally when a high SpO2 value is measured, a low SpO2 value would be displayed on the monitoring device. | Employee error | An "Urgent: Advisory Notice," dated 09/30/2014, was sent to customers. The letter informed customers of the problem identified and the actions to be taken. Customers were instructed to evaluate their inventory of Oximetry Cable(s) using the following method: Use the Oximetry Cable to take a SpO2 measurement on a healthy individual with an expected SpO2 value above 90% or review SpO2 measurements take with the Oximetry Cable and verify that measurement results were consistent with expected SpO2 values. If the SpO2 measurement(s) is consistent with expected SpO2 value, the Oximetry Cable is not affected and can continue to be used. Customers are instructed to complete and return Attachment 2, Tracking/Verification Form, to Masimo in order to document that the Oximetry Cable(s) is functioning as intented. If the assessment results in SpO2 measurement that is not consistent with expected SpO2 value, then the Oximetry Cable should be removed from service, segregated and returned to Masimo. Customers are instructed to call 1-800-326-4890 and select option 2 for Technical Services. Customers are to determine with Technical Services whether they will return the affected Oximetry Cable(s) to Masimo for replacement or refund. If returning for replacement customers are suppose to obtain a Return Material Authorization Number (RMA number). |  | 21754 | 21,754 units (21,222) in the U.S. and Puerto Rico) | Worldwide distribution including; nationwide (U.S. and Puerto Rico) and the countries of Argentina, Austria, Brazil, Canada, China, Denmark, France, Gernany, Hong Kong, India, Italy, Japan, Jordan, Kuwait, Mexico, Netherlands, Poland, Portugal, South Korea, Spain, Switzerland, Turkey, and United Kingdom. | Terminated | March 28, 2016 | 545 | 21754 | Undetermined
Z-0209-2015 | 69536 | No | JCW | Penile inflatable implant. | Gastroenterology/Urology | Gastroenterology/Urology | PMA | AMS 700Âż Series Inflatable Penile Prosthesis Accessory Kit (AMS 700 Line Accessory Kit), 72401850. |  |  | October 10, 2014 | 2014 | American Medical Systems, Inc. | AMS' supplier has issued a recall of certain lots of Lightning Bolt Keith needles supplied in AMS 700 Accessory Kits. Based on this, AMS is recalling AMS 700 Accessory Kits which contain these lots of needles.The supplier made the decision to issue this recall after an internal investigation that was requested by AMS. The investigation identified a potential nonconformity that may lead to a bent | Nonconforming Material/Component | Consignees were sent via either Fed Ex mailing or hand delivery by AMS Field Personnel beginning on 10/10/2014 an AMS "Urgent Medical Device Recall" letter dated October 9, 2014. The letter described the problem and the product involved in the recall. Requested consignees to complete and return the Product Recall Acknowledgement Form. For questions consignees can contact AMS Customer Service at 1-800-328-3881 , FAX 1-800-366-9035, ore mail salesoperations@ammd.com. |  | 184 | 184 (110 US, 74 OUS) | Worldwide Distribution US (nationwide) including PR and the states of AL. AZ. AR, CA, CO, CT, IL, IN, KS, KY, MD, MA,MI, MN, MO, NY,NC, OH,OR, TN, TX, VA, and WA, and the countries of ARGENTINA, AUSTRALIA, BRAZIL, CANDA, GERMANY, FRANCE, GREAT BRITAIN, LEBANON, PANAMA, SOLOMON ISLANDS, SWEDEN, and COLOMBIA. | Terminated | September 28, 2015 | 353 | 184 | Undetermined
Z-0205-2015 | 69488 | No | DQY | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | Medtronic Launcher Coronary Guide Catheter; MODEL: LA6JR40K |  |  | October 13, 2014 | 2014 | Medtronic Vascular | Two reports have been received on the 6F launcher Guide Catheter lot (Model Number LA6JR40K; lot No 0007244836} for incorrect usable length of the Guide Catheter. It was reported that the catheters had a usable length of 90 cm instead of 55 cm, which was specified on the packaging label. Usable length is defined as the total length of catheter shaft from the distal end of the strain relief to the | Process control | Beginning October 13, 2014, Medtronic field representatives will being hand delivery of the customer communication to affected customers, notifying them of the issue, and requesting return of unaffected product. Customers will be asked to sign a certificate confirming they have received the notification and have taken the requested action.1. Identify and quarantine all listed, unused product in your inventory.2. Return all listed product in your inventory to Medtronic. Contact Customer Service at 888-283-7868 to initiate a product return and credit. Your Medtronic sales representative can assist you in the return of this product as necessary.3. Complete the attached Customer Confirmation Certificate and email it to RS.CFQFCA@Medtronic.com or fax it to Medtronic at 651-367-0612 to the attention of Customer Focused Quality. |  | 11 | 11 devices | US Distribution including the states of NC, IA, IL and KS. | Terminated | February 18, 2015 | 128 | 11 | Undetermined
Z-0219-2015 | 69685 | No | HSB | Intramedullary fixation rod. | Orthopedic | Orthopedic | 510(k) | 10 mm/135 degree Titanium Cannulated Trochanteric Fixation Nail 170 MM - Sterile: intended to treat stable and unstable fractures of the proximal femur |  |  | October 29, 2014 | 2014 | Synthes, Inc. | The 10 mm/135 degree Titanium Cannulated Trochanteric Fixation Nail 170 MM (Sterile) was assembled incorrectly. If the non conformance goes undetected, it may impact the locking function post-operatively and compromise the bone reduction and construct stability which may lead to non-union/malunion. | Employee error | An urgent medical device recall notice, dated October 29, 2014, was sent to end users that identified the product, problem and action to be taken. Customers were instructed to return the device to Synthes along with completing the verification section of the notice. |  | 6 | 6 | US Distribution in the states of: AK, FL, OH, OK, and TX. | Terminated | December 18, 2015 | 415 | 6 | Undetermined
Z-0001-2015 | 69192 | No | MWI | Cardiac monitor (including cardiotachometer and rate alarm). | Cardiovascular | Cardiovascular | 510(k) | Philips Expression MR200 MRI Patient Monitoring System Model 866120 |  |  | September 08, 2014 | 2014 | Invivo Corporation | During production, excess metal (aluminum) debris was found stuck to the cling plastic that covers the device's pole assembly. The metallic debris could affect system performance due to the potential for a failure to power up or loss of power. This could result in the loss of monitoring and a delay in diagnosis of the patient. | Material/Component Contamination | The firm, Philips Healthcare, sent an "URGENT- Medical Device Recall" letter dated June 17, 2014, to its customers informing them of the product, problem and actions taken. The firm informed the customers that there are no additional actions required on their part. The Philips/Invivo Account Manager will contact them in order to arrange for replacement of the device. If you need any further information or support concerning this issue, contact your local Philips/Invivo representative: 877-468-4861 option 1, then option 2. |  | 10 | 10 units | Worldwide Distribution: US distribution in state of: CA and countries of: Poland, Singapore, and Sweden. | Terminated | January 09, 2017 | 854 | 10 | Undetermined
Z-2731-2014 | 69213 | No | NAY | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | ASSEMBLY, PATIENT SIDE CART, PS4000 used with the da Vinci Xi Surgical System. |  |  | September 11, 2014 | 2014 | Intuitive Surgical, Inc. | Correction due to the detection of a motor sensor failure in the Patient Cart instrument arms of the da Vinci Xi (IS4000) Surgical Systems after a mid-procedure restart. | Other | Intuitive sent an Urgent Medical Device Correction letter dated September 11, 2014 to all consignees via trackable mail system. Letters identified the issue and provides information for users should a error code be displayed. For questions contact your Clinical Sales Representative or Intuitive Surgical Customer Service |  | 75 | 75 devices | Worldwide Distribution - US nationwide and the countries of: Belgium, Germany, China(Hong Kong), India, Israel, Italy, and the United Kingdom. | Terminated | October 23, 2014 | 42 | 75 | Undetermined
Z-2736-2014 | 68014 | No | PEN | Multiplex nucleic acid assay for identification of microorganisms and resistance markers from positive blood cultures. | Microbiology | Microbiology | 510(k) | Verigene BC-GN Kit (Catalog number 20-005-021); 1) 20 Verigene BC-GN Test Cartridges; are qualitative multiplexed in vitro diagnostic tests for simultaneous detection and identification of selected gram-positive (BC-GP) and gram-negative (BC-GN) bacteria and resistance markers. |  |  | June 17, 2013 | 2013 | Nanosphere, Inc. | There is a specific Extraction Tray lot containing Tips that may slightly increase the occurrence of Process SP Processing Error when used due to the Processor SP experiencing difficulty releasing the Tip to the final location in the Tip Assembly. The difficult releasing the Tip can result in a Processing Error and occasionally the Tip to be dropped into the Processor SP. | Process change control | Undated recall notification letters were sent to all direct accounts (customers) June 17 - 19, 2013. The letters included instructions for customers to: 1) check their inventory for the recalled products; 2) set any recalled products aside for return; and, 3) contact Nanosphere to obtain replacement tip holders. Customers with questions about this recall can contact Nanosphere via e-mail at technicalsupport@nanosphere.us or by calling 888-837-4436. |  | 20 | 20 extraction trays | Worldwide Distribution:US (Nationwide) including states of: AZ, CA, FL, IL, KY, MD, MN, MT, OH, PA, TN, TX, VA, WA, WI; and countries of: Belgium, China, and Saudi Arabia. | Terminated | March 14, 2016 | 1001 | 1680 | Undetermined
Z-2265-2014 | 68738 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | Synthes 3.0 MM TI Cannulated Screw Instrument and Implant Set |  |  | June 30, 2014 | 2014 | Synthes, Inc. | A screw set was returned from the field for destruction after being exposed to flood water, which was erroneously reprocessed and returned to the field. Initial investigative testing indicated the components within the set may be cytotoxic. | Employee error | An Urgent Field Safety Notification, dated July 21, 2014, was sent to surgeons explaining the product, problem, and action to be taken with regard to routine post-operative follow-up on patients exposed to the affected product. Questions or concerns should be directed to Leonie Rynn at 610-719-5408 or lrynn1@its.jnj.com. |  | 4 | 4 | Distributed in the states of WI and MI. | Terminated | December 18, 2015 | 536 | 4 | Undetermined
Z-2238-2014 | 68848 | No | JMG | Ferritin immunological test system. | Immunology | Clinical Chemistry | 510(k) | Siemens Dimension Chemistry Wash (RD701) is an accessory of the flex cartridge system used on the Siemens Dimension RxL and Xpand assays. The wash buffer is used to clean the probe sensor between uses and prepares it for taking the next reading/measurement. |  |  | July 15, 2014 | 2014 | Siemens Healthcare Diagnostics, Inc. | Siemens Healthcare Diagnostics has received complaints for the Dimension Chemistry Wash (RD701) of QC and patient sample result accuracy shifts on Troponin I, CTNI (RF421C), LTNI (RF521C); Free T4, FT4 (RF410); Thyroid Stimulating Hormone, TSH (RF412); and N-terminal pro-brain natriuretic peptide, PBNP (RF423A) and LPBN (RF523A) when using Dimension&reg; Chemistry Wash lot RD41041.Internal inv | Material/Component Contamination | Siemens sent an Urgent Medical Device Recall letter dated July 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Actions to be taken by the Customer:1. Immediately discontinue use and discard all inventory of Chemistry Wash lotRD41041. If this lot was never in use on the instrument, no further action is needed. If this lot has been in use, continue with steps 2 through 5.2. Place an alternate lot of Chemistry Wash on the Dimension system.3. Prime the new bottle of Chemistry Wash 30 times.4. Recalibrate the CTNI, LTNI, TSH, FT4, PBNP and/or LPBN assays using the alternate lot of Chemistry Wash on board.5. For the listed assays, process Quality Control every 8 hours.Actions to be taken by Siemens:Your Customer Service Engineer will be making a visit to decontaminate the system if you have used lot RD41041. This process will take approximately two hours. Once decontamination has been completed, process and evaluate quality control. Resume routine QC testing frequency according to your standard laboratory practice.Siemens is shipping replacement product at no charge for all orders filled with lot RD41041. If you have questions about your replacement shipment, please contact Customer Service at (888) 588-3916.Please complete the attached form and fax it to (302) 631-8467 to indicate that you have received this information.Please retain this letter with your laboratory records. Please forward this letter to anyone to whom you may have provided this product.We apologize for the inconvenience this situation has caused. If you have any questions, please contact your Siemens Customer Care Center or your local Siemens technical support representative. |  | 3200 | 3200 | USA (nationwide) Distribution and Internationally to Canada. | Terminated | August 26, 2015 | 407 | 3200 | Undetermined
Z-2242-2014 | 68931 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Vanguard 360 Femoral Boss Reamer, |  |  | July 01, 2014 | 2014 | Biomet, Inc. | The Vanguard 360 Femoral Boss Reamers were manufactured oversized to the print. The instrument may not fit into the mating component. | Process change control | Urgent Medical Device Recall Notices dated July 15, 2014 were mailed with verification of delivery via FedEx.Response forms were provided to all consignees in order forBiomet to track returns. The letters informed the consignees to return the devices to Biomet and conduct a sub-recall sending the recall letter to all customers. |  | 461 | 461 | California, IA, FL, MD, UT, SC, AL, IL, CO, MA, KS, GA, WA, NV, AR, NH, OK,Pennsylvania, New York, Texas, Virginia, New Jersey,North Carolina, Kentucky, Michigan, Ohio, Wisconsin, Louisiana,Arizona, Indiana, South Dakota, Missouri &amp; Montana. In addition to Korea, B.V., Japan, Puerto Rico, UK, Chile, New Zealand, Australia, &amp; Canada | Terminated | October 30, 2015 | 486 | 461 | Undetermined
Z-2241-2014 | 68527 | No | LZO | Hip joint metal/ceramic/polymer semi-constrained cemented or nonporous uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Apex Arc Hip Stem, Size 2, HA Coated; Product Code: HC-21002-HA |  |  | May 09, 2014 | 2014 | Omnilife Science Inc. | Staining on hip implant. | Manufacturing material removal | All product was consigned to US sales Agents until implanted or returned to OMNI. All US Agents holding the inventory were notified on May9, 2014 by phone and follow up email that the product is being recalled. Customers have been instructed to immediately discontinue use and return the product to OMNIlife science. |  | 13 | 13 devices | US: CA, OK, VA, FL, MA, NY, TX | Terminated | August 19, 2014 | 102 | 13 | Undetermined
Z-2258-2014 | 68792 | No | JDS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | TRIGEN(R) INTERTAN(TM) 10MM X 18 CM NAIL, 125o, REF 71675201, QTY (1), STERILE R, Smith &amp; Nephew, Memphis, Inc., TN 38116 |  |  | July 02, 2014 | 2014 | Smith &amp; Nephew, Inc. | There was an error in manufacturing which resulted in the proximal lag screw hole being positioned incorrectly in several batches of TRIGEN(R); INTERTAN(TM); NAILS. | Under Investigation by firm | Smith &amp; Nephew Issued an Urgent Product Recall dated July 2, 2014 to all affected customers via E-mail and/or FedEx . The notification identified the effective product, problem and actions to be taken. Customers were instructed to: Inspect inventory and locate any unused affected devices, cease distribution and use of those products, quarantine them immediately, complete the last two columns in the Inventory Return Certification Form, and return any affected product to the address provided. For questions call 901-396-2121 |  | 601 | 601 units | Worldwide Distribution - US Nationwide and the countries of AUSTRALIA, CANADA, CHINA, INDIA, UAE, JAPAN, SOUTH AFRICA, SOUTH KOREA, UNITED KINGDOM. AUSTRIA, NORWAY, POLAND, PORTUGAL | Open, Classified |  |  | 601 | Undetermined
Z-2263-2014 | 68870 | No | GEX | Laser surgical instrument for use in general and plastic surgery and in dermatology. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | WaterLase iPlus Dental Laser System, Part No. 7200854-XX |  |  | June 12, 2014 | 2014 | Biolase Technology Inc | Biolase is recalling the WaterLase iPlus Dental Laser System because of a discrepancy present in the user interface software (version 1.7). If an error occurs in the procedure screen, the air setting will change to 1% and OFF. | Under Investigation by firm | An Urgent Medical Device Correction letter dated 6/12/2014 was sent to all customers who purchased the WaterLase iPlus Dental Laser System. The letter informs the customers of the problems identified and the actions to be taken. Customers are informed that a Field Service Engineer can install the software update. Customers with are instructed to contact Customer Service at (800) 321-6717. |  | 117 | 117 units | Nationwide in US and worldwide to: Canada, UK, Russia, Korea, Hong Kong, Libya, Iran, Iraq, Kuwait, Singapore, Armenia, Egypt, Pakistan, Sweden, New Zealand, Poland, Germany, Italy, UAE. | Terminated | October 06, 2014 | 116 | 117 | Undetermined
Z-2259-2014 | 68935 | No | MHY |  |  | Neurology | PMA | Medtronic, Activa PC, Model 37601, Method of Sterilization: Ethylene Oxide, Single Use Only, Rx Only. |  |  | July 14, 2014 | 2014 | Medtronic Neuromodulation | Medtronic is recalling six Activa PC (model 37601) Implantable Neurostimulators due to the potential for a damaged electrical component during manufacturing. | Process control | The firm, Medtronic, notified their Consignees on 07/14/2014 via telephone of the recall. Medtronic representative used telephone script to convey the information. The script was directed to Risk Management or Inventory Management. The caller was to inform consignees of the problem and product being recalled. Advised consignees to quarantine the product and provided the Medtronic Device Removal Reply Form to the consignees via e-mail or Fax. The Reply Form included contact information which was to call 1-800-633-8766 in case they needed to contact a Medtronic representative. The completed form is to be faxed back to 1-800-897-3899 or e-mail a PDF to neuro.quality@medtronic.com.If you have any questions, call 763-526-1294. |  | 6 | 6 | Distributed in the states of: MA, NC, OH, and TX. | Terminated | October 06, 2014 | 84 | 6 | Undetermined
Z-2260-2014 | 68980 | No | KRA | Continuous flush catheter. | Cardiovascular | Cardiovascular | 510(k) | Uni*Fuse INFUSION SYSTEM WITH COOPER WIRE, 5F x 90 cm x 10 cm Infusion Pattern, Catalog No./REF 12401806, Product No. H787124018065, STERILE -- AngioDynamics, 603 Queensbury Avenue, Queensbury, NY 12804 |  |  | August 06, 2014 | 2014 | Angiodynamics, Inc. | AngioDynamics became aware that Uni*Fuse Infusion Systems with Cooper Wire have the potential to contain an incorrect sized guidewire (too long). | Component design/selection | AngioDynamics sent an URGENT MEDICAL DEVICE RECALL Letter dated August 6, 2014, to all affected customers via Federal Express. Consignees were instructed to segregate and return all affected devices to AngioDynamics,Inc. and complete and return the Reply Verification Tracking Form, provided in the recall notification. Those customers with the affected product were also instructed to call AngioDynamics Customer Service at 1-800-772-6446.For questions regarding this recall call 718-742-4571. |  | 97 | Domestic: 70 units; Foreign: 27 units | Worldwide Distribution - USA (nationwide) and Internationally to Austria, Australia, Ireland, Malaysia, Sweden, and Switzerland. | Terminated | March 27, 2017 | 964 | 227 | Undetermined
Z-2268-2014 | 68735 | No | DRT | Cardiac monitor (including cardiotachometer and rate alarm). | Cardiovascular | Cardiovascular | 510(k) | Blood pressure (BP) circuit of the ANSAR ANX-3.0 Autonomic Monitor system. |  |  | July 01, 2014 | 2014 | Ansar Group, Inc. Dba Ansar Medical Technologies, Inc | The blood pressure (BP) circuit occasionally fails to inflate the blood pressure cuff and the BP circuit fails to activate, resulting in no BP measured. | Unknown/Undetermined by firm | Recall notification letters sent to customers on July 8-9, 2014 by registered mail. |  | 351 | 351 | Nationwide Distribution. | Terminated | August 10, 2015 | 405 | 351 | Undetermined
Z-2270-2014 | 68956 | No | MWI | Cardiac monitor (including cardiotachometer and rate alarm). | Cardiovascular | Cardiovascular | 510(k) | Philips IP5 Information Portal |  |  | June 30, 2014 | 2014 | Invivo Corporation | The wireless function of the IP5 radio could be interrupted causing loss of wireless communication between the IP5 and host monitoring system. | Process design | The firm, Philips Healthcare, sent an "URGENT - Field Safety Notice" dated June 30, 2014 to its customers. The notice describes the product, problem and actions to be taken. The notice also provides the customers with instructions for use while waiting for the update to the IP5. A Philips representative will contact you regarding your affected device.If you need any further information or support concerning this issue, please contact your local Philips representative: 877-468-4861 option 1, then option 2. |  | 573 | 573 devices. | Worldwide distribution: US (nationwide) including states of: AL, AR, AZ, CA, CO, CT, DE, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, NC, NE, NH, NJ, MN, NV, NY, OH, OK, OR, PA, RI, SC, TN, TX, UT, VA, VT, WA, WI, WV, Hawaii, Alaska, and District of Columbia; and countries of: Australia, Austria, Denmark, France, Germany, Ireland, Netherlands, Norway, Poland, Saudi Arabia, Singapore, Slovenia, Spain, Sweden, Switzerland, United Arab Emirates, and United Kingdom. | Open, Classified |  |  | 573 | Undetermined
Z-2295-2014 | 68973 | No | GAX | Nonpneumatic tourniquet. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Port Access Kit with MicroClave Clear Connector, Item No. B79104 |  |  | August 05, 2014 | 2014 | ICU Medical, Inc. | ICU Medical is recalling the BD PosiFlush Sterile Field Saline Flush Syringe because some unit packages may exhibit open seals which impacts package sterility and potentially product sterility. | Other | An Urgent recall letter dated 8/5/2014 was sent to all customers who purchased the BD PosiFlush SF Saline Flush Syringe 10ml. The letter informs the customers of the problems identified and the actions to be taken. Customers with questions are instructed to contact ICU Customer Service Monday through Friday between 8:30AM to 4:00PM PST at (866) 829-9025 and select option 8, or email customerservice@icumed.com. Customers are instructed to complete the attached Medical Device Recall Response Form and fax to (801) 264-1755 or email to recall@icumed.com. |  | 400 | 400 units | Nationwide Distribution - USA including AL and MD. | Terminated | February 26, 2015 | 205 | 1000 | Undetermined
Z-2562-2014 | 68650 | No | KNC | Obstetric table and accessories. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(K) Exempt | AffinityÂż Four Birthing Bed, |  |  | June 26, 2014 | 2014 | Hill-Rom, Inc. | The attachment latch mechanism can become bent depending on user handling. Latch mechanism bending could lead to improper engagement of the foot section and this could result in injury to the user due to a fall. | Nonconforming Material/Component | Hill-rom issued an Urgent Field Safety Notice letter dated June 26, 2014 to affected customers. The Notification identified the affected product issue and instructions for product inspection. The letter states that the firm is working on making changes to the Lift-Off foot section that will mitigate risk even if damage was to occur. If you have any questions concerning the notice or requested procedure, please contact Hill-Rom Technical Support at 800-445-3720. |  | 13000 | 13,000 units | Worldwide Distribution - US Nationwide and the following countries: Argentina , Aruba, Australia, Austria, Bahamas, Bahrain, Belgium, Belize, Bermuda, Brazil, Brunei, Bulgaria, Cayman Islands, Chile, China, Colombia, Costa Rica, Croatia, Cuba, Cyprus, Czech Republic, Denmark, Dominican Republic, Egypt, Equatorial Guinea, Finland, France, Germany, Greece, Guam, Guyana, Hong Kong, Hungary, India, Indonesia, Iraq, Ireland, Israel, Italy, Japan, Jordan, Kenya, Kuwait, Lebanon, Libya, Malaysia, Mauritius, Mexico, Morocco, Netherlands, New Zealand, Nigeria, Norway, Oman, Pakistan, Palestine, Philippines, Poland, Portugal, Puerto Rico, Qatar, Romania, Russian Federation, Saudi Arabia, Serbia, Singapore, Slovenia, South Africa, South Korea, Spain, Sweden, Switzerland, Syria, Taiwan, Thailand, Trinidad and Tobago, Tunisia, Turkey, Turks and Caicos Islands, United Arab Emirates, United Kingdom, US Virgin Islands, Venezuela and Vietnam. | Terminated | July 14, 2017 | 1114 | 13000 | Undetermined
Z-2192-2014 | 68443 | protector, skin pressure | FMP | Skin pressure protectors. | General Hospital | General Hospital | 510(K) Exempt | Gel-E Donut, Squishon 2 gel pillows |  |  | May 28, 2014 | 2014 | Childrens Medical Ventures | The gel-filled Gel-E Donut and Squishon product line has received a number of complaints about visible mold. The mold detected was determined to be Cladosporium and Penicillium Fungi, which are commonly found molds. | Material/Component Contamination | Childrens Medical Ventures sent an Urgent Medical Device Recall letter dated May 28, 2014, to affected customers via UPS. The letter identified the affected product, problem, and actions to be taken. Customers were instructed to review entire inventory and inspect for signs of visible mold, do not open the packaging to do the inspection, the mold should be seen through the transparent outer packing, complete and return the reply form provided on the last page of the letter, dispose of any product containing mold per your facilitys environmental guidelines, do not return any product to your Distributor or to Philips/Childrens Medical Ventures.and credit for disposed product will be given. If you need any further information or support concerning this problem, please contact your Distributor at the following phone number: |  | 332649 | 332,649 | Worldwide Distribution - USA (nationwide) including the country of Canada. | Terminated | November 29, 2016 | 916 | 332649 | Undetermined
Z-2291-2014 | 68966 | No | LZE | Nonpowered goniometer. | Orthopedic | Orthopedic | 510(K) Exempt | GE Healthcare, Discovery XR656P/N 5374989/5397837 is intended to generate digital radiographic images of the skull, spinal column, chest, abdomen, extremities and other body parts in patients of all ages. |  |  | July 25, 2014 | 2014 | GE Healthcare, LLC | GE Healthcare has become aware of a potential safety issue involving the nylon hooks which support the FlashPad" detector on the Discovery XR656 wall stand systems.There have been reported incidents of detector falls, but no injuries have been reported. | Nonconforming Material/Component | Consignees were sent a GE Healthcare "Urgent Medical Device Correction" letter dated July 25, 2014. The letter was addressed to Hospital Administrators / Risk Mangers, Managers of Radiology / Cardiology and Radiologists / Cardiologists. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction, and Contact Information. For questions in the USA call 800-437-1171. For other countries contact your local GE Healthcare Representative. |  | 334 | 334 (162 US, 172 OUS). | USA: AL, CA, CO, FL, GA, HI, IL, IN, KS, LA, MD,MI, MN, MO, MT, NJ, NM, NY, NC, OH, OR, PA, SC, TN, TX, UT, VT, VA, WA, WV, WI.OUS: ARGENTINA, AUSTRALIA, AUSTRIA, BELGIUM, BERMUDA, CANADA, CHILE CHINA, DENMARK, FINLAND, FRANCE, GERMANY, HUNGARY, INDIA, ITALY, JAPAN, REPUBLIC OF KOREA, KUWAIT, MOROCCO, NEW ZEALAND, NORWAY, PAKISTAN, RUSSIA, SAUDI ARABIA, SLOVENIA, SPAIN, SWEDEN,THAILAND, UNITED ARAB EMIRATES, UNITED KINGDOM, URUGUAY. | Terminated | February 20, 2015 | 210 | 334 | Undetermined
Z-2542-2014 | 68730 | No | JCF | Lymphocyte separation medium. | Hematology | Hematology | 510(K) Exempt | BD Vacutainer CPT Cell Preparation Tube with Sodium Heparin, 8 mL, 16 x 125 mm, Catalog #REF 362753, Lot #4063409 with expiry 03/2015, intended for the collection of whole blood and the separation of mononuclear cells. |  |  | July 03, 2014 | 2014 | Becton Dickinson &amp; Company | BD has initiated the recall of its Vacutainer CPT Cell Preparation Tube due to poor separation, resulting in a reduced number of recovered mononuclear cells. | Under Investigation by firm | BD issued an "Urgent Voluntary Product Recall" Letter dated July 3, 2014 and a "Notice of Return" Form to all affected customers. |  | 296160 | 296,160 tubes | US and Foreign | Open, Classified |  |  | 296160 | Undetermined
Z-2554-2014 | 68961 | No | KMI | Bed-patient monitor. | General Hospital | General Hospital | 510(K) Exempt | CURBELL MEDICAL CareSense Wireless Advanced LCD Bed and Chair Monitor, MODEL # CSM-BC400; --- COMMON/USUAL NAME: Bed Patient Monitor; CLASSIFICATION NAME: Bed Patient Monitor --- Monitor is connected to a sensor pad that monitors the patient if they attempt to arise from a bed, chair, floor mat, or toilet seat. The monitor will sound an alarm. |  |  | August 05, 2014 | 2014 | Curbell Medical, Inc. | The firm became aware of a potential problem that was initiated by a customer complaint. After consultation with the manufacturer, it was discovered that a resistor was incorrectly placed within the circuit board on the monitor. This change to the resistor was a planned change to address a product improvement (improve battery drain). However, this modification rendered the monitor's alarm via t | Component design/selection | Urgent Medical Device Recall Letters (dated 8/01/14) and Response Forms were sent to the consignees via e-mail (w/read receipt) on 8/05/14. |  | 323 | Domestic: 112 units; Foreign: 211 units | Nationwide, Canada, Ireland, Italy | Open, Classified |  |  | 323 | Undetermined
Z-2545-2014 | 68934 | No | KRR | Knee joint patellofemoral polymer/metal semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Zimmer Gender Solution Patello-Femoral Joint (PFJ) System Patello-Femoral Trochlea Component, Precoat, Size 4, Left. Knee prosthesis. |  |  | July 30, 2014 | 2014 | Zimmer, Inc. | Internal review found that five (5) PFJ implant forgings were identified as have forged laps during receiving inspection process for raw material. These five forgings were not segregated and were subsequently issued to final manufacturing under Part 00-5926-014-01; lot 62455332. A forged lap is a defect that forms when metal folds over itself during forging. | Nonconforming Material/Component | Zimmer sent an URGENT MEDICAL DEVICE RECALL letter dated July 2014 to affected distributors and hospital accounts. The letter identified the affected product, problem and the actions to be taken. The letter instructed customers to 1. Review the notification and ensure affected personnel are aware of the contents, 2. Locate and quarantine any affected product listed above and notify your Zimmer sales representative, and 3. Your Zimmer sales representative will remove the recalled product from your facility. If after reviewing this notification you have further questions or concerns please call the customer call center at 1-877-946-2761 between 8:00 am and 5:00pm EST. |  | 15 | 15 units | Worldwide Distribution - US (nationwide) in the states of HI, CA, GA, MI, and IL and the countries of TAIWAN, BELGIUM, AUSTRALIA, FRANCE, JAPAN, and UK. | Terminated | January 20, 2015 | 174 | 15 | Undetermined
Z-2555-2014 | 68986 | No | OJR | Powered suction pump. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | Enforcement Discretion | The Thermedx Fluid Management System is designed to provide irrigation fluids to irrigate, distend, and debride while warming the fluid to assist in the prevention of peri-operative hypothermia. The fluid is warmed via an internal infrared mechanism. Additionally the system will hold suction canisters for the collection of used irrigant. |  |  | February 10, 2014 | 2014 | Thermedx LLC | A vendor changed the manufacturing process of a component used in the canister ring that may affect the performance of the ring. The deficit display may be higher than the actual deficit when the fluid collection canister is fully loaded. | Component design/selection | Thermedx sent an Urgent Product Correction letter dated March 15, 2014, to all affected consignees. The letter identified the product, the problem, and the action to be taken by the consignee. Consignees were notified that the devices were in need of correction, and that it would be performed at their facilities by Thermedx. Customers with questions were instructed to call a Thermedx technical service representative at 1-888-542-9276.For questions regarding this recall call 440-542-0883. |  | 21 | 21 | Nationwide Distribution including HI, MS, NY, OH, PA, VA, WA, WV, and IL. | Terminated | August 04, 2015 | 540 | 21 | Undetermined
Z-2559-2014 | 68710 | No | NTI | Drug metabolizing enzyme genotyping system. | Toxicology | Toxicology | 510(k) | Verigene CYP2C19 Nucleic Acid Test (CYP2C19 Test). Each CYP2C19 kit consists of 20 Verigene CYP2C19 Nucleic Acid Test Cartridges and 20 Verigene CYP2C19 Nucleic Acid Extraction Trays (with Tip Holder Assemblies). Clinical laboratory test. |  |  | March 18, 2014 | 2014 | Nanosphere, Inc. | Nanosphere has recieved reports related to an unacceptable increase in initial no-call rate when performing the Verigene CYP2C19 Nucleic Acid Test (2C19), which has been confirmed through internal testing. | Under Investigation by firm | Nanosphere sent an Field Notification Letter dated March 18, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.The letters included instructions for customers to: 1) discontinue use of the affected products, 2) discard the affected product, and 3) complete and return the attached for to obtain free-of-charge replacement kits. Customers with any questions about this recall should contact their Nanosphere technical support representative at 1 888 837-4436.. |  | 78 | 78 kits | Worldwide Distribution - US including the states of AL, CA, MD, ND, NJ, OK and TX., and internationally to Bulgaria, China and New Zealand. | Terminated | July 02, 2015 | 471 | 78 | Undetermined
Z-2536-2014 | 68739 | Yes | BTR | Tracheal tube. | Anesthesiology | Anesthesiology | 510(k) | 4001PS Neonatal/Paed.Intubation Kit 2.5mm Murphy, LDS 15mm Conn |  |  | July 07, 2014 | 2014 | Smiths Medical ASD, Inc. | The tube fastener on the ET Tube Holder may compress the tracheal tube when used which may cause difficulty when locking the ET Tube Holder in place or difficulty when passing asuction catheter. | Process change control | Smiths Medical sent an Urgent Medical Device Recall Notice dated July 11, 2014, to all affected customers. The letter informed customers of the recall and instructed them to inspect their inventory and segregate any unused affected products and to complete and return the Recall Confirmation Form, by Fax 1-800-621-2590 or email to endotrachealtubeholders@smiths-medical.com.Customers with questions were instructed to call Smiths Medical's Customer Service Department at 1-800-258-5361.For questions regarading this recall call 866-216-8806. |  | 135518 | 135,518 (110,246 nationwide) | Worldwide Distribution - USA including AK, AR, AZ, CA, CO, CT, DC, FL, GA, IA, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, MT, NC, NE, NH, NJ, NM, NY, OH, OK, OR, PA, SC, SD, TN, TX, UT, VA, WA, WI, WV and Internationally to CA, CY, GB, IE, IT, JE, KW, MT, PL, RU, SA, ZA. | Open, Classified |  |  | 135518 | Undetermined
Z-2557-2014 | 68834 | No | LOX | Percutaneous Transluminal Coronary Angioplasty (PTCA) Catheter. | Cardiovascular | Cardiovascular | 510(k) | Cordis EMPIRA RX PTCA Dilatation Catheter |  |  | June 23, 2014 | 2014 | Cordis Corporation | The balloon may exhibit resistance to the removal of the balloon cover. Weakening of the affected product in the region of the proximal balloon seal may occur when excessive force is applied to remove the balloon cover. This may lead to balloon burst, shaft separation or deflation difficulties, which could result in procedural delay, vessel trauma or surgery. | Environmental control | Cordis Corp sent Urgent Medical Device Correction Letters dated June 23, 2014, to all affected customers. Those in the Foreign Countries were sent via E-mail dated June 23, 2014. The letter identified the product the problem and the action needed to be taken by the customer.Cordis and Creganna-Tactx are not removing the products, but rather providing additional labeling (customer letter) advising the user not to use the product if unusual resistance is felt during removal of the balloon cover. Cordis and Creganna-Tactx explained in the letter to the consignees that is is known that balloons with hydrophilic coating are susceptible to expansion of the coating at extremely elevated humidity, therefore, it is important for the User to follow the current labeling, to store the product in a cool, dark and dry place.Action requested on your part:"Pass on this letter to anyone who needs to be informed in your facility and in any facility where potentially affected devices may have been transferred."Maintain awareness of this communication until the information has been incorporated into the Cordis EMPIRA" NC RX and EMPIRA" RX PTCA dilatation catheter labeling."Retain this letter with affected product."Report all balloon cover removal difficulties through the standard complaint process."For questions related to the issue please contact Cordis Endovascular Medical information Line 877-DEVICE5 (877-338-4235), Monday through Friday from 9:00AM to 5:00 PM EDT."For questions related to the Acknowledgment Form and its return, please contact Cordis QA at (786) 313-8730. |  | 250000 | 250,000 units. | Worldwide Distribution - USA including Puerto Rico and the states of AR, AZ, CA, CO, CT, DC, FL, GA, IL, IN, KY, LA, MA, MD, MI, MN, MO, MS, NC, ND, NE, NH, NJ, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, WA, WV, WY and Hawaii., and the countries of Argentina, Brazil, Canada, Chile, Colombia, Mexico, Panama and Venezuela. | Open, Classified |  |  | 250000 | Undetermined
Z-2232-2014 | 68747 | No | DAP | Fibrinogen/fibrin degradation products assay. | Hematology | Hematology | 510(k) | VIDAS D-Dimer Exclusion II, an automated quantitative test for use on instruments of the VIDAS family for the immunoenzymatic determination of fibrin degradation products in human plasma using the Enzyme Linked Fluorescent Assay technique. Also indicated for use in conjunction with clinical pretest probability assessment model to exclude DVT and PE disease. |  |  | April 17, 2013 | 2013 | BioMerieux SA | Customer complaints about not repeatable results below the detection limit (&lt;45 ng/mL). Therefore false negatives are being reported for results around the medical decision point of 500 ng/mL that initially tested at &lt;45 ng/mL. | Under Investigation by firm | A recall letter was issued on 4/17/2013 to US customers via FED Ex mailing. The letter included a fax back acknowledgment form. A recall notice was also sent via email to bioMerieux Subsidiaries and international distributors on 4/11/2013. |  | 2160 | 2160 units | AL, AR, CA, CO, CT, DE, FL, GA, ID, IL, IN, IA, KS, KY, LA, ME, MA, MI, MN, MO, MT, NE, NH, NJ, NY, NC, OH, OK, OR, PA, RI, SC, TN, TX, VT, VA, WA, WV, WI, WY | Terminated | October 08, 2014 | 539 | 2160 | Undetermined
Z-2236-2014 | 68831 | No | PEN | Multiplex nucleic acid assay for identification of microorganisms and resistance markers from positive blood cultures. | Microbiology | Microbiology | 510(k) | Verigene BC-GN Test Kit (Catalog number 20-005-021) includes: 1) 20 Verigene BC-GN Test Cartridges. Each Test Cartridge comes preloaded with all required reaction solutions, including wash solutions, oligonucleotide probe solution and signal amplification solutions required to generate a test result. BC-GN is indicated for use in conjunction with other clinical and laboratory findings to aid in the diagnosis of bacterial bloodstream infections; however, is not used to monitor these infections. |  |  | July 09, 2014 | 2014 | Nanosphere, Inc. | Nanosphere Inc. has recently determined through four customer complaints that a small number of Verigene Gram-Negative Blood Culture Nucleic Acid Test (BC-GN) Extraction Trays in Lot 042914021B have low levels or are missing Magnetic Microparticles (MMPs) required for DNA Extraction. | Under Investigation by firm | Nanosphere sent an Urgent Product Recall letter dated July 9, 2014 to all affected customers. All customers were initially notified of the recall by telephone on July 9, 2014. The customers were instructed to quarantine and return all affected product to Nanosphere. The letter included instructions: 1) on how to return the recalled products, and 2) to review any test results that were generated while using the recalled products to ensure that there has been no negative impact on patient care. Customers with questions were instructed to contact Nanosphere at 888-837-4436, option 2 or via e-mail at Technicalsupport@nanosphere.us. |  | 1045 | 1045 extraction trays (840 affected by recall) | Nationwide Distribution including CA, CO, IA, MT, OH, OK, OR, TX, and UT | Terminated | September 01, 2015 | 419 | 1045 | Undetermined
Z-2155-2014 | 68806 | No | MCW | Intraluminal artery stripper. | Cardiovascular | Cardiovascular | 510(k) | CSI Cardiovascular Systems, Inc., Diamondback 360 Peripheral Orbital Atherectomy System, Model Number DBP-125MICRO145, Part Number 7-10003. |  |  | May 23, 2014 | 2014 | Cardiovascular Systems, Inc. | CSI has initiated a recall on the Diamondback 360 Peripheral Orbital Atherectomy Device because it may contain a saline sheath that may experience cracking, fracture, and release particulate during use. | Process control | Cardiovascular System, Inc. sent consignees an Urgent Medical Device Recall letter dated May 23, 2014. The letter described the Affected Product, Recall Description, Instructions which included to remove the affected product and return it to CSI and to complete and return the Customer Acknowledgement Form. For further Information they customers were instructed to contact Customer Service, Cardiovascular System, Inc., 877-274-0901.For questions regarding this recall call 877-274-0901. |  | 48 | 48 | Nationwide Distribution including AZ, AR, CA, CT, FL, IL, IA, MD, MI, NY, NC, PA, TN,and TX. | Terminated | October 01, 2014 | 131 | 48 | Undetermined
Z-2212-2014 | 68651 | No | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Device Brand Name: Distal Radius Fracture Repair System |  |  | June 23, 2014 | 2014 | Biomet, Inc. | The firm received complaints indicating that the drill guide becomes stuck in the jig. The drill guide can be placed in the jig about halfway before it becomes stuck. | Nonconforming Material/Component | On June 23, 2014 Biomet issued recall notification to their consignees via letter. Notification included affected product issue and instructions for return were provided. Distributors who had affected hospitals in their regions were notified and provided with a copy of the notice that was sent to the hospital. Consignees are being contacted on an individual basis in an attempt to reconcile the product 100%. A response form was provided to all consignees in order for Biomet to track returns. Non-responders will be contacted on an individual basis. |  | 37 | 37 | The product was distributed domestically and to internationally. Domestically to CA, PA, IA, CO, NY, MD, NC,AL, AZ, IL, GA, VA.Netherlands, Argentina, | Terminated | July 14, 2015 | 386 | 37 | Undetermined
Z-2138-2014 | 68774 | No | NXX | Staphylococcus aureus serological reagents. | Microbiology | Microbiology | 510(K) Exempt | AdvanDx Quick FISH Fixation Kit for BC ( Blood Cultures), 50 tests. In vitro diagnostic. The QuickFISH Fixation Kits are components that are sold separately for the Staphylococcus QuickFISH BC test |  |  | July 04, 2014 | 2014 | Advandx Inc | Positive result in the Internal negative control well for the Staphylococcus QuickFISH BC test, failing QC | Nonconforming Material/Component | AdvanDx notified customers by e-mail on July 4, 2014 of the problem and advised to discontinue use of the QuickFISH Fixation Kits with Lot#04914. Replacement kits were immediately shipped to all of the affected customers. Questions contact local AdvanDx representative or Technical Support 1-866-376-0009, fax 1-781-376-0111 or e-mail: techsupport@advandx.com.An acknowledgement of Receipt and Destruction of Materials requested to be completed. |  | 10 | 10 kits | Distributed in the states of AZ, ID, FL, MO, OH, MI, NJ, and VA. | Terminated | September 30, 2014 | 88 | 10 | Undetermined
Z-2136-2014 | 68863 | No | FZX | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Serrated Soft Tissue Shields. Orthopedic surgical instruments, part of the Silex Sacroiliac Fusion System. The soft tissue shields are surgical instrument guides in the Silex Sacroiliac Joint Fusion System for spinal surgery. |  |  | July 11, 2014 | 2014 | X Spine Systems Inc | Revision B of the tissue shields of the device was 7 mm longer than the previous design. | Employee error | X-Spine systems contacted both customers via telephone and instructed to return the affected product back to X-Spine.All product has been recovered. The distributors have already returned the recalled product. No product was utilized in any surgical procedure.Immediate Action Taken:X-spine inventory was reviewed and two pieces of each size were recovered from finished inventory. The X-spine sales person returned the product. The distributor has returned the remaining product. The devices were recovered prior to use in any surgical procedure.For further questions please call (937) 847-8410 |  | 5 | A total of 5-pieces of each part number distributed | US Distribution including the states of FL and MO. | Terminated | August 14, 2014 | 34 | 5 | Undetermined
Z-2137-2014 | 68406 | No | KPR | Stationary x-ray system. | Radiology | Radiology | 510(k) | KODAK DirectView DR 9000 System, Catalog Numbers 1966688 (US and Canada) and 8161937 (outside the US and Canada) --- A permanently installed diagnostic x-ray system composed of 3 main components: an operator console, U-Arm assembly with x-ray tube including a collimator, and a stationary generator. |  |  | May 30, 2014 | 2014 | Carestream Health Inc. | A control system failure may cause the U-Arm to move downward unexpectedly when the Radiology Technologist is positioning the equipment using the "Up" or "Down" control buttons. Downward movement occurs at three times the normal rate of speed and will continue until the device control button is released, the Emergency Stop button is activated or the equipment Bump Sensor contacts an object. | Nonconforming Material/Component | Urgent Medical Device Recall Letters (dated 5/29/2014) were sent to the consignees on 5/30/2014. Carestream Health recommends that the DR 9000 U-arm be moved before positioning the patient under the U-arm for a decubitus exam. This will remove any potential safety hazard posed to the patient as a result of any unexpected movement. For questions or concerns, please contact the Carestream Customer Care Center in the US at 1-800-328-2910, available 7 days per week on a 24 hour basis; Outside of the US, please call your local Service support number. |  | 146 | Domestic: 100 units; Foreign: 46 units | Worldwide Distribution-USA (nationwide) and the countries of Canada, China, Italy, Germany, Greece, Kuwait, Spain, South Africa, and United Arab Emirates. | Terminated | December 17, 2014 | 201 | 146 | Undetermined
Z-2144-2014 | 68862 | No | DZE | Endosseous dental implant. | Dental | Dental | 510(k) | Tapered-Screw Vent Implant, HA, 4.7mmD x 8mmL x 4.5mm, Part No. TSVWH8 Dental implant component. |  |  | July 17, 2014 | 2014 | Zimmer, Inc. | Zimmer Dental is recalling the Tapered Screw-Vent Implant because the inner vial state "03.7 x 8mm" instead of "04.7 x 8 mm." | Process design | A recall letter dated 7/22/14 was sent to all customers who purchased the Tapered Screw-Vent Implant to inform them of the recall by Zimmer Dental. The letter informs the customers of the problems identified, risks involved, and the actions to be taken. Customers with questions are instructed to contact Zimmer Dental Customer Service at (800) 854-7019, 7-5pm, Mon-Fri. Customers are instructed to complete the customer response form and e-mail a copy to corporatequality.postmarket@zimmer.com or fax to (574) 372-4265. |  | 291 | 291 units | Worldwide Distribution-USA (nationwide) and the countries of Holland, Italy, and Japan. | Terminated | February 06, 2015 | 204 | 291 | Undetermined
Z-2154-2014 | 68746 | No | FCM | Urological catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | RUSCH Foley Urinary Catheter Insertion Tray with 10cc syringe, BZK Swabs, Sterile, Rx only, Teleflex Medical Product Code: 76730. |  |  | July 10, 2014 | 2014 | Teleflex Medical | Labeling error: The manufacturing date is the same as the expiration date. The manufacturing date is correct, however the expiration date incorrectly indicates that the product is expired. | Process control | Consignees were notified via letter on/about 07/10/2014. |  | 720 | 720 | Nationwide distribution in CA, CT, FL, MA, MO, NJ, OH, OR and TX. | Terminated | November 25, 2015 | 503 | 720 | Undetermined
Z-2153-2014 | 68686 | [6 empty] | Panthera-Puncher 9 . |  |  | July 01, 2014 | 2014 | PerkinElmer Health Sciences, Inc. | Risk of incorrect results from samples prepared with Panthera-Puncher 9 | Process design | Perkin Elmer sent an Urgent Medical Device Recall letter dated July 1, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. The Recall Letter delivered to the end customer by email facilitated through our Perkin Elmer Technical Support group located in Akron Ohio.Technical Support will be asked to make sure the customers receive the Recall Letter and implement the recommended precautions: "Removing the plates from the Panthera-Puncher 9 when not actively punching and processing samples in the instrument."Opening the plate track hatch if plates must reside in the Panthera-Puncher 9 while the instrument is idle or in sleep mode. This will interrupt power to the ionizer and eliminate any ionizer impact on the plate wells.For further information, please contact your local Perkin Eimer representative or GSTechnicaiSupport@PerkinEimer.com. For further information please call (781) 663-6900. |  | 144 | 144 units | Worldwide Distribution - USA (nationwide) and the countries of : Canada, Australia, Brazil, China, Denmark, Great Britain, Italy, France, Mexico, Qatar and Korea. | Terminated | October 21, 2015 | 477 | 144 | Undetermined
Z-2160-2014 | 68971 | No | DRS | Extravascular blood pressure transducer. | Cardiovascular | Cardiovascular | 510(k) | Trifurcated Transpac IV Monitoring Kit w/Safeset Reservoir, 03 ml Squeeze Flush and 6 Needleless Valves, Item No. 011-46106-22 |  |  | July 18, 2014 | 2014 | ICU Medical, Inc. | ICU has become aware of an issue with a component where the connection on an arterial monitoring kit may crack and cause a fluid leak. | Nonconforming Material/Component | ICU Medical, Inc. sent an email dated July 18, 2014, to their customers, The email inform the customer of the problem and the action to be taken. The email stated that the US Sales Representative would remove the recalled product on July 22, 2014. If product was further distribution customers should notify their customers. Customer was informed of the RGA to have the recalled product returned.For questions regarding this recall call 801-264-1732. |  | 60 | 60 units | Worldwide Distribution - USA including Oregon and Internationally to Australia. | Terminated | September 25, 2014 | 69 | 3570 | Undetermined
Z-2169-2014 | 68617 | No | JJH | Clinical sample concentrator. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | 5 mL Reaction Vessels (a consumable of the Abbott m2000sp Automated Fluid Handling System). The Abbott m2000 RealTime System is intended for use in performing nucleic acid amplification testing by polymerase chain reaction in clinical laboratories. The system is comprised of the m2000rt and the m2000sp instruments. |  |  | June 19, 2014 | 2014 | Abbott Molecular | Abbott Molecular identified that some of the 5 ml Reaction Vessels (RVs) contained in lots 56944001, 57026001, 57059001, and 57673001 (original vendor lot 2198077) may have an insufficient rim. An RV with an insufficient rim may not be held securely in the 1 mL Subsystem Carrier. This may cause the RV to be seated incorrectly in the carrier or drop through it. | Employee error | Abbott sent a Field Correction Recall / Urgent Field Safety Notice dated July 3, 2014, to all affected customers. June 19, 2014 were sent to all customers. The letters instructed customers to: 1) examine their inventories, 2) remove any of the affected lots, and 3) contact their Abbott Molecular Representativeat 1-800-553-7042 to coordinate replacement.Customers were also instructed to review this information with laboratory personnel and retain the communication for future reference.For questions regarding this recall call 224-361-7274 |  | 1688000 | 1,688,000 reaction vessels | Worldwide Distribution - USA (nationwide) and Internationally to:AUSTRALIAAUSTRIABELGIUMBENINCAMEROONCHINAETHIOPIAFINLANDFRANCEGERMANYINDIAINDONESIAIRELANDISRAELITALYKENYAMALAYSIAMALAWIMALINETHERLANDSNEW ZEALANDNORWAYPAKISTANPHILIPPINESPOLANDPORTUGALRUSSIASENEGALSINGAPORESOUTH AFRICASOUTH KOREASPAINSWEDENSWITZERLANDTAIWANTHAILANDTURKEYUKRAINEUNITED KINGDOMVIETNAM | Terminated | February 23, 2016 | 614 | 1688000 | Undetermined
Z-2170-2014 | 68177 | No | PEN | Multiplex nucleic acid assay for identification of microorganisms and resistance markers from positive blood cultures. | Microbiology | Microbiology | 510(k) | Verigene Gram-Negative Blood Culture Nucleic Acid Test (BC-GN), performed using the sample-to-result Verigene System, is a qualitative multiplexed in vitro diagnostic test for the simultaneous detection and identification of selected gram-negative bacteria and resistance markers. BC-GN is performed directly on blood culture media using blood culture bottles identified as positive by a continuous monitoring blood culture system and which contain gram-negative bacteria as determined by gram stain. BC-GN is indicated for use in conjunction with other clinical and laboratory findings to aid in the diagnosis of bacterial bloodstream infections; however, is not used to monitor these infections. Sub-culturing of positive blood cultures is necessary to recover organisms for antimicrobial susceptibility testing (AST), for identification of organisms not detected by BC-GN, to detect mixed infections that may not be detected by BC-GN, for association of antimicrobial resistance marker genes to a specific organisms, or for epidemiological typing. The BC-GN test is performed on the Verigene System platform, which is a fully automated, bench-top, molecular diagnostics workstation consisting of Verigene Reader and a bank of up to 32 Verigene Processor SP units. The System enables the detection of bacterial DNA from blood culture, unformed stool, or nasopharyngeal swab, depending on the test, through automated nucleic acid extraction and hybridization. |  |  | April 17, 2014 | 2014 | Nanosphere, Inc. | Nanosphere has determined that a portion of BC-GN Test Cartridge Lot number 040214021A has an elevated risk for No Call results and False Negative results. The BC-GN Test Cartridge (20-006-021) is distributed as part of the Verigene BC-GN Test Kit (20-005-021), which is labeled as "IVD". Therefore, Nanosphere has initiated a voluntary recall to retrieve the following BC-GN Test Cartridges: Test Ca | Under Investigation by firm | All customers were initially notified of the recall by telephone on April 17, 2014. The customers were instructed to quarantine and return all affected product to Nanosphere. Nanosphere sent an Voluntary Product Recall Notification letter dated April 17, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. The letters included instructions: 1) on how to return the recalled products, and 2) to review any test results that were generated while using the recalled products to ensure that there has been no negative impact on patient care. Customers with questions can contact Nanosphere at 888-837-4436, option 2 or via e-mail at Technicalsupport@nanosphere.us.. |  | 580 | 580 test cartridges | US Distribution including the states of AZ, CA, FL, MN, MT, NE, NJ, NY, TX and WI | Terminated | March 14, 2016 | 697 | 580 | Undetermined
Z-2168-2014 | 68708 | No | OZN | Clostridium difficile toxin gene amplification assay. | Microbiology | Microbiology | 510(k) | Nanosphere Verigene |  |  | October 16, 2013 | 2013 | Nanosphere, Inc. | Nanosphere has received report of an increased rate of "No Call - INT CTL 2" results for the Internal Control 2 (INT CTL 2) when negative samples are tested with certain lots of CDF Amplification Trays (20-011-022), which are contained in the CDF Amplification Reagent Kit (20-012-022); specifically for lots 061913022D and 071813022E. | Nonconforming Material/Component | Nanosphere sent an Field Notification Letter dated October 16, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.The letters included instructions for customers to: 1) discontinue use of the affected product, 2) discard any of the affected product, and 3) complete the attached form and return it to Nanosphere. Customers with questions can contact Nanosphere at 847-400-9000 or 888-837-4436 x 2. |  | 280 | 280 sets of trays | US Distribution including the states of CA and WA | Terminated | May 22, 2015 | 583 | 280 | Undetermined
Z-2172-2014 | 68838 | No | LKK |  |  | General Hospital | PMA | Medtronic Ascenda Intrathecal Catheter models 8780 and 8781, and Ascenda Intrathecal Catheter Pump Segment Revision Kit, model 8784. |  |  | July 11, 2014 | 2014 | Medtronic Neuromodulation | Medtronic is removing specific lots of the Ascenda Intrathecal Catheters and Revision Kits, which are used with the implantable SynchroMed drug infusion pump. This recall is being conducted due to a single component of the catheter, the retainer ring, not meeting specification criteria. There is a possibility of unintentional disconnection of the catheter from the pump, or difficulty in disconne | Nonconforming Material/Component | Medtronic sent an Urgent Medical Device Removal letter dated July 11, 2014, to all affected customers. The letter described the issue, identified affected product, provided patient management recommendations for patients implanted with an affected device, asked for Risk Management to quaratine unused inventory within the hospital, and to contact Medtronic Customer Service at 1-888-638-7627 to facilitate return of devices. |  | 1200 | 1200 (551 US, 649 OUS) | Worldwide Distribution - USA (nationwide) including the states of AL,AR,AZ, CA, CO, CT, DC, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, Ml, MN, MO, MS, MT, NC, NE, NJ, NM, NY, OH, OK, OR, PA, SC, TN, TX, UT, VA, VT, WA, WI and WV., and the countries of Austria, Belgium, Cyprus, Denmark, France, Germany, Iraq, Italy, Lebanon, Netherlands, Panama, Poland, Saudi Arabia, South Africa, Spain, Sweden, Switzerland, United Kingdom and Uruguay. | Terminated | July 17, 2015 | 371 | 1200 | Undetermined
Z-2173-2014 | 68633 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Medtronic MiniMed Implantable Insulin Pump, Model No. MMT-2007D |  |  | June 16, 2014 | 2014 | Medtronic MiniMed Inc. | Medtronic MiniMed is recalling the Implantable Insulin Pump because there have been reports of a pump malfunction resulting in the disconnection of the delivery of insulin prior to the end of the 7 year useful life of the device. | Under Investigation by firm | An Urgent Field Safety Notice Letter dated 6/16/14 was sent to all customers who purchased the Medtronic Implantable Insulin Pump. The letter informs the customers of the problems identified and the actions to be taken. Customers are informed that Medtronic MiniMed is working to make new unaffected pumps available as soon as possible. Customers are instructed to contact Elisabeth Andrieu at + 33 6 89 71 9056 or Dr. Francine Kaufman at +1 818.576.5331. |  | 764 | 764 units | New Jersey and Worldwide: Belgium, France, Netherland, and Sweden | Terminated | September 25, 2014 | 101 | 764 | Undetermined
Z-2178-2014 | 68844 | No | IWZ | Radiographic film. | Radiology | Radiology | 510(K) Exempt | Radiomat M+ NIF 14 x 17; Radiomat SG 14 x 17 NIF; RX-B 8 x 10 in./IBF-004355; Radiomat B+ NIF 14 x 17; Radiomat B+ NIF 8 x 10; Radiomat B+ NIF 9 1/2 x 9 1/2; CX Ultra UV-G Plus NIF 11 x 14; Radiomat G Plus 14 x 17 NIF |  |  | July 16, 2014 | 2014 | AGFA Healthcare Corp. | Some sheets of the medical screen film are fogged. | Process design | On July 16, 2014, an "URGENT FIELD SAFETY NOTICE" letter was sent to the US and Canadian dealers (consignees). The letter described the safety alert and the request to quarantine affected product. Acknowledgment, via FAX-Back or email, that the information was received and understood was requested from the dealers. A customer Urgent Field Safety Notice was also provided to those dealers who had shipped affected product to end users so that it could be sent to their customers. |  | 2617 | 2617 | Worldwide Distribution and US Nationwide in the states of AL, AZ, CA, FL, IL, LA, MI, MN, MS, NE, NH, NJ, NC and country of Canada | Terminated | December 09, 2014 | 146 | 2617 | Undetermined
Z-2174-2014 | 68507 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Siemens Linear Accelerators of type ARTISTE, ONCOR, and PRIMUS with Automatic Sequenced Cancer Treatment Delivery Option. |  |  | May 28, 2014 | 2014 | Siemens Medical Solutions USA, Inc | A safety risk exists with automatically sequenced treatment technique using the SIMTEC Auto Field Option where automatic movements of the gantry and the tabletop during an auto-sequenced treatment have the potential to lead to a collision with the patient. | Use error | An Urgent Medical Device Correction Field Safety Notice Letter was sent to end users on/about June 2014 which described the product, problem, and actions to be taken to avoid encountering the safety risk to patients and users. |  | 160 | 160 | Nationwide and ROW, including Canada and Mexico. | Terminated | July 23, 2015 | 421 | 160 | Undetermined
Z-2182-2014 | 68679 | No | NKB | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | Integra NewPort MIS System |  |  | June 27, 2014 | 2014 | Integra LifeSciences Corp. | Integra LifeSciences is recalling the Integra NewPort MIS System because there is a possibility that components from the demonstration set may have been mixed with the NewPort MIS pedicle screw system components used in surgery. | Under Investigation by firm | The firm, Integra, sent an "URGENT MEDICAL DEVICE CORRECTION FIELD SAFETY NOTICE" dated June 27, to its customers. The letter informs the customers that Integra Lifesciences is recalling the Integra NewPort MIS System because there is a possibility that the components from the demonstration set may have been mixed with the NewPort MIS pedicle screw system components used in the surgery. The letter also informs the customers of the problems identified and the actions to be taken. Customers are instructed to complete the attached response form and return it by email or fax as indicated on the form. Customers with questions regarding the recall letter instructions are instructed to contact the Marketing Manager at (760) 207-7050. |  | 4 | 4 units | US Distribution in TX only. | Terminated | September 16, 2014 | 81 | 4 | Undetermined
Z-2183-2014 | 68890 | No | NIG | Tumor-associated antigen immunological test system. | Immunology | Immunology | 510(k) | VITROSÂż Immunodiagnostic Products CA 19-9 Calibrators and Reagent Packs, Lot 1320 |  |  | July 16, 2014 | 2014 | ORTHO-CLINICAL DIAGNOSTICS | Customers may receive positively biased results when using VITROSÂż Immunodiagnostic Products CA 19-9 Total Reagent Packs, Lot 1320 and VITROSÂż CA 19-9 Calibrators, Lot 1320 on the VITROS ECi/ECiQ Immunodiagnostic Systems, the VITROS 3600 Immunodiagnostic System and the VITROS 5600 Integrated System. There is an unexpected observed change in CA 19-9 Calibrators, Lot 1320 calibrator values. | Other | Ortho Clinical Diagnostics sent an Urgent Product Correction Notification July 2014 to all affected customers. The Notice identified the product, the problem, and the action to be taken by the customer. Customers were instructed to Immediately discontinue using and discard Lots 1320 of VITROS CA 19-9 Calibrators and Reagent Packs.Review previously reported VITROS 25-OH Vitamin D Total results using Lot 1320. Review quality control results used to verify calibration events. Results lower than the target mean may indicate potentially biased assay results.Discuss any concerns regarding previously reported results with the Laboratory Medical Director or with the requesting physician.Complete and return the attached Confirmation of Receipt form. OCD will:credit their account for all VITROSÂż CA 19-9 Calibrators and Reagent Packs that were discarded or ship repalcment product uspon request.Customers were instructed to forward the information in this notification if they have distributed this product outside of their facility.Customers with questions were instructed to contact their Customer Technical Service representative or contact the Customer Technical Service representative at 1-800-421-3311. |  | 475 | 475 | Worldwide Distribution - USA (nationwide) and Internationally to France and Japan | Open, Classified |  |  | 475 | Undetermined
Z-2120-2014 | 68302 | No | GEX | Laser surgical instrument for use in general and plastic surgery and in dermatology. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | GYC-1000 Green Laser Photocoagulator system |  |  | May 16, 2014 | 2014 | Nidek Inc | Laser Aperture label was not applied to certain GYC-1000 Green Laser Photocoagulators devices in the field as per FDA regulation, | Radiation Control for Health and Safety Act | Nidek sent an Medical Device Correction Notifications dated May 16, 2014, informing doctors of the lack of the Laser Aperture Label for GYC-1000 Green Laser Photocoagulator System. Letters included a Aperture label and instruction/illustration on where to affix it to the device. A response form is to be completed and returned indicating that they read the correction information and have affixed the label on the GYC-1000 device as instructed. Questions may be directed to Nidek Customer service at 1-800-223-9044 Monday - Friday from 8 am to 5 pm PST.CDRH approves the CAP subject to the following conditions:1. Notification of all dealers and purchasers is to be made within 15 working days of receipt of this letter in the manner specified in 21 CFR 1003.21 and 1003.22. This office and the Food and Drug Administration (FDA) district office coordinator noted below are to be included in the notification.You may proceed or continue with implementation of the CAP (subject to the conditions noted above). |  | 567 | 567 units | USA (nationwide) Distribution | Terminated | October 16, 2014 | 153 | 567 | Undetermined
Z-2214-2014 | 68840 | No | JIX | Calibrator. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | VITROSÂż Chemistry Products CDM PROM Calibration Data Module Rev. No. 0168 and 0169 (Product Code 1999077) are used in conjunction with VITROS Chemistry Products NH3 DT Slides (Product Code 1532589) GEN 59 and VITROSÂż Chemistry Products DT Calibrator Kit lot 44 (Product Code 1957927) and tested on the VITROS DT60/DT60 II Chemistry System: IVD, Made in USA by Ortho-Clinical Diagnostics, Inc. Rochester, NY |  |  | July 15, 2014 | 2014 | Ortho-Clinical Diagnostics | Inability to calibrate or use VITROS NH3 DT Slides, GEN 59 with VITROS Chemistry Products DT Calibrator Kit, Lot 44 when CDM PROM 0168 or 0169 is installed on VITROS DT60/DT60 II Chemistry Systems. | Process control | Ortho Clinical Diagnostics sent an Urgent Product Correction Notification letter dated July 15, 2014,by Federal Express overnight to all affected customers. All US end-user consignees who had been shipped VITROS NH3 DT slides, GEN 59, Product Code 1532589 were informed of the issue. Customers were instructed to: a) use VITROS DT NH3 slides, GEN 59 with Calibration Data Module, Rev. 0168 or 0169 with VITROS DT Calibrator, Lot 42 if available; or b) If VITROS DT Calibrator, lot 42 is not available, the user is instructed to request CDM, Rev. 0170 to be used with VITROS NH3 DT slides, GEN 59 and VITROS DT Calibrator, Lot 44.Customers were instructed to forward this notice if the product was further distributed outside of their facility.OCD has manufactured a new CDM, Rev. 170, containing the necessary data/calibration mathematics for VITROS DT Calibrator, Lot 44 to be used to calibrate VITROS NH3 DT slides, GEN 59.Customers with questions were instructed to call Customers Technical Services at 1-800-421-3311.For questions regarding this recall call 908-218-8776. |  | 5906 | 5906 total: 240 (US); 5666 (OUS) | Worldwide Distribution - USA including MA, TX, KY, IL, MS, OH and Puerto Rico; Internationally to: Australia, Brazil, Canada, Chile, China, Colombia, England, France, Germany, India, Italy, Japan, Mexico, Panama, Singapore, Spain, and Venezuela. | Open, Classified |  |  | 5906 | Undetermined
Z-2211-2014 | 68763 | No | LGX | Operating tables and accessories and operating chairs and accessories. | General &amp; Plastic Surgery | General Hospital | 510(K) Exempt | Medical Positioning Rapid Response HUT table, Powered Medical Examination table, Model # 1219. |  |  | July 10, 2014 | 2014 | Medical Positioning Inc | Excessive weight on the table will cause tables with the affected component to drop into the Trendelenburg position more rapidly than normal. The product was manufactured with a component that can cause the device to malfunction. | Component design/selection | Medical Positioning Inc. sent Potential Product Hazard Bulletin letter dated July 8, 2014 to affected customers, via USPS certified return receipt mail. The letter identified the affected product, problem and actions to be taken. Customers were informed a Product Support Specialist will contact them to make arrangements for a Field Service Technician to perform a test designed to identify if the table component affects the table mechanics. If the table is identified as having the affected component, the component will be replaced. |  | 1 | 1 bed | Worldwide Distribution - US Nationwide in the states of CA, GA, FL, IL, IN, KS, KY, MD, MN, MO, MS, NC, NY, OH, TN, TX, VA and WI. | Terminated | January 03, 2017 | 908 | 40 | Undetermined
Z-2560-2014 | 69050 | No | MBI | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | ACL Sleeve Installation Tool Sizes: 6mm, 6.5mm, 7.0mm, 7.5mm, 8.0mm, 8.5mm, 9.0mm, 10.0mm |  |  | July 25, 2014 | 2014 | OrthoPediatrics Corp | Weld fractures on the Sleeve Installation Tool for ACL reconstruction system. | Nonconforming Material/Component | OrthoPediatrics sent a Notification letter dated July 25, 2014 to all affected customers. The letter Identified the affected product, problem and actions to be taken. The letter informed customers to take immediate possession of all affected product within their control and at any of their health care facilities and immediately return the recalled device(s) to OrthoPediatrics Corp. This notice and the attached Reply Form should be forwarded if the affected product was further distributed. . |  | 56 | 56 units | US Nationwide Distribution. | Terminated | May 19, 2015 | 298 | 56 | Undetermined
Z-2149-2014 | 68665 | No | NBC | B-type natriuretic peptide test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | SIEMENS ADVIA Centaur CP Immunoassay System BNP kit; SMN 10309044 - 100 test |  |  | June 23, 2014 | 2014 | Siemens Healthcare Diagnostics, Inc | Complaints of failed calibration. | Other | Siemens sent an Urgent Medical Device Correction (UMDC) letter to all affected Siemens Healthcare Diagnostics customers in the United States on June 23, 2014 via Federal Express. An Urgent Field Safety Notice (UFSN) was sent out for distribution to all customers outside the United States on June 23, 2014 via hard copy, email, and/or fax per each country's local regulations. These notices inform the customer of the potential for low calibrator deviation failures when using ADVIA CentaurÂż Systems BNP kit lots ending in 172 on the ADVIA CentaurÂż CP due to an error on the master curve card. Siemens informs customers if a successful calibration is achieved, the results obtained are valid. If customers are unable to obtain a successful calibration on the ADVIA CentaurÂż CP, they should switch to an alternative lot of BNP assay. The notices should be reviewed with the facility's laboratory or medical director.Fax-back forms will be used to confirm that the customers received the UMDC. For those customers that do not return a fax-back form, a Siemens representative will follow up with those customers to ensure they received the Urgent Medical Device Correction and completed the required actions. Customers with questions were instructed to contact their Siemens Customer Care Center or their local Siemens technical support representative.For questions regarding this recall call 877-229-3711. |  | 1191 | 1,191 kits | Worldwide Distribution - USA (nationwide including Puerto Rico) and Internationally to Canada, Albania, Argentina, Australia, Bosnia Herzeg., Brazil, Chile, Colombia, Egypt, Georgia, Hong Kong, India, Indonesia, Iran, Iraq, Israel, Japan, Kazakhstan, Lebanon, Malaysia, New Caledonia, Oman, P.R. China, Pakistan, Philippines, Qatar, Rep. of Yemen, Republic Korea, Russian Fed., Saudi Arabia, Serbia, Singapore, South Africa, Taiwan, Thailand, U.A.E., Ukraine, Venezuela, Vietnam, Austria, Bulgaria, Croatia, Cyprus, Czech Republic, Denmark, Estonia, Finland, France, Germany, Greece,Hungary,Iceland, Ireland, Italy, Latvia, Lithuania, Netherlands, Norway, Poland, Portugal, Romania,Slovakia, Slovenia, Spain, Sweden, Switzerland, Turkey, and United Kingdom. | Terminated | March 31, 2016 | 647 | 2152 | Undetermined
Z-2720-2014 | 69063 | No | JIX | Calibrator. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Siemens Dimension VistaÂż CHEM 1 CAL, lot 3GM081. |  |  | August 18, 2014 | 2014 | Siemens Healthcare Diagnostics, Inc. | Siemens Healthcare Diagnostics has received complaints for low recovery of Magnesium (MG) Quality Control (QC) and patient samples following calibration of MG with the Dimension Vista&reg; CHEM 1 CAL, lot 3GM081. Investigation by Siemens confirms that QC and patient samples under-recovers MG by 0.25 mg/dL [0.10 mmol/L] across the assay range following calibration with 3GM081. | Process design | Siemens sent an Urgent Medical Device Dorrection letter, dated August 2014, to all affected consignees. The letter identified the product, the problem, and the actions to be taken. Consignees were instructed if they still have an active calibration of MG with this lot of CHEM 1 CAL, they should immediately calibrate MG with an alternate lot of CHEM 1 CAL. Consigness were instructed to retain the letter with their laboratory records, and to forward the letter to those who may have received the affected product. Customers with questions were instructed to contact their Siemens Customer Care Center - Techhnical Solutions or their local Siemens technical support representative. |  | 9769 | 9769 Total (6000 US) | Worldwide Distribution - USA (Nationwide) and Internationally to AustraliaAustriaBelgiumCanadaCzech RepublicFinlandFranceGermanyGreat BritainItalyJapanNetherlandsNew ZealandNorwayPortugalSlovak RepublicSloveniaSouth KoreaSpainSwitzerland | Terminated | February 02, 2015 | 168 | 9769 | Undetermined
Z-2691-2014 | 69140 | No | GDW | Implantable staple. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Justright Surgical, JustRight 5mm Reload, REFJR-REL25-2.0, Sterile, EO, Rx Only |  |  | August 27, 2014 | 2014 | Justright Surgical, LLC | JustRight Surgical is recalling specific lots of JustRight 5mm Stapler and Reload products because they may deliver poor distal staple formation, past the end of the knife cut. | Process design | Consignees were notified via letter on August 27, 2014 in addition, the sales rep will hand deliver the Recall Notification on 8/28/14. Acknowledgement Form will be provided and returned products will be sent to STD Med where it will be held in quarantine. |  | 39 | 3 cases containing 36 units | Distributed in the states of Colorado, California and New York. | Terminated | October 06, 2014 | 40 | 60 | Undetermined
Z-2656-2014 | 69057 | No | FNL | AC-powered adjustable hospital bed. | General Hospital | General Hospital | 510(K) Exempt | Zenith Slide W-I-D-E Long Term Care Bed, Model(s) APS98674, APS9867406, APS986741, APS986742, ZG58674, ZG78674, ZZ78674, ZZ98674, and ZZ98774. |  |  | July 21, 2014 | 2014 | GF Health Products, Inc./d.b.a.Basic American Metal Products | The bed head deck may fail to elevate as expected during normal operation. When the non-operation occurs, the head deck remains in the lowered horizontal / flat position and will no longer function as intended. | Nonconforming Material/Component | Consignees were sent a Basic American Medical Products "Product Field Correction Notice" dated July 21, 2014. The letter provided the background information on the recall and the " Required Action" to be taken which includes the replacement of the head deck. Consignees need to contact 1-888-974-4347 to scheduled the replacement order. The replaced head decks should be disposed and disposal documented on the enclosed Letter of Disposal Agreement. The completed letter needs to be returned to Basic American Medical Products. For additional questions they can contact the Field Correction Coordinator at 1-888-974-4347 or electronically: fieldcorrection@grahamfield.com. |  | 1941 | 1941 | Worldwide Distribution: USA (nationwide) including the states of AL, AZ, AR, CA, CO, CT, FL, GA, IL, IN, IA, KY, LA, ME, MD, MA,MI, MN, MS, MO,MT, NE, NH, NJ, NM, NY,M NC, ND, OH, OK, OR, PA, RI, SC, SD, TN, TX, LT, VA, WA, WV, and WI and PUERTO RICO and the country of CANADA. | Terminated | August 11, 2015 | 386 | 1941 | Undetermined
Z-2658-2014 | 69145 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | ZOLL Pro-padz Sterile Adult Multi-Function Electrodes |  |  | August 26, 2014 | 2014 | Bio-Detek, Inc. | Sterility of device is not assured | Process control | BioDetek sent an Urgent Device Correction letter dated August 26, 2014, to all affected customers via UPS overnight delivery. There has also been notifications addressed to the 4 International accounts on August 27, 2014. Customers were requested to return the suspect product(s) in their possession and complete the Device Correction Response ReturnForm provided to each customer. Customer inventory will be replaced as soon as replacement electrodes are processed. Customers with questions can contact the technical support numbers at 1 (800) 348-9011 or +1 (978) 421-9460 24/7.For questions regarding this recall call 401-729-1400. |  | 198 | 198 cases | Worldwide Distribution - USA (nationwide) and Internationally to Australia, Canada, and Germany. | Completed |  |  | 310 | Undetermined
Z-2619-2014 | 69088 | No | FMG | Intravascular administration set. | General Hospital | General Hospital | 510(k) | STAT 2 Flow Controller; 80" (203 cm) Appx 14.8 mL, 20 Drop Primary Set |  |  | August 19, 2014 | 2014 | ICU Medical, Inc. | ICU is recalling the ConMed Stat 2 flow controller because it may deliver a significantly higher flow rate than was intended. | Nonconforming Material/Component | ICU Medical sent an Urgent Medical Device Recall letter dated August 19, 2014 to all affected customers. The letter identified the affected product, problems identified and the actions to be taken. Customers were instructed to complete the recall response form and fax to ICU Medical at (801) 264-1755 or by email: recall@icumed.com. Customers with questions or require assistance relating to this recall were instructed to contact ICU Customer Service Monday through Friday between the hours of 8:30 AM and 4:00 PM Pacific time: 866-829-9025 and select option 8, or e-mail the following address:customerservice@icumed.com.For questions regarding this recall call 801-264-1732. |  | 3750 | 3750 units | Worldwide Distribution - US (nationwide) in the state of WA, OR and the country of South Africa. | Terminated | November 18, 2014 | 91 | 6250 | Undetermined
Z-2593-2014 | 69020 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | CareFusion EnVe and ReVel Ventilators. |  |  | August 08, 2014 | 2014 | CareFusion 203, Inc. | CareFusion has identified a potential for damage to the power connector on the EnVe and ReVel ventilators due to misconnection by the operator. The input port on the ventilators must have the adapters correctly aligned to function properly. Incorrect alignment may cause the pins in the connector to bend, causing the PTV AC power adapter and/or PTV Auto Lighter Power Cord to be shortened, and pre | Component design/selection | CareFusion sent a Product Recall Notice dated August 8, 2014, to all affected customers. The letter listed the product involved in the recall, described the Issue, Potential Risk , Actions to be Taken by Carefusion, Action to be Taken by the Customer and Other information. Customers were not required to return the devices. Customers were asked to perform the following actions:Once the kit(s) has/have been received, these kits must be fitted and a confirmation must be returned to CareFusion for each kit fitted, identifying the ventilator serial numbers and/or number of AC Power Adapters, Auto Lighter Power Cords to which the kit has been fitted. The confirmation card will be in each kit.To avoid possible delays in providing ventilation and for future compatibility, customers must update their entire inventory. Warranty is void if vent-side LEMO adapter is not properly used in accordance with its instructions and ventilator is damaged due to misuse.Contact Information: CareFusion Support Center, SupportCenter@CareFusion.com Phone 1-888-562-6018, Fax 1-858-617-4851 (Monday through Friday 7AM to 4PM Pacific); Customer Service &amp; Technical Support, ACVentCare@CareFusion.com, Phone: 1-800-754-1914; FAX CS: 1-763-398-8403 (24 hours and 7 days a week between the hours of CST 8:00 am - 5:00pm. After hours and on weekends expect a one hour call back.For questions regarding this recall call 888-562-6018. |  | 1729 | 1729 | Nationwide Distribution including AL, AZ, CA, CO, CT, FL, GA, IL, IN, KS, KY, LA, MA, MI, MN, MS,M0, NE, NV,NH,NH, NY NMC, OH, OR,PA, RID, SC, TN, TX, UT VT, VA, WA, and WI. | Terminated | October 07, 2015 | 425 | 1729 | Undetermined
Z-2672-2014 | 67912 | No | FMI | Hypodermic single lumen needle. | General Hospital | General Hospital | 510(k) | EZ-IO 45mm (40kg &amp; up) Needle Set, part number 9079-VC-005; lot code 3606455 |  |  | March 31, 2014 | 2014 | Vidacare Corporation | The voluntary recall is due to the needle set not meeting its specifications. | Material/Component Contamination | Vidacare sent an Urgent Medical Device Recall Notification to all of its customers on March 31, 2014.The letter identified the product the problem and the action to be taken by the customers. The firm asked its customers to return the affected lot for a replacement.Customers were instructed:1. If you have affected stock, immediately discontinue use and quarantine any product with the part number and lot number as identified above2 To return product. complete the enclosed Medical Device Recall Acknowledgment and Return Response Form and fax it to 1 866-431-6194, Attn "Customer Service" This will allow us to document the amount of product you have on hand for return. A customer service representative will contact you with a Return Goods Authorization (RGA)Number and will provide instructions for the return of product to Vidacare Corporation and to facilitate replacement product if one has not already been provided3. If you have already used the affected stock. please complete the enclosed Medical Device Recall Acknowledgement and Return Response Form and fax it to 1-866-431- 6194, Attn" Customer Service". This will allow us to document your receipt of the letter.Vidacare Corporation is committed to providing high quality, safe and effective products. We sincerely apologize for any inconvenience this action may cause your operations. If you have any other questions, feel free to contact your local sales representative or Customer Service at 1-866-479-8500. |  | 172 | 172 | Worldwide Distribution-US (nationwide) including the states of LA, NY, CA, CT, WA, AK, NJ, TX, GA, MI, MO, MD, FL, KY, VA, OH, UT, SC, NC, IN, AZ, MN, IL, TN, OK, KS, NE, MT, MS, AR, ID, WV and PA., and Internationally to Canada | Terminated | January 30, 2015 | 305 | 172 | Undetermined
Z-2674-2014 | 69178 | No | KWQ | Spinal intervertebral body fixation orthosis. | Orthopedic | Orthopedic | 510(k) | DePuy Synthes TI Vectra Plates: |  |  | September 02, 2014 | 2014 | Synthes, Inc. | DePuy Synthes is initiating a voluntary medical device recall of certain lots of the TI Vectra Plates, which is a part of the DePuy Synthes Vectra Anterior Cervical Plate Family of Systems. The Vectra Systems are intended for anterior screw fixation to the cervical spine (C2-C7). It was discovered in certain lots, that the clips are potentially missing from the TI Vectra Plates resulting in the i | Employee error | DePuy Synthes sent an Urgent Notice Medical Device Recall letter dated September 2, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Please take the following actions:If you DO have any of the identified devices, please take the following steps: Call DePuy Synthes at 1-800-479-6329 to obtain a Return Authorization (RA) Number. Complete the Verification Section (page 3 of this letter) by checking the appropriate box indicating affected product has been located. Also, please indicate the number of devices found and note the Return Authorization Number. Please include your name, title, address, telephone number and signature in the spaces provided. Return the Verification Section (page 3 of this letter) with the product to: Credit/Returns, DePuy Synthes, 1101 Synthes Avenue, Monument, CO 80132. Send a copy of the completed Verification Section to DePuy Synthes by: Fax: (610) 430-7083 or Scan/email: Fieldaction@synthes.comIf you DO NOT have the identified product, please take the following steps: Complete the attached Verification Section (page 3 of this letter) by checking the appropriate box indicating that no affected product has been located. Please include your name, title, address, telephone number and signature in the spaces provided. This return documentation acknowledges your receipt of medical information. Return the documents to DePuy Synthes by: Fax: (610) 430-7083 or Scan/email: Fieldaction@sythes.com Note: If the Verification Section is answered on behalf of more than one facility and/or individual, please clearly indicate the name and address of the facility and/or individual on page 3 of the notification.If you have any questions, please call 610-719-5450, or contact your DePuy Synthes sales consultant. |  | 16 | 16 | US Distribution including the states of IL, SC, VA, TX, MA, NC, LA, CA, AL, CO, NE, OR and KS., and Internationally to Canada. | Terminated | September 02, 2015 | 365 | 16 | Undetermined
Z-2677-2014 | 67900 | No | NKB | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | Circular Lock Screws, part # G826, lot # 11406. The Lock Screws are part of the Genesys Spine TiLock Pedicle Screw System. |  |  | April 02, 2014 | 2014 | Genesys Orthopedics Systems, LLC | The recall is being initiated because MK Precision (the Circular Lock Screw supplier) notified Genesys Spine that a portion (less than 20%) of the Circular Lock Screws in lot 11406 were manufactured using Stainless Steel in the center Post instead of the specified titanium alloy. | Nonconforming Material/Component | Genesys Spine sent an Urgent Medical Device Recall letter dated April 2, 2014, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to immediately examine their inventory and quarantine any affected product. Customers were instructed to notify their customers (end users) of the recall and request that they return any affected product to Genesys Spine by completing the attached Recall Tracking Form to Genesys Spine, 1250 South Capital of Texas Hwy, Building Three, Suite 600, Austin, Texas 78746.Customers with questions were instructed to call 512-381-7070. |  | 642 | 642 | Nationwide Distribution including NV, CA, NY, MO, TX, FL, NJ, GA, TN. | Terminated | May 15, 2015 | 408 | 642 | Undetermined
Z-2676-2014 | 68721 | No | GAT | Nonabsorbable poly(ethylene terephthalate) surgical suture. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | The synthetic non-absorbable surgical sutures CARDIOFLON Evolution are composed of polyester [poly(ethylene terephtalate)] braids, coated, with or without PLEDGETS of different sizes. |  |  | June 25, 2014 | 2014 | Peters Surgical | A potential splitting of the surface treatment of the thread ends on certain sutures has recently been identified. | Process control | The sole US distributor (MedAlliance Inc.) was notified of the recall via e-mail on July 16, 2014. IMPORTANT AND URGENT - BATCH RECALL PROCEDURE letters dated June 25, 2014 under the Peters Surgical header were sent to all MedAlliance Inc. consignees. The letters included instructions for consignees to contact MedAlliance Inc. by phone (630-933-8661) or by e-mail (agoos@medalliancesolutions.com) for information on what to do with the recalled products. Consignees were also notified that they can contact Peters Surgical directly at qualite@peters-surgical.com for answers to any additional questions. MedAlliance Inc. sent VOLUNTARY PRODUCT RECALL/WITHDRAWAL LETTER[s] dated June 25, 2014 to all consignees. The letters included instructions for consignees to: 1) discontinue use of and segregate recalled product in a secure location for return to MedAlliance Inc., and 2) to complete and return the enclosed RETURNED GOOD AUTHORIZATION (RGA) FORM to MedAlliance Inc. to arrange for the return of the recalled products. Consignees with any questions about the recall can contact Ms. Alaina Goos, Senior Product Manager, by phone at 630-933-8661 x 104 or via e-mail at agoos@medalliancesolutions.com. |  | 22190 | 22,190 units | Worldwide Distribution- US: Nationwide (CA, FL, GA, IL, IN, LA, MD, MI, MO, MS, NC, NY, OH, OR, PA, SC, TX, VA, WA); and countries: ALGERIA, ARGENTINA, AUSTRALIA, AUSTRIA, BANGLADESH, BELGIUM, CAMBODIA, CHINA, CYPRUS, CZECH REPUBLIC, DOMINICAN REPUBLIC, EGYPT, FRANCE, GEORGIA, GERMANY, HUNGARY, INDONESIA, IRELAND, ISRAEL, ITALY, JAPAN, KAZAKHSTAN, LEBANON, LIBYA, MALAYSIA, MAURITIUS, MOLDOVA, MOROCCO, MYANMAR, NETHERLANDS, NEW ZEALAND, NORTH KOREA, PAKISTAN, PERU, POLAND, ROMANIA, RUSSIA, SAUDI ARABIA, SOUTH AFRICA, SOUTH KOREA, SPAIN, SRI LANKA, SWITZERLAND, TAIWAN, THAILAND, UNITED KINGDOM, VENEZUELA, VIETNAM, YEMEN | Open, Classified |  |  | 27166 | Undetermined
Z-2679-2014 | 67727 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | Mini Cannulated Titanium Headed and Headless Screw Set and 2.5, 3.0, and 4.0 mm countersink/depth gauges in the Mini Screw Set/System |  |  | April 25, 2013 | 2013 | Instratek, Incorporated | Faded and wrong markings on bone screw Countersink/Depth Gauge instruments in Mini Screw Set. | Nonconforming Material/Component | The firm plans to notify their customers of the recall via email, telephone and visit. The firm then plans to scrap and discard any affected product. |  | 96 | 96 | Nationwide: Maine, Oregon, Florida, Texas, Virginia, Arizona, Nebraska, California, Ohio, Illinois, Oklahoma, Washington, North Carolina, Arkansas, Tennessee, New Mexico, Connecticut, Colorado | Terminated | January 26, 2016 | 1006 | 96 | Undetermined
Z-2686-2014 | 69158 | Yes | KPO | Hemodialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Fresenius Naturalyte Liquid Bicarbonate Concentrate |  |  | June 23, 2014 | 2014 | Fresenius Medical Care Holdings, Inc. | Product was held at temperature above the labeled recommended storage temperature | Storage | Fresenius Medical North America contacted customers via telephone on 6/23/14 by Customer Service and follow-up with formal letter notification, Urgent Recall by certified mail with signature confirmation and faxback form Customers instructed to examine their inventory to determine whether they have any of the affected Naturalyte Liquid Bicarbonate. If customers have the affected product, they are instructed to contact FMC-RTG Customer Service to have the product replaced.A revised/clarification letter dated 7/10/14 issued to state only products from the identified lots that were delivered by RTG LLC to the specified facilities on May 5, 2014 are affected by this recall as products were exposed to temperatures higher than their recommended storage temperature during transportation on May 5, 2014. If you have any additional questions, please contact your FMCNA Customer Service Team at 1-800-323-5188. |  | 72 | 72 cases | Distributed in Texas. | Open, Classified |  |  | 72 | Undetermined
Z-2566-2014 | 67547 | No | LZD | Total temporomandibular joint prosthesis. | Dental | Dental | PMA | Biomet Microfixation Temporomandibular Joint (TMJ) Patient Matched Left Fossa Part# CP751506 |  |  | September 20, 2013 | 2013 | Biomet Microfixation, LLC | Product mix occurred prior to final package and labeling process. | Mixed-up of materials/components | Biomet Microfixation, Inc. sent an Urgent Medical Device Field Notice dated September 20, 2013 to its customer. The customer was instructed to: 1) Immedicately quarantine the device. 2) Discontinue use and return to Biomet Microfixation. 3) If the product was further distributed, the letter must be forwarded to the firm's customers. 4) A replacement device will be shipped priority upon completion. Questions related to this notice should be directed to 904-741-9448 or 904-239-6671.For questions regarding this recall call 904-741-4400. |  | 1 | 1 | Worldwide Distribution to Denmark only. | Terminated | October 03, 2014 | 378 | 1 | Undetermined
Z-2680-2014 | 69148 | No | KPR | Stationary x-ray system. | Radiology | Radiology | 510(k) | The Ysio Systems with software version VC10 |  |  | August 27, 2014 | 2014 | Siemens Medical Solutions USA, Inc | Siemens discovered that an unlikely error may occur on the Ysio system with fixed detector in the wall stand. This error may result in line artifacts in the image. If these artifacts appear in the region of interest, the examination may need to be repeated. | Under Investigation by firm | The firm, Siemens, sent a "Safety Advisory Notice" dated August 27, 2014, to end users/customers. The notice identified the product, problem, and action to be taken. The customers were instructed to follow the instructions in the notice; promptly notify and instruct all the staff at your organization who have to be aware of this problem; forward this safety information to other organizations affected by this measure; and if the device has been sold and therefore no longer in your possession, forward to new owner.Siemens prepared a field modification of the Ysio systems with a hardware upgrade. This field modification will be available in September 2014.If you have any questions, please contact Regulatory Specialist at 610-219-4334 or Customer support at 610-219-6300. |  | 62 | 62 | Nationwide Distribution. | Terminated | March 30, 2015 | 215 | 62 | Undetermined
Z-2681-2014 | 69034 | No | PEN | Multiplex nucleic acid assay for identification of microorganisms and resistance markers from positive blood cultures. | Microbiology | Microbiology | 510(k) | The Verigene Gram Negative Blood Culture Nucleic Acid Test (BC-GN), performed using the sample-to-result Verigene System, is a qualitative multiplexed in vitro diagnostic test for the simultaneous detection and identification of selected gram-negative bacteria and resistance markers. BC-GN is performed directly on blood culture media using blood culture bottles identified as positive by a continuous monitoring blood culture system and which contain gram negative bacteria as determined by gram stain. |  |  | October 10, 2013 | 2013 | Nanosphere, Inc. | Nanosphere has received several reports relating to an increased rate of Processing Errors during procedures performed on the Processor SP caused by Tips in Tip Holder Assemblies for the following products and lots of Extraction Trays. | Nonconforming Material/Component | Nanosphere sent a Field Notification Letter dated October 10, 2013, to all affected customers. The letter identified the product, the problem and the action to be taken by the customer. Instructions were included in the letter for customers to determine whether or not the extraction tray lots should be returned. If the extraction tray lots are to be returned, customers were to complete and return the attached form and replacement tip holder assemblies will be supplied at no cost. Customers could then dispose of the recalled unused tip holder assemblies. Customers with questions were instructed to contact their Nanosphere technical support representative.For questions regarding this recall call 1-888-837-4436, ext 3. |  | 400 | 400 units | Worldwide Distribution - USA including AZ, CA, FL, GA, IA, IL, IN, KY, MA, MD, MN, MO, MT, NC, NE, NJ, NY, OH, OR, PA, TX, UT, VA, WA, WI, OH and Internationally to Bulgaria, Germany, Italy, Japan, South Korea, and Spain. | Terminated | June 16, 2015 | 614 | 5920 | Undetermined
Z-2687-2014 | 69107 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | UniCel 600/800 Access Immunoassay Systems are in vitro diagnostic devices used for the quantitative, semi-quantitative, or qualitative determination of various analytes concentrations found in human body fluids. The cables and kits are used in UniCel DxI 600, 800 instruments, and are distributed to Beckman and Non-Beckman Service Engineers for replacement of worn and/or damaged parts during field services calls. |  |  | August 21, 2014 | 2014 | Beckman Coulter Inc. | Beckman Coulter is recalling the Sample Probe Cable because nonconforming sample probe cables may been installed on newly manufactured and/or recently serviced UniCel DxI 600 and 800 Access Immunoassay Systems, UniCel DxC 660i, 680i, and 860i, and UniCel DxC 880i Synchron Access Clinical Systems instruments. Use of a nonconforming sample probe cable can casuse a delay in reporting results, but wi | Under Investigation by firm | A customer notification letter dated 8/21/14 was sent to all customers who purchased the Sample Probe Cable for use with the UniCel DxI Family of Immunoassay Systems. The letter informs the customers of the problems identified and the actions to be taken. Customer are instructed to complete and return the enclosed response form within 10 days. Customers are instructed to call (800) 854-3633 in US and Canada. |  | 236 | 236 units total (70 units in US) | Nationwide in US and worldwide: Canada, China, Colombia, Croatia, Czech Republic, France, Germany, India, Italy, Japan, Korea, Kuwait, Malaysia, Mexico, Philippines, Puerto Rico, Russia, Singapore, Slovakia, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, United Kingdom, United States, Vietnam. | Terminated | April 30, 2015 | 252 | 236 | Undetermined
Z-2692-2014 | 69105 | No | JGJ | Magnesium test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | ABX PENTRA Magnesium RTU Reagent, Part No. A11A01646 |  |  | August 11, 2014 | 2014 | Horiba Instruments, Inc dba Horiba Medical | Horiba Medical is recalling the ABX PENTRA Magnesium RTU Reagent because the on-board stability claims has been updated. Horiba has updated the ABX PENTRA Magnesium RTU package inserts by removing the 7-day on-boardstability claim and replacing it with a 1-day on-board stability claim. | Other | Horiba sent an Urgent Safety Notice letter dated August 11, 2014 to all customers running the Magnesium RTU assay on the ABX PENTRA 400 Chemistry analyzers that the On-Board Stability has been updated. The letter informs the customers of the problems identified and the actions to be taken. Customers are instructed to complete and return the enclosed response form within 10 days. For questions call 888-903-5001, ext 4266. |  | 359 | 359 units | US Nationwide Distribution. | Terminated | November 25, 2014 | 106 | 359 | Undetermined
Z-2695-2014 | 69035 | No | IOR | Mechanical wheelchair. | Physical Medicine | Physical Medicine | 510(k) | Fuze Manual Tilt Wheelchair. |  |  | June 12, 2012 | 2012 | Pdg Product Design Group Inc | The hex bolt head might shear and cause the axle plate to detach from the lower frame of the wheelchair and could collapse the wheels while in use. | Component design/selection | PDG Product Design group sent all Canadian customers via an e-mail dated August 9, 2012. The letter identified the product the problem and the action needed to be taken by the customer. Immediately call PDG Customer Service @ 1-888-858-4422. Advise the Customer Service Representative that you have received this letter. Our Customer Service Representative will be happy to organize for you to receive a brand new chair.In the United States, we initially notified our Comfort Company who is our only United States Sales Representative. Comfort Company then notified U.S. Consignees via a letter which is provided for your reference. Also a letter was sent to our biggest customer in the USA, which is Cair Medical. It was sent on December 12, 2012 (PDG Letter to Cair Medical Fuze recall) to explain and outline how PDG handled the Fuze hex bolts breaking. A copy of this letter is also provided for your reference.Our Hong Kong customers were informed thru e-mail. Non- respondents were re- contacted by either e-mail or telephone.Effectiveness checks were done thru telephone. Defective product on hand at our facility has been scrapped and destroyed. Defective product that was removed from wheelchairs in the field was destroyed at the time the bolts were replaced. |  | 153 | 153 | Worldwide Distribution- US Distribution including the states of : AZ, CA, GA, KY, MI, MN, MO, MS, NC, NE, NY, OH, PA, SC, TX, VA, and WA., and Internationally to Canada and Hong Kong. | Terminated | July 23, 2015 | 1136 | 153 | Undetermined
Z-2696-2014 | 68521 | Yes | MSD | Blood access device and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | UltraStream Chronic Hemodialysis Catheter Standard Kit; Model Number 32101524 (24 cm), 32101528 (28cm), 32101532 (32 cm), and 32101536 (36 cm) and 32102540 (40 cm). The Rex Medical UltraStream Chronic Hemodialysis Catheter is a chronic hemodialysis catheter that can achieve high flow rates at low arterial pressure. |  |  | June 05, 2014 | 2014 | Argon Medical Devices, Inc | UltraStream Chronic Hemodialysis Catheter may develop cracks on the red (arterial) luer during attachment or manipulation. | Component design/selection |  |  | 3783 | 3,783 units | Nationwide Distribution. | Terminated | April 08, 2015 | 307 | 3783 | Undetermined
Z-2707-2014 | 69223 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Spacelabs Medical Patient Monitor is a Physiological, Patient monitor With Arrhythmia Detection or Alarms. The Models are 91367, 91369, 91370, 91387, and 91393 are sold with Perioperative Mode (option D), which includes Start case/End case functionality, Upgrade Kit 040-1548-00. |  |  | August 26, 2014 | 2014 | Spacelabs Healthcare Inc | The clinician did not realize the monitor was in END CASE mode and was expecting audible alarms. Visual alarms continue to work normally.The firm performed a retrospective analysis of all recalls since 2011 and discovered 222 additional serial numbers were missed and were not included in the previous recall under RES 59456 (recall number Z-3032-2011). | Environmental control | Spacelabs sent an "URGENT - MEDICAL DEVICE CORRECTION (Expanded)" letter dated 5 September 2014 to all consignees in the US. This letter identified recalled model 91367, 91369, 91370, 91387, and 91393, Upgrade Kit 040-1548-00, Equipped with the Perioperative Option (-D), problem, and actions to be taken.On September 12, 2014, a recall notification letter (translated asnecessary) will be emailed to all international subsidiaries and distributors of record.The notification letter stated that the Operating Room staff may not be aware that alarm tones are off if the device has been left in the END CASE state. If the customers make use of the Perioperative START CASE / END CASE feature, they should advise the clinicians to familiarize themselves with this feature as described in the Perioperative section of the Operations Manual. In particular,they should know how to confirm the state of the alarm tones prior to starting a case.The firm's Field Service Engineers will contact customers to schedule a time to come out to update affected monitors at no cost.Customers have any questions about this corrective action program, please contact Spacelabs Healthcare at 1-800-522-7025. |  | 222 | 82 units distributed in US and 140 units distributed outside US | Worldwide Distribution- US (nationwide) including the states of: FL, GA, IA, KY, LA, MA, MI, MN, MO, MS, NJ, NY, OH, OK, PA, SC, TN, TX, WA, Guam and Puerto Rico and to the following countries: Argentina, Australia, Canada, China, Czech Republic, Finland, France, Germany, Greece, India, Israel, Italy, Netherlands, New Zealand, Panama, Paraguay, Poland, Singapore, Spain, Sri Lanka, Switzerland, Tunisia, United Kingdom, and Uruguay. | Terminated | June 26, 2015 | 304 | 222 | Undetermined
Z-2706-2014 | 69116 | No | BTT | Respiratory gas humidifier. | Anesthesiology | Anesthesiology | 510(k) | Vapotherm Precision Flow,Disposable Patient Circuit |  |  | August 13, 2014 | 2014 | Vapotherm, Inc. | Water may leak into the center gas lumen of the Patient Circuit | Process control | Vapotherm issued an Initial Notification on August 13, 2014. The letter was updated on August 20, 2014 to include two additional lots. The recall letter notified users that this defect is easily detectable during normal use and theprobability of serious injury or death is remote. The letter was sent to raise awareness of the recent increase in incidence rate of this issue and reminded users of the requirement in the operating manual that requires that the cannula be connected to the patient only after the set temperature is reached. The letter also advised patient care personnel not to use, or to discontinue use of any Disposable Patient Circuit if water is observed leaking into the center lumen or from the cannula. They were instructed to replace the device and report the defect to Vapotherm immediately for further investigation.Should you have any questions please contact our technical support team atEmail: ts@vtherm.com(US) 855.557.8276(INTL) 1-603.658.5121 |  | 207 | 207 units | Worldwide Distribution - USA (nationwide) and Internationally to Canada, France, Swizerland, Korea, Sweden, Germany , Belgium, Spain, Mexico, Scotland, UAE, Australia, Israel, Hong Kong, Ireland, kuwait, and Italy. | Terminated | October 02, 2015 | 415 | 48329 | Undetermined
Z-2712-2014 | 68932 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | The Symbia T Series consists of the T16, T6, T2, and T. The Symbia T integrates state-of-the-art SPECT and high quality spiral CT to provide precise attenuation correction and anatomical mapping. The system can also be used as a complete clinical SPECT system for general purpose, WB and SPECT applications. |  |  | July 28, 2014 | 2014 | Siemens Medical Solutions USA, Inc. | Siemens Medical Solutions USA, Inc. is performing a field correction because we recently detected a rubber isolator with unfavorable aging properties. The function of the rubber isolator, which is situated within the rotate motor mounting assembly, may degrade over time and may impact motor support. The event does not pose a hazard by itself but could lead to a rupture of the drive belt which in t | Component design/selection | Siemens sent an Undated Customer Safety Advisory Notice letter to all direct accounts (customers). The letters included instructions for customers to: 1) continue to use the system(s) while waiting for the field correction to be scheduled and performed unless the system exhibits abnormal noise or vibration in which case the system(s) should not be used and the local Siemens Medical Solutions USA service engineer should be contacted; 2) ensure that the letter is placed in the Systems Owner's Manual and share it with support staff; and, 3) if the product was sold to another party, forward the letter to that party. A local service engineer will contact the customer to schedule the servicing of the affected systems by December 2014 which will consist of replacing the mounting plate. If customers have any questions about this issue, customer can contact their local service representative by calling 800-888-7436. |  | 1039 | 1039 systems | Worldwide Distribution - USA including AL, AK, AR, AZ, CA, CO, CT, FL, GA, IA, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, NC, NH, NJ, NM, NV, NY, OH, OR, PA, SC, TN, TX, UT, VA, VT, WAM WI, WY, District of Columbia, Puerto Rico and Internationally to Algeria,Argentina,Australia,Austria,Bangladesh,Belgium,Brazil,Bulgaria,Canada,Chile,China, Colombia,Croatia,Czech Republic,Denmark,Ecuador,Egypt,Finland,France,Germany,Hong Kong,Hungary,India,Indonesia,Iran,Ireland,Israel,Italy,Japan,Kazakhstan,Korea, Kuwait,Latvia,Libya,Lithuania,Malaysia,Mexico,Morocco,Myanmar,Netherlands,New Zealand,Norway,Oman,Pakistan,Peru,Philippines,Poland,Portugal,Qatar,Romania,Russia,Saudi Arabia,Singapore,Slovakia,Slovenia,South Africa,Spain,Sweden,Switzerland,Taiwan,Thailand,Tunisia,Turkey,United Arab Emirates,Ukraine,United Kingdom,Vietnam | Open, Classified |  |  | 1087 | Undetermined
Z-2651-2014 | 69112 | No | GIF | Blood cell diluent. | Hematology | Hematology | 510(K) Exempt | Coulter DxH Diluent, Catalog No. 628017 |  |  | August 25, 2014 | 2014 | Beckman Coulter Inc. | Beckman Coulter is recalling Coulter DxH Diluent because it may be associated with elevated platelet (PLT) background counts between 3-7 x 10(3) cells/uL. This will result in failed Daily Checks on the UniCel DxH 800 and DxH 600 Coulter Cellular Analysis Systems. | Under Investigation by firm | Beckman Coulter sent an Urgent Medical Device Recall letter dated August 25, 2014, to all customers that purchased the Coulter DxH Diluent. The letter informed the customers of the problems identified and the actions to be taken. Customers were instructed to complete and return the enclosed response form within 10 days. Customers with questions were instructed to contact Beckman Coulter's Customer Support Center at http://www.beckmancoulter.com, or call (800) 526-7694 in the US and Canada. |  | 46050 | 46,050 units total (40,078 units in US) | Worldwide Distribution - USA (nationwide) and Internationally to Canada, Chile, Mexico, Panama, and Thailand.. | Terminated | February 25, 2015 | 184 | 46050 | Undetermined
Z-2650-2014 | 69070 | No | KFF | Heparin assay. | Hematology | Hematology | 510(k) | Stago IVD, Liquid Anti-Xa, NDC 3607450003116. |  |  | July 18, 2014 | 2014 | Diagnostica Stago, Inc. | Diagnostica Stago Inc. received some customer complaints regarding a misreading of Reagent 1 (substrate) barcode on STA analyzers. | Under Investigation by firm | Diagnostica Stago sent a "Quality Information" letter/Customer Acknowledgement form to its customers. The letter identified the product, problem, and actions to be taken by the customers. For additional information and further support contact the Diagnostica Stago Hotline at 800-725-0607. |  | 514 | 514 kits | Distributed USA (nationwide) and the country of Canada. | Terminated | January 16, 2015 | 182 | 514 | Undetermined
Z-2648-2014 | 69136 | No | GDF | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Bard DuaLok Breast Lesion Localization Wire, |  |  | August 26, 2014 | 2014 | Bard Peripheral Vascular Inc | Bard Peripheral is recalling the Bard DuaLok Breast Lesion Localization Wire because it may be at risk of having portions of the wire protruding from the packaging, thus representing a breach of sterile barrier. | Other | A customer notification letter was sent on 8/27/14 to all customers who purchased the Bard DuaLok Breast Lesion Localization Wire. The letter informs the customers of the problems identified and the actions to be taken. Customers are instructed to call Bard Peripheral's Recall Coordinator at (800) 321-4254 option#2 Ex 2727. Customers are instructed to fax the completed recall and effectiveness check form to BPV at (800) 994-6772. |  | 177760 | 177,760 units | Worldwide Distribution-USA (nationwide), Australia, Canada, and European, Latin American and Asian Pacific countries. | Terminated | March 06, 2015 | 192 | 177760 | Undetermined
Z-2643-2014 | 68871 | No | JIX | Calibrator. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | ADVIA Chemistry A1c_3 CAL (Calibrator) for use on the ADVIA Chemistry Systems, REF/Catalog Number/Siemens Material Number (SMN) 10491408, IVD |  |  | August 06, 2014 | 2014 | Siemens Healthcare Diagnostics | The ADVIA Chemistry Systems A1c_3 method may exhibit a percent bias of -9% to +11% for hemoglobin (Hb)A1c samples when using A1c_3 Calibrator lots 3HD044 and 3LD068. Depending on quality control limits, this issue may not have been detected. NOTE: The percent bias range is a percentage of the HbA1c result. It is NOT absolute HbA1c units. | Under Investigation by firm | Siemens sent an Urgent Field Safety Notice dated August 2014 to all affected consignees on August 6, 2014 via Federal Express. The letter instructed consignees to discontinue use and discard calibrator lots 3HD044 and 3LD068. Customers were requested to complete and return the Effectiveness Check questionnaire attached to the Urgent Medical Device Correction letter within thirty (30) days. Distribution of the Urgent Field Safety Notices to foreign consignees was determined at the country level.For questions regarding this recall call 914-524-2955. |  | 1358 | Domestic: 1122 kits; Foreign: 236 kits | Worldwide Distribution - USA (nationwide) and Internationally to Argentina, Austria, Brazil, Canada, China, Colombia, Croatia, Cyprus, Denmark, Estonia, Finland, France, Germany, Greece, Guatemala, India, Indonesia, Istanbul, Italy, Lithuania, Mexico, Paraguay, Peru, Poland, Romania, Russian Federation, Serbia, Singapore, Slovakia, South Korea, Spain, Switzerland, Thailand, Uruguay, Vietnam, and Zambia. | Open, Classified |  |  | 1358 | Undetermined
Z-2565-2014 | 68970 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | Medsolutions4U Multi-Function Electrodes to be used with PhysioControl Lifepak Defibrillators, models 9, 10C, 11, 12, 15, 20, 500 and are indicated for Defibrillation, Cardioversion, Pacing, and Monitoring |  |  | March 12, 2013 | 2013 | Bio-Detek, Inc. | Medsoultions4U Multi-Function Lifepak Electrodes, manufactured with an incorrect wire connector which could not interface with the defibrillators | Nonconforming Material/Component | Bio-Detek issued an Urgent Field Safety Notice to Medsolutions4 U on 3/12/13, informing them of the problem and actions that should be taken to remove the affected products from service. |  | 30 | 30 cases (300 electordes) | Worldwide Distribution in Germany and Slovenia. | Terminated | June 17, 2015 | 827 | 30 | Undetermined
Z-2573-2014 | 69033 | No | KWP | Spinal interlaminal fixation orthosis. | Orthopedic | Orthopedic | 510(k) | STREAMLINE CT CERVICO-THORACIC SYSTEM; SET SCREW |  |  | September 10, 2012 | 2012 | PIONEER SURGICAL TECHNOLOGY, INC. | Some set screw thread profiles were found not to be within specifications. | Nonconforming Material/Component | The recall took place at the distributor/sales representative level. Communication was done through phone and email notifications on 9/10/2012.All affected set screws were to be traded out through a scheduled process. This was coordinated through Pioneer Surgical's customer service. Product was scrapped and no reconditioning took place. All affected set screws were swapped out. |  | 11 | 11 sets | Distributed in the states of NC, NH, NY, CA, MI, IL, and FL. | Terminated | September 30, 2014 | 750 | 11 | Undetermined
Z-2571-2014 | 69138 | No | ITX | Diagnostic ultrasonic transducer. | Radiology | Radiology | 510(k) | V5Ms Transesophageal transducers; |  |  | August 13, 2014 | 2014 | Siemens Medical Solutions USA, Inc. | Reports of deterioration of material covering the articulating section of the V5Ms transesophageal transducer. | Under Investigation by firm | Siemens sent a Customer Safety Advisory notification to all affected customers on August 13, 2014. The notification identified the product, the problem, and the action to be taken by the customer. Customers were advised that if they notice any damage or wear on their transducer, or if the transducer fails to pass the leakage current test, immediately discontinue use of the transducer and contact their local Customer Service Engineer. If they have anyquestions, customers should contact their local service support person for information regarding timelines and status.Customers were instructed to share this information with all personnel within their organization who need to be aware of this issue.For questions regarding this recall call 650-417-8185. |  | 321 | 321 units | Worldwide Distribution - USA (nationwide) and Internationally tAustralia, Austria, Belgium, Bolivia, Brazil, Canada, China, Columbia, France, Germany, Great Britain, India, Ireland, Italy, Japan, Mexico, Netherlands, Norway, Peru, Poland, Portugal, Russian Federation, Saudi Arabia, Slovak Republic, South Africa, South Korea, Spain, Sweden, Taiwan, Thailand,and United Arab Emirates. | Terminated | March 06, 2015 | 205 | 321 | Undetermined
Z-2572-2014 | 69087 | Yes | LWS |  |  | Cardiovascular | PMA | Ellipse Implantable Cardioverter Defibrillators (ICDs), Model No. CD1275, CD1311, CD2275, CD2311, CD1411, CD2411 |  |  | August 19, 2014 | 2014 | St Jude Medical Cardiac Rhythm Management Division | An anomaly may result in delayed, partial, or no delivery of high voltage therapy. Internal damage to the high voltage capacitors, which may result in extended charge time, occurring during capacitor maintenance or charging for high voltage therapy. It may present as a vibratory patient notification and upon interrogation, a programmer or Merlin.net alert. | Component design/selection | A customer letter dated 8/19/14 was sent to physicians with the affected Ellipse ICDs on 8/20/14. The letter identifies the problems and the actions to be taken.Customers with any questions or concerns are instructed to contact their local St. Jude Medical representative or St. Jude Medicals Technical Services Department. |  | 51492 | 51,492 units (27,396 units in US) | Worldwide Distribution: US (nationwide), Canada, Australia, New Zealand, Japan, EU, Middle East, Latin America, and Asia-Pacific. | Terminated | October 08, 2015 | 415 | 51492 | Undetermined
Z-2588-2014 | 68939 | No | DYB | Catheter introducer. | Cardiovascular | Cardiovascular | 510(k) | GALT VALVED TEARAWAY INTRODUCER SET GALT VTI RX ONLY UNITS 7 STERILE EO |  |  | July 25, 2014 | 2014 | Galt Medical Corporation | The adhesive used in the product could dislodge when objects are inserted through the valve which could pose an embolic risk to the patient. | Nonconforming Material/Component | Galt Medical sent an Urgent Medical Device Recall letters to all affected customers. The letter identified the affected product, reason for recall and actions to be taken. Customers were asked to identify and segregate the recalled lot (s) that is in their possession, complete and return the enclosed Recall Reply form, and return the recalled product to Galt Medical Corp. For questions contact your local sales representative or Recall Coordinator at 214-778-1306. |  | 11155 | 11,155 units in total | Worldwide Distribution - US (nationwide) Canada, Germany, the Netherlands, Ireland, Switzerland, France, Italy, and Greece. | Terminated | February 26, 2016 | 581 | 11155 | Undetermined
Z-2590-2014 | 69021 | No | HMX | Manual ophthalmic surgical instrument. | Ophthalmic | Ophthalmic | 510(K) Exempt | iTRACK250A Canaloplasty Microcatheter Kit; Catalogue number: iT-250A; Rx only, sterile. |  |  | August 06, 2014 | 2014 | Ellex iScience, Inc. | One lot may not have been properly sealed, resulting in a non-sterile device. | Under Investigation by firm | Urgent Medical Device Recall letters were sent to all customers on August 6, 2014 by trackable mail. Letters advised on the potential lack of sterility of iTRACK250A, Canaloplasty Microcatheter Kit, lot 1404-03 due to improperly sealed pouches. Lack of sterility may pose the risk of infection to the patient. Customers are identify and segregate any affected product in their inventory. The affected product should be returned following the instructions provided. Replacement product will be provided at no charge. Questions should be directed to Customer Care at 1-888-846-4724, M-F 8 am to 4:30 Pm Eastern Time. |  | 148 | 148 devices | Worldwide Distribution -- US (nationwide) and countries of: Hong Kong and Switzerland. | Terminated | October 28, 2014 | 83 | 148 | Undetermined
Z-2594-2014 | 68915 | No | NHA | Endosseous dental implant abutment. | Dental | Dental | 510(k) | BIOMET 3i Certain Bellatek Encode Healing Abutment |  |  | June 25, 2014 | 2014 | Biomet 3i, LLC | Identifiers on the label are correct but the label has the incorrect color coding. | Mixed-up of materials/components | Biomet 3i sent an Urgent - Medical Device Removal letter / and by phone call on and dated June 25, 2014, to all affected customers. The letter identified the product the problem and the action to be taken by the customer.If you have any Certain BellaTek Encode Healing Abutment(s) of the affected lot in your practice, please return the abutment(s) to BIOMET 3i in exchange for Cettain BellaTek Encode Healing Abutments, IEHA458 with a proper color stripe. Alternatively, you may return the Certain BellaTek Healing Abutment(s) for full credit.We also ask that you confirm receipt of this notification by completing the attached Recall Return Response Form, indicating your preference for replacement or credit and/ or notifying us if the product was used.Thank you for your support - we are committed to maintaining high standards of quality for all our products and apologize for any inconvenience this may have caused you.If you have any further questions, please contact Customer Service at 1-800-342-5454 for assistance. |  | 6 | 6 units | US Distribution to NH and TX. | Terminated | September 23, 2014 | 90 | 6 | Undetermined
Z-2598-2014 | 69119 | No | NRY | Percutaneous catheter. | Cardiovascular | Neurology | 510(k) | Merci Retriever |  |  | June 25, 2014 | 2014 | Concentric Medical Inc | Merci Retrievers were shipped to US customers past their expiration date, even though they were labeled with the correct expiration date. | Process control | Concentric sent an Urgent Medical Device Recall Notification dated June 25, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Letters were sent to Risk Manager- Director of Material and requested that the subject devices be quarantined and that Concentric be notified if the product has been distributed to other organizations. If so Concentric should be notified so that the additional customers can be notified.The response form should be completed and returned to Neuro Complaints @Stryuker.com. upon receipt of the response form, a representative will contact the firm to arrange for return and/or credit for the devices. Fax it to 1-866-876-4355. For further questions please call ( 650) 810-1707 Ext. 1716. |  | 9 | 9 devices | US Distribution including the states of PA, CA, TX, NM, and OH. | Terminated | October 29, 2014 | 126 | 9 | Undetermined
Z-2610-2014 | 69109 | No | GAG | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | PROXIMATE Skin Stapler. PROXIMATE RH Rotating-Head Skin Stapler (Model PRR35), PROXIMATE PX Fixed-Head Skin Stapler (Model PXR35), Bulk/Non-Sterile PROXIMATE RH Rotating-Head Skin Stapler (Model CK 117), and Bulk/Non-Sterile PROXIMATE PX Fixed-Head Skin Stapler (Model CK 158). Skin stapler for use in a single patient, which is designed to apply rectangular stainless steel staples for routine wound closure. |  |  | August 12, 2014 | 2014 | Ethicon Endo-Surgery Inc | The device may have a non-conforming component that may cause the device to fire an unformed staple. | Nonconforming Material/Component | The firm, Ethicon Endo-Surgery, sent an "URGENT: MEDICAL DEVICE RECALL" letter dated August 25, 2014 to its customers via overnight, UPS to allow tracking of the receipt. The letter described the product, problem and actions to be taken. The customers were instructed to immediately examine your inventory to determine if you have recalled product on hand; remove the recalled product and communicate the issue to relevent staff; complete and return Business Reply Form (BRF) confirming receipt of this notice within three (3) business days, indicating whether they have affected product, how much and whether they will/have returned the product or whether they no longer have any of the recall product. The BRF may be sent to Stericycle by fax at 1-855-328-6659 or by email at ethicon3022@stericycle.com. Product replacement is available for customers who return unopened affected product.If you need clinical or product support, please contact your local sales representative or call our Customer Support Center at 1-877-ETHICON (1-877-384-4266). If you need additional shipping labels or a communications package, contact Stericycle at 1-855-328-0398 and reference Event 3022. |  | 42798 | 42,798 | Worldwide Distribution: US (nationwide) including states of: AK, AL, AR, AZ, CA, CO, CT, DC, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MD, ME, MI, MN, MO, MS, MT, NC, NE, NJ, NM, NV, NY, OH, OK, OR, PA, SC, SD, TN, TX, UT, VA, WA, WI, WV, WY; and countries of: Armenia, Belgium, Germany, Denmark, Finland, France, UK, Italy, Japan, South Korea, Netherlands, Norway, Russia, Sweden, Singapore, Taiwan, Tanzania, and South Africa. | Terminated | July 10, 2015 | 332 | 42798 | Undetermined
Z-2617-2014 | 68877 | No | CCL | Oxygen gas analyzer. | Anesthesiology | Anesthesiology | 510(k) | CARESCAPE Respiratory Modules E-sCO, E-sCOV, E-sCAiO, E-sCAiOV, E-sCAiOE, E-sCAiOVE; the Airway Gas Option N-CAiO and respective service exchange units. |  |  | August 08, 2014 | 2014 | GE Healthcare, LLC | GE Healthcare has recently become aware of a potential safety issue due to a failure of an O2 sensor component associated with the CARESCAPE respiratory modules, the Airway Gas Option and their respective service exchange units. | Nonconforming Material/Component | Consignees were sent on 8/8/2014 a GE Healthcare "Urgent Medical Device Correction" letter Ref# 36106 dated August 8, 2014. The letter was addressed to Healthcare Administrator / Risk Manger, Chief of Nursing and Director of Biomedical Engineering. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction and Contact Information. For questions you can contact Technical Support at 1-800-558-7044 or your local Service Representative. |  | 748 | 748 (307 units US, 441 units OUS) | US: CA, CO, FL, GA, IN, KY, LA, MD,MI, NY,NC,OH, PA, SC, TN, TX, VA,WI.OUS: UNITED KINGDOM, TURKEY, THAILAND, SWITZERLAND, SWEDEN, SPAIN, SOUTH AFRICA, SINGAPORE, SAUDI ARABIA, PHILIPPINES, PANAMA, PAKISTAN, NORWAY, NICARAGUA, NEW ZEALAND, NETHERLANDS, MEXICO, LEBANON, KUWAIT, REPUBLIC OF KOREA, JAPAN, ITALY, INDONESIA, INDIA, GERMANY, FRANCE, FINLAND, EGYPT, CZECH REPUBLIC, COLOMBIA, CHILE, BULGARIA, BELGIUM, AZERBAIJAN, AUSTRALIA. | Terminated | September 08, 2015 | 396 | 748 | Undetermined
Z-2596-2014 | 68805 | cover, mattress (medical purposes) | FMW | Mattress cover for medical purposes. | General Hospital | General Hospital | 510(K) Exempt | Prime Care Âż Transcend Mattress, Prime Care Âż Transcend Perimeter Mattress, Prime Care Âż Transcend Advanced Mattress, Prime Care Âż Transcend Advanced Perimeter Mattress, Prime Plus Âż Bariatric Mattress and Prime Plus Âż Bariatric Perimeter Mattress, individual units packaged seperately. |  |  | August 11, 2014 | 2014 | Primus Medical LLC | The firm found that the top cover of the mattresses were delaminating. | Component design/selection | The firm sent out Urgent: Voluntary Field Correction/Removal letters, dated August 11, 2014, to all consignees with instructions for correcting the issue. Customers were to immediately examine their inventory and quarantine any affected product. Customers are to complete and return the response form provided. Questions can be directed to 1-877-638-2776. |  | 1557 | 1,557 units | Worldwide Distribution -- USA, including the states of TX, OH, PA, KS, FL, NV, IL, WA, MN, NJ, NC, UT, NY, CA, MD, CO. NE VA, and AZ; and, the country of Canada. | Terminated | July 19, 2017 | 1073 | 1557 | Undetermined
Z-2614-2014 | 68843 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | ***Baxter***BURETROL Solution Set***105" (2.7 m) |  |  | July 14, 2014 | 2014 | Baxter Healthcare Corp | Baxter Healthcare Corporation is voluntarily issuing a recall for specific lots of Buretrol Solution set product codes due to complaints for separation between the burette chamber and the drip chamber. Affected product codes are: 2C7564, 2C8864 and 2H8864. | Nonconforming Material/Component | Action to be taken if product was purchased directly from Baxter:Baxter is requesting that you take the following actions: 1. Locate and remove all affected product from your facility (the product code and lot number can be found on the individual product package or shipping carton).2. Contact Baxter Healthcare Center for Service to arrange for return and credit. The Center for Service can be reached at 888-229-0001 between the hours of 7:00 am and 6:00 pm Central Time. Please have your Baxter eight-digit ship-to account number ready when calling.3. Complete the enclosed customer reply form, and return it to Baxter by either fax Or scanned e-mail.4. If you are a dealer, wholesaler, or distributor/reseller that distributed any product to other facilities, please conduct a recall with your end-user customers in accordance with your customary procedures. Baxter distributed this product to customers between September 27, 2013, and April 25, 2014.Action to be taken if the products were purshased from a distributor or re-seller.1. Locate and remove all affected product from your facility (the product code and lot number can be found on the individual product package or shipping carton).2. Contact Baxter Healthcare Center for Service to arrange for return and credit. The Center for Service can be reached at 888-229-0001 between the hours of 7:00 am and 6:00 pm Central Time. Please have your ship-to account number ready when calling.3. Follow your supplier's reply and recall process. Please do not return the customer reply form to Baxter. |  | 107232 | 107232 | Nationwide Distribution. | Terminated | March 15, 2016 | 610 | 352612 | Undetermined
Z-2618-2014 | 68991 | No | NQX | Antimicrobial susceptibility test powder. | Microbiology | Microbiology | 510(k) | The BD MAX MRSA Assay, catalogue #442953. |  |  | July 23, 2014 | 2014 | Becton Dickinson &amp; Co. | The performance of the assay no longer conforms to this statement in the Analytical Specificity section of the Package Insert:One hundred-eleven (111) out of 111 MSSA strains tested at extremely high concentrations (&gt; 106 CFU/swab), produced negative results with the BD MAX MRSA Assay.One of the samples tested with each of the CAP surveys MRS5-A 2014 and MRS5-B 2014 yielded false positive results. | Under Investigation by firm | Becton Dickinson sent an Field Correction Notification initiated on July 23, 2014. US customers were contacted via letters sent by UPS. The letter identified the product the problem and the action needed to be taken by the customer.The acknowledgement form included with the follow up letter that is faxed back to BD will be reconciled with the customer list. BD will contact 10% of non-respondents to verify receipt of the customer letter. Any product returned to BD will be quarantined and discarded.Nonetheless, you may observe some false positive results for MSSA strains when using the above referenced product. Please contact the BD Technical Service Department at 1-800-638-8663 if you have observed an increase in MSSA false positive results or have failed the above referenced CAP surveys. |  | 585048 | 585,048 tests | US Distribution including the states of : CA, FL, HI, IL, IN, MD, MI, MN, MT, NC, ND, NJ, NM, NV, NY, OH, OK, OR, PA, SC, TN, TX, UT, VA, WA and WI. | Terminated | May 09, 2016 | 656 | 585048 | Undetermined
Z-2623-2014 | 69089 | No | DYE | Replacement heart valve. | Cardiovascular | Cardiovascular | PMA | Hancock Apical Left Ventricle Connector, Model No. H174Axx, Catalog No. 0174ACP012. |  |  | August 15, 2014 | 2014 | Medtronic Cardiovascular Surgery-the Heart Valve Division | Medtronic has decided to discontinue production and distribution of this low-volume product. | Component design/selection | An Urgent Medical Device Recall letter, dated August, 2014, was sent to customers on 8/15/14 for specific serial numbers of Hancock Apical Left Ventricle Connector, model H174Axx, devices. This recall is the result of a decision by Medtronic to discontinue sales of this low volume product, in lieu of fulfilling new regulatory requirements to maintain its distribution in the United States. Customers are instructed to quarantine and return all listed product in their inventory to Medtronic and contact customer service at (800) 854-3570. Customers are instructed to complete the attached customer confirmation certificate and e-mail it to RS.CFQFCA@Medtronic.com or fax to Medtronic at (651) 367-0612 to the attention of Customer Focused Quality. |  | 2 | 2 units | Distributed in the states of MD and TX. | Terminated | November 06, 2014 | 83 | 2 | Undetermined
Z-2624-2014 | 68954 | No | KCT | Sterilization wrap. | General Hospital | General Hospital | 510(k) | Integra Container Lid, ÂżÂż Size, Red; Container Lid, ÂżÂż Size, Blue; Container Lid, ÂżÂż Size, Red; Container Lid, ÂżÂż Size, |  |  | July 30, 2014 | 2014 | Integra York PA, Inc | Integra LifeSciences has identified through an internal investigation that some colored lids for Integra containersystems were not yet cleared by the FDA for sale and were inadvertently distributed. | Employee error | Integra sent an Urgent Voluntary Medical Device Recall letter dated July 30, 2014, to the affected customer. The letter identified the product, the problem, and the action to be taken by the customer. The customer was instricted if they have distributed the affected product, they should identify those customers adn forward a copy of the Integra Medical Device Recall Notice to them. The customer was also instructed to complete the attached form and return by email FCA@integralife.come or FAX to 1-609-275-9445. For questions the customer was asked to contact Customer Service at 1-855-532-1723.For questions regarding this recall call 717-840-9335. |  | 84 | 84 | US only to one repacker/relabeler and three medical facilities. The relabeler was identified as: KLS Martin, 11201 St. Johns Industrial Parkway S, Jacksonville, FL 32246 (904) 641-7746 | Terminated | March 10, 2015 | 223 | 84 | Undetermined
Z-2628-2014 | 69062 | No | JWY | Antimicrobial susceptibility test powder. | Microbiology | Microbiology | 510(k) | BD BBL Vancomycin Screen Agar Catalog number 222204 |  |  | July 28, 2014 | 2014 | Becton Dickinson &amp; Co. | These lots of Agar may have exhibited breakthrough growth of vancomycin susceptible Enterococcus species control organisms. Agar that allows susceptible enterococci to grow appears as a falsely resistant culture. If a lab does not QC, the error would be identified as falsely resistant enterococci would have further AST workup. There could be a delay in the identification of false-resistance | Under Investigation by firm | BD Diagnostic System sent an Urgent Product Recall letter dated July 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Our records indicate you may have been shipped the above-referenced lot numbers of product between March 5, 2014 and June 5, 2014. Please discontinue use of these lot numbers and discard any remaining packages. If further assistance is needed, please contact BD Technical Services Department at 1-800-638-8663.The acknowledgement form included with the follow up letter that is faxed back to BD will be reconciled with the customer list. BD will contact 10% of non-respondents to verify receipt of the customer letter. BD has asked distributors to send their customer list to BD for follow up. |  | 28540 | 28540 | Worldwide Distribution: US (Nationwide) including the states of AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, TN, TX, UT, VA, VT, WA, WI and WV; and Internationally to the countries of Belgium, Canada, Colombia, India, Mexico, Singapore and Taiwan. | Terminated | January 07, 2016 | 528 | 28540 | Undetermined
Z-2630-2014 | 68585 | No | JKA | Blood specimen collection device. | Clinical Chemistry | Clinical Chemistry | 510(k) | BD Vacutainer EDTA Tubes (LAVENDER TOP TUBE 2.0ML) |  |  | June 23, 2014 | 2014 | Fresenius Medical Care Holdings, Inc. | Products were held at improper storage temperature | Storage | Fresenius Medical North America contacted customers via telephone on 6/23/14 by Customer Service and follow-up with formal letter notification, Urgent Recall by certified mail with signature confirmation and fax-back form. Customers instructed to examine their inventory to determine whether they have any of the affected Naturalyte Liquid Bicarbonate or BD Collection Tubes. If customers have the affected product, they will be instructed to contact FMC-RTG Customer Service to have the product replaced. On July 10, 2014 Fresenius issued a clarification letter to users to state only products from the identified lots that were delivered by RTG LLC to the specified facilities on May 5, 2014 are affected by this recall as the products were exposure to higher temperatures than recommended.If you have any additional questions, please contact your FMCNA Customer Service Team at 1-800-323-5188. |  | 26 | 26 boxes | Distributed in the states of LA, OK, and TX. | Open, Classified |  |  | 55 | Undetermined
Z-2642-2014 | 68997 | No | JCX | Powered suction pump. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Vac-Assist Suction Units (SUA01 Suction Aspirator Unit w/800 cc Canister). |  |  | August 03, 2014 | 2014 | EMG Technology Co., Ltd. | suction canister lids may have occluded ports. | Nonconforming Material/Component | A recall letter, dated August 6, 2014, was sent via email which identified the product, problem, and actions to be taken. |  | 1452 | 1452 | US Distribution in IL. | Terminated | August 25, 2015 | 387 | 1452 | Undetermined
Z-2567-2014 | 68627 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | C-Series High Energy Linear Accelerator - Models: Novalis Tx, Trilogy, Trilogy Tx, Clinac iX, Clinac 2100 C/D, Clinac 2300C/D, Clinac 21EX, Clinac 23EX with versions 7, 8, 9 and 6 MV configuration and maximum dose rate of at least 6000 MU/min only. |  |  | June 17, 2014 | 2014 | Varian Medical Systems, Inc. | Trend in reports of unexpected decrease in beam output in C-series High Energy Linear Accelerators for 6MV photon treatment mode. | Radiation Control for Health and Safety Act | Urgent Medical Device correction letters were sent on June 16, 2017 to all affected users by certified mail. |  | 3416 | 3416 | Worldwide Distribution-USA (nationwide) and the countries of Algeria, Argentina, Australia, Austria, Azerbaijan, Bangladesh, Byelarus, Belgium, Bosniak, Brazil;, Bulgaria, Burma, Canada, Chile, China, Columbia, Costa Rica, Czech Republic, Denmark, Dominican Republic, Ecuador, Estonia, Finland, France, Germany, hong Kong, Hungary, Iceland, India, Indonesia, Iraq, Ireland, Israel, Italy, Japan, Luxembourg, Macau, Macedonia, Malaysia, Mauritania, Mexico, Netherlands, New Zealand, Norway, Oman, Pakistan, Panama, Paraguay, Peru, Philippines, Poland, Portugal, Qatar, Reunion, Romania, Russia, Saudi Arabia,, Singapore, Slovakia, South Africa, South Korea, Spain, Sri Lanka, Sweden, Switzerland, Taiwan, Thailand, Turkey, Ukraine, United Arab Emirates, United Kingdom, and Venezuela. | Open, Classified |  |  | 3416 | Undetermined
Z-1470-2014 | 67788 | No | KLT | Neuroleptic drugs radioreceptor assay test system. | Toxicology | Toxicology | 510(k) | VITROS Chemistry Products CRBM Slides, REF 889 2382, IVD -- Ortho-Clinical Diagnostics, Inc. For in vitro diagnostic use only. VITROS Chemistry Products CRBM Slides quantitatively measure carbamazepine (CRBM) concentration in serum and plasma using VITROS 250/350/950/5,1 FS and 4600 Chemistry Systems and the VITROS 5600 Integrated System. |  |  | March 05, 2014 | 2014 | Ortho-Clinical Diagnostics | The firm identified a potential for biased carbamazepine (CRBM) results to be generated when using VITROS CRBM Slides, Lot 3920-0080-8403. | Under Investigation by firm | The firm, Ortho Clinical Diagnostics, sent an "URGENT PRODUCT CORRECTION NOTIFICATION" letter (Ref. CL14-086; dated March 5, 2014) via FedEx Overnight Courier and/or US Postal Service Priority Mail (for PO Boxes only) to its US consignees (OCD Direct, Drop-Ship and US Federal Government). In addition, OCD proactively communicated with the 3 affected customers via phone to inform them of the potential issue and expedite product replacement. The letter described the product, problem and actions to be taken. The customers were instructed to immediately discontinue using and discard all remaining inventory of VITROS CRBM Slides, Lot 3920-0800-8403; review previously reported results using this lot; discuss any concerns regarding previously reported results with their Laboratory Medical Director or with the requesting physician; complete the attached Confirmation of Receipt form no later than March 14, 2014; and forward the information in this notification if you have distributed this product outside of your facility. If you have any additional questions, please contact Customer Technical Services at 1-800-421-3311. |  | 20 | 20 units | US Distribution to states of: IN, MD, and TN. | Open, Classified |  |  | 20 | Undetermined
t | o | o |  | b | i | g | [18 empty] | Undetermined
Z-0994-2014 | 67406 | [6 empty] | Integra Cusa Excel Rx Only |  |  | December 05, 2013 | 2013 | Integra LifeSciences Corp. | The blue COAG button on the CUSA Excel CEM nosecone could stick in the 'activated' position. | Material/Component Contamination | Integra sent an Urgent: Medical Device Correction Letter/Acknowledgement Return Form dated December 5, 2013 via traceable courier to their customers. The notification: advises customers of the nature of the issue and how it could affect CEM Nosecone use, advises customers what to do if the issue occurs and requests customers complete and return the Acknowledgment Return Form. The form should be faxed to rc@integralife.com or 1-609-275-9445.Any questions can be directed to customer service at 855-532-1723. |  | 17253 | 17253 boxes | Worldide Distribution - USA Nationwide | Open, Classified |  |  | 17253 | Undetermined
Z-0973-2014 | 67363 | No | OXO | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Medtronic O-arm Imaging System Mobile X-ray system. |  |  | July 25, 2013 | 2013 | Medtronic Navigation, Inc. | Review of on-system electronic product labeling and Instructions for comply with applicable Use in response to a customer inquiry uncovered gaps in compliance FDA standards with the requirements for information to be provided to users in 21CFR 1020.30(h) | Radiation Control for Health and Safety Act | Medtronic Navigation Inc. issued Notiifcation on 3/4/14 and included thee errata sheet to all affected customers updating the missing and incorrect information and ensure that labeling is now in compliance with 21 CFR Part 1020.30(h).This will be followed by a software update that contains the same information on the errata sheet, which serves to correct the electronic version of the Instructions for Use.The paper errata sheet will be sent via traceable means and the software upgrade will be performed by Medtronic Service either through remote access or by site visit of Medtronic service personnel.For further questions please call (978) 698-6000. |  | 449 | 449 in US | USA (nationwide) Distribution | Open, Classified |  |  | 449 | Undetermined
Z-0804-2014 | 67125 | No | JAA | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Video Monitor Suspensions that may be on the following Fluoroscopic and Radiographic X-Ray systems: Precision 500D, Legacy, Prestilix, Monitrol, RFX, SFX, Tilt-C, Prestige VH, Prestige SI, and Prestige SII |  |  | November 08, 2013 | 2013 | GE Healthcare, LLC | Missing screws in the overhead Video Monitor Suspension. GE discovered that the 4 outer screws that hold the overhead Video Monitor Suspension to the carriage were missing since installation of the system. GE is not aware of any fall of the overhead monitor suspension. | Employee error | GE Healthcare Systems sent an " Urgent Medial Device Correction" letter dated November 7, 2013. The letter was addressed to Hospital Administrators / Risk Mangers, Managers of Radiology /Cardiology and Radiologists / Cardiologists. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction and Contact Information (USA 800-437-1171, Japan 0120-055-919, For other countries, please contact your local GE Healthcare Service Representative). |  | 5651 | 5651 | Worldwide Distribution - US Nationwide including Washington D.C., Guam, and Puerto Rico. Outside the US to include: Afghanistan, Algeria, Argentina, Australia, Austria, Azerbaijan, Belarus, Belgium, Bolivia, Bosnia and Herzegovina, Brazil, Bulgaria, Burkina faso, Cameroon, Canada, Chile, China, Colombia, Costa Rica, Cote D'Ivoire, Croatia, Czech Republic, Denmark, Dominican Republic, Ecuador, Egypt, Estonia, Finland, France, French Polynesia, Germany, Ghana, Greece, Guadeloupe, Guatemala, Guyana, Honduras, Hong Kong, Hungary, Iceland, Indonesia, Iran, Ireland, Israel, Italy, Jamaica, Japan, Jordan, Kazakhstan, Korea (Republic of), Kuwait, Lebanon, Libyan Arab Jamahiriya, Macedonia, Malaysia, Mali, Malta, Martinique, Mexico, Morocco, Netherlands, New Caledonia, New Zealand, Niger, Norway, Palestine, Panama, Paraguay, Peru, Philippines, Poland, Portugal, Romania, Russia, Saudi Arabia, Serbia, Slovakia, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, Tunisia, Turkey, Ukraine, United Arab Emirates, United Kingdom, Uruguay, Venezuela, Vietnam, Yemen. | Terminated | November 03, 2014 | 360 | 5651 | Undetermined
Z-1126-2014 | 67038 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | *** 1) List number 126970465 labeled in part: 12697-65 LifeShield LATEX-FREE HEMA; Y-TYPE BLOOD SET, Nonvented, 88 Inch with 210 Micron Filter Pump, Prepierced Injection Site and Secure Lock; Made in Costa Rica; Product Inquiries should be directed to Hospira, Inc., Lake Forest, IL 60045 USA; |  |  | October 25, 2013 | 2013 | Hospira Inc. | Several lots of the LifeShield Latex-Free HEMA Y-Type Blood Set IV Administration Set were assembled incorrectly. | Process design | The firm, Hospira, sent "URGENT MEDICAL DEVICE CORRECTION" letters dated October 25, 2013 to all direct accounts/customers.The letter described the product, problem and actions to be taken. The letters instructed customers to: discontinue use of the recalled product lots; search storage areas, supply carts, and other patient care areas for affected lots and to remove any discovered recalled products; quarantine all recalled products; complete and return the attached RESPONSE REQUIRED form via fax at 1-888-356-6430 or e-mail to: Hospira6107@stericycle.com; return affected products using the labels supplied with the letter; if you have not received a return label or reply form or require additional assistance; and, if the recalled products were further distributed, to forward the recall information to them and have them call Stericycle at 866-229-7775 (Monday - Friday, 8:00 AM - 5:00 PM EST) to receive a reply form. The letter further instructs customers to ensure that the facility's protocol(s) for administering fluids and blood products are followed and to report any backflow, inability or difficulty priming, or no flow issues to Hospira at 800-441-4100 (Monday - Friday, 8:00 AM - 5:00 PM CST) or via e-mail at ProductComplaintsPP@hospira.com. For further inquiries, please contact Hospira using the information provided: Reports of adverse events or product complaints can be made by contacting Hospira Global Complaint Management at 800-441-4100 (Monday - Friday, 8:00 AM - 5:00 PM CST) of via e-mail at ProductComplaintsPP@hospira.com. General medical inquiries can be made by contacting Hospira Medical Communications at 800-615-0187 (24 hours a day, 7 days a week) or via e-mail at medcom@hospira.com. This information was also posted to Hospira's Internet website. |  | 173689182 | *** 1) List Number 126970465: 21,860 units; *** 2) List Number 127200465 : 20,256 units; *** 3) List Number 046419701 : 26,900 units | Worldwide Distribution: US (nationwide) including states of: AL, AR, AZ, CA, FL, GA, IL, IN, KS, KY, LA, MI, MN, MO, MS, MT, NC, NE, NJ, NM, NV, NY, OH, OR, PA, TN, TX, VA, WI, WV and Puerto Rico; and countries of Canada, Espana, France, Ireland, and Slovenia. | Open, Classified |  |  | 173689182 | Undetermined
Z-0972-2014 | 67227 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | TSX-301C Aquilion One (Vision Edition) Computed Tomography X-Ray System. |  |  | December 11, 2013 | 2013 | Toshiba American Medical Systems Inc | Toshiba America Medical Systems, Inc. initiated this recall due to the failure of the parts used in the X-ray high-voltage generator. In some cases, X-ray exposure cannot be performed. | Other | Toshiba America Medical Systems, Inc. initiated this recall by sending "Urgent Medical Device Correction," letter dated December 11, 2013, informed customers of the recall by providing, the device name, serial number, problem statement, corrective action statement, and contact information. A customer reply form accompanied the recall notification letter.Strategy:The majority of the letters are sent via USPS return receipt mail. The letter contains a return reply form that is to be faxed to TAMS for retention. After two weeks contact will be made with each non responding location and a copy of the letter will be faxed to that location and the fax receipt will serviceas proof of notification. Upon release of the new part it will be assigned to a TAMS employee with instructions for implementation at each site. Each incident is assigned a unique task number and each task is trackedthrough to completion. In the event that a customer refuses the installation of the corrected software/hardware the customer would be requested to sign a service case record stating that the corrected software/hardware was refused.CDRH approves the CAP subject to the following condition:Notification of all dealers and purchasers is to be made within 15 working days of receipt of this letter in the manner specified in 21 CFR 1003.21 and 1003.22. For further questions please call (714) 730-5000. |  | 8 | 8 units | US Distribution including the states of: MD, WV, OH, CA, MN, IN, NY and WI. | Terminated | October 16, 2014 | 309 | 8 | Undetermined
Z-1144-2014 | 67436 | [6 empty] | Simplexa Flu AlB &amp; RSV Direct Assay Kits, intended for use as an aid in the differential diagnosis of influenza A, influenza B, and RSV viral infections in humans and is not intended to detect influenza C. Model Number MOL2650. |  |  | February 04, 2014 | 2014 | Focus Diagnostics Inc | Focus Diagnostics initiated the recall of the certainSimplexa Flu AlB &amp; RSV Direct assay kits due tothe potential of false positive Flu A, Flu B and RSV signals. | Under Investigation by firm | Focus Diagnostic initiated this recall on February 4, 2014 by sending recall notification letters to customers via email. The letter dated February 4, 2014, titled "URGENT: SAFETY NOTICE REMOVAL", informed customers of the recall (removal). The letter provided the following information: product description with code, reason for recall, the issue, the risk to health, recommendation, background, actions by customers, contact information, labeling, and response sheet. |  | 2175 | 2,175 kits | Worldwide Distribution: US (nationwide) and to countries of:Canada, Sweden, Belgium, Spain, Portugal, Slovenia, and Israel. | Terminated | November 01, 2014 | 270 | 2175 | Undetermined
Z-1137-2014 | 67454 | No | OMP | Powered suction pump. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | PRE1055: Prevena Dressing, 5 Pack |  |  | January 30, 2014 | 2014 | KCI USA, INC. | Prevena Incisional Management System Kit/Dressing Connector - the male connector on the tubing is difficult to connect to the female connector on the Prevena canister tubing | Nonconforming Material/Component | KCL sent an Urgent Voluntary Medical Device Correction letter dated February 2014 to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to locate, quarantine and return the affected device. For a replacement of the affected product and questions customers were instructed to contact KCL Customer Technical Support at 1-800-275-4524. A Product Reconciliation Form is attached to be completed and returned. |  | 20850 | 20,850 units total | USA Nationwide Distribution | Terminated | April 06, 2015 | 431 | 20850 | Undetermined
Z-1146-2014 | 67332 | No | JIX | Calibrator. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | ABX PENTRA Multical. |  |  | January 20, 2014 | 2014 | Horiba Instruments, Inc dba Horiba Medical | HORIBA Medical is informing all ABX PENTRA MultiCal, N Control, and P Control customers running theCreatinine Rate Blank Method assay on the ABX PENTRA 400 and PC200 that the Target Values for Creatinine RateBlank Method have been updated. This issue is only applicable to Creatinine Rate Blank Method values reported on theABX PENTRA 400 and PC200 Chemistry Analyzers only. No other analyzers | Under Investigation by firm | HORIBA Medical initiated this recall by sending customers recall communication via FedEx week of January 20, 2014. The recall communication was a Field Safety Notice letter and Acknowledgement Form. The letter, dated 01/20/2014, titled "URGENT SAFETY NOTICE", informed customers of the recall by providing product description and codes, issue, impact, actiona/resolution, and contact information. The letter contained the following attachments: Field Safety Notice Acknowledgement Form and Updated Annexes with re-assigned values. |  | 2023 | 2,023 units | United States nationwide distribution. | Terminated | December 05, 2014 | 319 | 2023 | Undetermined
Z-1134-2014 | 67442 | No | HFX | Obstetric-gynecologic specialized manual instrument. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | DeRoyal ST. Circumcision Clamp, REF 32-1622, Size 1.3 cm, 1 Per Pack, Rx Only STERILE EO, distributed by DeRoyal Industries. |  |  | January 16, 2014 | 2014 | DeRoyal Industries Inc | Three lots of circumcision clamps were manufactured out of specification and as a result may not properly clamp at the surgical site completely. | Process control | DeRoyal initiated the recall by sending overnight letters to its direct accounts on 01/16/2014. If distributors inform DeRoyal of further distribution, DeRoyal will contact those consignees in the same manner. If necessary, further communications will be sent in the form of another mailing, phone calls, and emails to materials management personnel when possible. |  | 2504 | 2504 units | US Distribution to states of: NJ, AL, TX, WA, IL, OK, VA, FL, OR, and KS. | Terminated | October 13, 2015 | 635 | 2504 | Undetermined
Z-0799-2014 | 67178 | light, microscope | IBJ | Microscopes and accessories. | Pathology | Pathology | 510(K) Exempt | Compensator Slider Lambda 6x20 component (part number 473704-0000-000) utilized on all upright ZEISS microscopes (Axio Imager, Axio Scope. A1, Axio Lab. A1). |  |  | December 20, 2013 | 2013 | Zeiss, Carl Inc | Carl Zeiss Microscopy, LLC has recalled the Compensator Slider Lamda 6x20 component utilized on Standard Pathology Microscopes due to misalignment of the lambda plate. | Nonconforming Material/Component | Carl Zeiss, Inc. sent an Important Medical Device Corrective Action letter dated December 20, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were advised that their service provider would contact them to schedule a visit in which a factory trained service technician would install a new Compensator Slider Lambida 6x20. Customers should return the old Compensator Slider to the factory for testing. Customers with technical questions were instructed to call 914-681-7815. |  | 335 | 335 units Total (14 domestically &amp; 321 internationally) | Worldwide Distribution - Nationwide including NY, TX, NC, FL, ME, MA and MI; and Internationally to Arab Emirates, Australia, Austria, Bangladesh, Belgium, Brazil, Canada, China, Czech Republic, Egypt, France, Germany, Hong Kong, India, Italy, Kosovo, Lithuania, Mexico, The Netherlands, New Zealand, Peru, Poland, Spain, and Taiwan. | Terminated | April 08, 2015 | 474 | 335 | Undetermined
Z-1131-2014 | 67435 | No | IOL | Powered exercise equipment. | Physical Medicine | Physical Medicine | 510(K) Exempt | Accessory shorts for the AlterG Anti-Gravity treadmill. |  |  | January 28, 2014 | 2014 | AlterG, Incorporated | Unapproved material used by vendor in subset of shorts causing them to stretch more than usual and are uncomfortable for patient. | Component change control | The firm, AlterG, Inc, sent "Medical Device Recall Notification Anti-Gravity Short Exchange" letters dated February 6, 2014 to all customers beginning February 7, 2014 by Federal Express. The letter describes the product, problem and actions to be taken. The customers were instructed to review your inventory of shorts to identify any that are labeled ZIN1212 or ZIN0413 and return any shorts with these labeled production date using the return label provided.If you have any questions about this shorts exchange program or if there are any other ways we can be of assistance to you, please email us at support@alterg.com or call the support team at any time at 510-2705369. |  | 2221 | 2221 devices | Worldwide Distribution:US (nationwide) and to countries: Australia, Netherlands, United Kingdom, Spain, Germany, France, Denmark, Switzerland, Poland, Slovak Republic, Austria, and United Arab Emirates. | Terminated | May 13, 2014 | 105 | 2221 | Undetermined
Z-1125-2014 | 67498 | No | JDP | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Synthes 5.0mm Variable Locking Screw |  |  | February 05, 2014 | 2014 | Synthes, Inc. | Labeling correction: The package insert in one 5.0mm Variable Locking Screw package was a GP0006 (Anterior Cervical Vertebrae Plate System) insert, which should have been a GP2848 (4.5mm VA-LCP Curved Condylar Plate System insert). | Employee error | Synthes sent an Urgent Field Safety Notification/Medical Device Labeling Correction letter to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to:Exchage the previously shipped package insert (GP0006) with the correct package insert (GP2848 provided with the notification).Review the package insert.Forward the Field Safety Notification to anyone in their facility that needs to be informed.If the package insert has been forwarded to another facility, contact that facility.Maintain awareness of this notice until all package inserts have been exchanged.Maintain a copy of the notice.Customers with questions were instructed to maintain a copy of the notice.For questions regarding this recall call 610-719-5000. |  | 45 | 45 | Nationwide Distribution including GA, CA, TX, MO, OH, PA, WA, AL, IN, AZ, MI, IA, TN, KY, NC, VA, NY, LA, IL, UT, and KS. | Terminated | August 31, 2015 | 572 | 45 | Undetermined
Z-1127-2014 | 67349 | No | EHD | Extraoral source x-ray system. | Dental | Radiology | 510(k) | Model Kodak 2100 Intraoral X-Ray System, Catalog Numbers 5303144, 8011579 -- Carestream Health, Inc. 150 Verona Street Rochester, NY 14608 -- Made in France -- TROPHY 77435 Marne La Vallee --- COMMON/USUAL NAME: Dental X-Ray Systems -- CLASSIFICATION NAME: Unit, X-Ray, Extraoral with timer --- There are four major components: x-ray tube, x-ray generator, collimator and image receptor. These systems are permanently installed diagnostic dental x-ray systems for generation of x-rays for examination of the teeth, jaw and oral structures. |  |  | January 24, 2014 | 2014 | Carestream Health Inc. | Carestream received problem reports relating to apparent early failure of Kodak 2X00 Intraoral X-ray device arms leading to arm breakage. | Component design/selection | Urgent Medical Device Recall Letters (dated 1/24/2014) were sent to the consignees on 1/24/2014 via Certified Mail, Return Receipt Requested. |  | 347 | Domestic: 347 units (total for both) | Worldwide distribution. USA (nationwide) and the country of Canada. | Terminated | June 18, 2015 | 510 | 347 | Undetermined
Z-0807-2014 | 67085 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | SOLAR 8000 AND TRANSPORT PRO WITH PATIENT DATA MODULE (Marketed As GE Healthcare CARESCAPE Patient Data Module) Product Usage: The Patient Data Module (PDM) is intended to provide uninterrupted acquisition of physiologic parameter data on adult, pediatric and neonatal patients during non-transport/bedside and transport patient care episodes. |  |  | December 05, 2013 | 2013 | GE Healthcare, LLC | Intermittent noise on ECG and/or Respiration waveforms associated with movement of the ECG cable connector on the DASH, PDM, and/or Tram Module. May result in reduced ECG, Respiration and arrhythmia detection performance. | Pending | GE Healthcare sent an Urgent Medical Device Correction letter dated December 5, 2013, to all affected customers. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction and Contact Information. Customers with questions were instructed to contact Technical Support at 1-800-558-7044. For questions regarding this recall call 262-513-4122. |  | 2516 | 2516 (2,321 US, 195 OUS) | Worldwide Distribution: US (nationwide) including DC and PR except SD, VT and WY; and Internationally to: ARGENTINA, AUSTRALIA, AUSTRIA, BAHRAIN, BELGIUM, BRAZIL, CHINA, COLOMBIA, COSTA RICA, CROATIA, CZECH REPUBLIC, EGYPT, GERMANY, UNITED KINGDOM, IRELAND, IRAQ, ISRAEL, JAMAICA, JAPAN, JORDAN, KOREA, KUWAIT, LEBANON, LITHUANIA, MARITIUS, MALDOVA, NEW ZEALAND, NIGERIA, PAKISTAN, PANAMA, PARAGUAY, PERU, POLAND, QATAR, RUSSIA, SLOVAKIA, SOUTH AFRICA, SWEDEN, SWITZERLAND, TUNISIA, UNITED ARAB EMIRATES, VENEZUELA, CHILE, CANADA, and BERMUDA. | Open, Classified |  |  | 51589 | Undetermined
Z-0813-2014 | 67255 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | AMS-530 7 Microbore Trifurcated Extension sets |  |  | January 07, 2014 | 2014 | Churchill Medical Systems, Inc. | Defective component may cause extension sets to leak. | Nonconforming Material/Component | Vygon sent an Urgent Medical Device Recall letter dated January 7, 2014 to each customer via courier service and email. The letter identified the affected product, problem and actions to be taken. The recall letter instructs customers to: 1. Check your stock for the products included within the scope of the recall. Immediately cease use and distribution, and quarantine all affected product immediately. 2. Count and document your affected inventory on the Recall Acknowledgement and Inventory Return Form 3. Fax or email the form to Vygon including customer contact information. 4. Vygon customer service will contact the customer and issue a Return Merchandise Authorization number. 5. Return the product as well as the original signed Recall Acknowledgement and Inventory Return Form using the instructions from Customer Service involving product pick-up. 6. All inventory will be credited or replaced at no cost to the customer. Distributors were provided with instructions on how to handle product that was further distributed by them. For questions contact your local sales representative or Vygon's Customer Service Department at 1-800-473-5414, or by email at customerservice@vygonus.com. |  | 3850 | 3850 | US Nationwide Distribution in the states of: NY, VA, IL, MO, OR, FL, DE, OH | Terminated | November 03, 2014 | 300 | 3850 | Undetermined
Z-1123-2014 | 67385 | No | HQL | Intraocular lens. | Ophthalmic | Ophthalmic | PMA | AMO TECNIS CL Foldable Silicone IOL with OptiEdge Design, Model Z9002, diopter size: 16.0 and 21.0. |  |  | January 28, 2014 | 2014 | AMO Puerto Rico Manufacturing, Inc. | Certain Silicone TECNIS IOLs may be labeled with the incorrect diopter power. | Nonconforming Material/Component | Beginning on January 28, 2014, AMO began distribution of the Recall Notification Letter to the 72 AMO TECNIS IOL customer accounts within the United States who had received lenses from the 2 affected P/O's. 100% delivery of the recall communication to US customers was confirmed on January 31, 2014. Included with the Recall Notification Letter is a facsimile delivery confirmation which each customer is instructed to fax back to AMO. |  | 169 | 169 units | Worldwide Distribution-USA (nationwide) and the countries of Australia, Great Britain, Canada and Japan. | Terminated | April 19, 2017 | 1177 | 169 | Undetermined
Z-0816-2014 | 66875 | No | JGS | Sodium test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | ABL90 FLEX Analyzer Radiometer Medical ApS, Âżkandevej 21 |  |  | October 15, 2013 | 2013 | Radiometer America Inc | RADIOMETER become aware that if the ABL90 FLEX is tilted, this may potentially cause a negative bias on sodium (Na+) results. | Other | Radiometer sent an Urgent Medical Device Recall letter dated November 5, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to check all of their ABL90 FLEX analyzers and ensure that the analyzer is placed on a horizontal platform/table. If the analyzer is placed on a roll stand, please ensure that the support plate is aligned horizontally and locked in this position. Customers were asked to complete and fax the accompanying acknowledgment form in order to confirm their receipt of this field safety notice. Customers with questions were instructed to contact Radiometer America Technical Support at 1-800-736-0600, option 4.For questions regarding this recall call 440-871-8900. |  | 3414 | Total 3414 (US 524, CA 31, ROW 2859) | Worldwide Distribution - USA including AL, AZ, CA, CO, CT, DC, FL, GA, IA, ID, IL, IN, KY, LA, MA, MD, ME, MI, MN, MO, MS, NC, ND, NE, NH, NJ, NM, NY, OH, OK, OK, OR, PA, SC, TN, TX, UT, VA, VT, WA, WI &amp; WV and Internationally to Australia, Austria, Canada, China, Czech Republic, Denmark, Finland, France, Germany, Hungary, Italy, Japan, Netherlands, Norway, Poland, Russia, Singapore, South Africa, Spain, Sweden, Switzerland, Turkey &amp; United Kingdom. | Terminated | June 19, 2014 | 247 | 3414 | Undetermined
Z-1119-2014 | 67392 | No | HWA | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Greatbatch Medical, Offset Cup Impactor. Part Numbers: P9485, T10243, T10281, T10287, T10484, T10491, T10507, T10600, T10604, T10753, T10861, T11209, T11340, T12121, T12230, T12288, T12305, T12391, T12699, T12767, T12861, T13344, T13480, T13642, T13722, T13999, T14155, T14384, T15311, T15752, T15822, T15949, T16117, T16427, T16604, T16611, T16661, T16829, T16934, T17062, T17203, T17238, T17321, T6318, T7821, T8042, T8043, T8044, T8087, T8088, T8093, T8095, T8177, T8184, T8277, T8310, T8333, T8468, T8487, T9109, T9196, T9205, T9316, T9348, T9360, T9556, T9557, T9747, T9894, T9954, T9955, T9999. |  |  | January 31, 2014 | 2014 | Greatbatch Medical | Greatbatch Medical has initiated a global Field Corrective Action for the Standard Offset Cup Impactor since the product was first released in 2004.Greatbatch has identified through internal Sterility Assurance Level validation testing that a sterility assurance level of 10-6 is not achieved on the Standard Offset Cup Impactor when it is subjected to the steam sterilization cycles and associated dry times that are specified in the current Instructions For Use (IFU) provided with the device. The Standard Offset Cup Impactor is provided non-sterile and must be sterilized prior to use in a surgery. Greatbatch has not received any reports of deaths, illnesses, injuries, or other adverse effects associated with this issue. | Process design | Greatbatch sent an Urgent Medical Device Field Corrective Action letter dated January 31, 2014, to all affected consignees. The letter was addressed to "Dear Valued Customer". The letter described the Reason for Field Corrective Action, Risk to Health, and Instructions to Customers. The instructions to customers included to review the enclosed Field Corrective Action Response Form, to examine their inventory and identify all Standard Offset Cup Impactors that are in their possession and immediately inform those responsible for sterilization to utilize the updated instructions. Also, advised consignees to immediately notify all customers of the Field Corrective Action and to complete and return the Field Corrective Action Response Form. For questions consignees were instructed to contact their local Greatbatch Medical representative at 1-763-951-8235 or e-mail them at FieldActionCenter@Greatbatch.com.For questions regarding this recall call 763-951-8235.UPDATED 3-28-2014: An updated Urgent Medical Device Correction letter was sent out to consignees 3/28/2014 to tell them of an additional validated sterilization parameter in addition to additional devices that are affected by the recall. A notice was asked to be posted where sterilization of the Standard Offset Cup Impactors occurs so personnel responsible for the sterilization of the devices would see the 2 validated cycles. A response form was asked to be returned to Greatbatch. |  | 4870 | 4870 worldwide (2482 US), EXPANDED: 7245 devices | Worldwide Distribution - USA including CA, IN, KS, MA, NJ, TN,TX,UT and Internationally to Australia, Austria, Belgium, China, France, Germany, Greece, Ireland, Italy, Japan, Luxembourg, Netherlands, Norway, South Korea, Spain, Sweden, Switzerland, Taiwan, United Kingdom. | Terminated | October 31, 2016 | 1004 | 4870 | Undetermined
Z-1112-2014 | 67415 | No | JSK | Supplement for culture media. | Microbiology | Microbiology | 510(K) Exempt | Oxoid Legionella BCYE Growth Supplement, For 500 ml of medium, SR0110C, packaged 10 vials/box, lyophilized. The firm name on the label is Oxoid Ltd., Basingstoke, Hants, England. |  |  | January 24, 2014 | 2014 | Remel Inc | The product may contain high levels of microbial contamination | Unknown/Undetermined by firm | Thermo Fisher sent an Urgent Medical Device Recall letter dated January 24, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customers. Please notify any personnel who need to be aware of the potential for false presumptive Legionella spp. identification or masking of the recovery/isolation of Legionella spp. Review your inventory for the affected product and discard it. You should review results reported using these lots and consider retesting and/or seek appropriate expert advice. Return the attached Medical Device Recall Return Response acknowledging receipt of the notice and disposal of the affected product.If you have further distributed these product lots, please contact those entities, advise them of the situation and provide them with a copy of this letter. You should insert your contact information, email, and fax numbers in the acknowledgement form and request that they return it to you.If you have any questions, please contact our Technical Services Department at 800-255-6730 (US) or 913-888-0939 (International). |  | 440 | 44/10/500-ml. vial boxes | US Distribution including Puerto Rico and the states of FL, IL, MA, ME, MI, and PA. | Terminated | July 23, 2014 | 180 | 440 | Undetermined
Z-1118-2014 | 67487 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | Philips HeartStart MRx Monitor/Defibrillator; Models: |  |  | February 18, 2014 | 2014 | Philips Medical Systems, Inc. | ECG trunk cable and connector block of the MRx could be susceptible to accelerated wear, which could result in an interrupted ECG signal | Component design/selection | Philips Healthcare issued Urgent Medical Device Correction notification/Field Safety Notice on 2/18/14 informs customers of the issue and identifies what action Philips plans to take to remedy the issue.Philips will perform a hardware upgrade free of charge.Philips is asking customers to follow the Action to be Taken by Customer/User section of the Urgent Medical Device Correction notification/Field Safety Notice:You can continue to use your MRx prior to receiving the hardware upgrade, while remaining aware that if the problem were to occur you could experience the following symptoms:Loss of demand mode pacing, nability to perform synchronized cardioversion with paddle andDisruption of ECG monitoring or INOP messages.Not every serial number within the listed range is affected by this Medical Device Correction. If you have questions about whether your MRx is affected by this issue, please contact Philips at 1-800-722-9377. |  | 49929 | 49929 units | Worldwide Distribution-USA (nationwide) and the countries of AFGHANISTAN ALBANIA ARGENTINA ARUBA AUSTRALIA AUSTRIA AZERBAIJAN BAHRAIN BANGLADESH BELGIUM BOLIVIA BOSNIA AND HERZEGOWINA BRAZIL BRUNEI DARUSSALAM BULGARIA CANADA CHILE CHINA COLOMBIA COSTA RICA CROATIA (local name: Hrvatska) CYPRUS CZECH REPUBLIC DENMARK ECUADOR EGYPT ESTONIA ETHIOPIA FINLAND FRANCE GERMANY GREECE GUADELOUPE HONG KONG ICELAND INDIA INDONESIA IRELAND ISRAEL ITALY JAPAN JORDAN KENYA KOREA, REPUBLIC OF KUWAIT LAOS LATVIA LEBANON LIBYAN ARAB JAMAHIRIYA LITHUANIA LUXEMBOURG MACAU MACEDONIA MALAYSIA MALTA MEXICO MOROCCO NAMIBIA NETHERLANDS NETHERLANDS ANTILLES NEW CALEDONIA NEW ZEALAND NICARAGUA NIGERIA NORWAY OMAN PAKISTAN PANAMA PAPUA NEW GUINEA PERU PHILIPPINES POLAND PORTUGAL PUERTO RICO QATAR REUNION ROMANIA RUSSIAN FEDERATION SAUDI ARABIA SINGAPORE SLOVENIA SOUTH AFRICA SPAIN SRI LANKA SWEDEN SWITZERLAND TAIWAN, PROVINCE OF CHINA TANZANIA, UNITED REPUBLIC OF THAILAND TRINIDAD AND TOBAGO TUNISIA TURKEY UGANDA UKRAINE UNITED ARAB EMIRATES UNITED KINGDOM UZBEKISTAN VENEZUELA VIET NAM YEMEN. | Open, Classified |  |  | 49929 | Undetermined
Z-0788-2014 | 66493 | No | KGG |  |  | Neurology | PMA | TRUFILLÂż n-Butyl Cyanoacrylate (n-BCA) Liquid Embolic System; Two 1 g tubes of n-BCA, one 10 mL vial of Ethiodized Oil and one 1 g vial of Tantalum Powder |  |  | October 11, 2013 | 2013 | Codman &amp; Shurtleff, Inc. | Incorrect instructions for use (IFU). | Employee error | The firm, CODMAN Neuro, sent a "Medical Device Correction Notice" letter dated October 21, 2013 to its consignees/customers. The letter describes the product, problem and actions taken. The letter informed the consignees/customers of the incorrect statement in the IFU. The Corrected IFU was provided. Codman Neuro is in the process of updating the Instructions for Use to include the corrected information. Consignees/customers were instructed to complete and return the Correction Notice Acknowledgement Form and fax it to Codman Neuro at 508-977-6665.For additional information, please contact Codman Neuro representative at 1-866-491-0974. |  | 3860 | 3,860 | Worldwide Distribution: US (nationwide) including Puerto Rico; and internationally to: Costa Rica and Russia. | Terminated | July 28, 2015 | 655 | 11843 | Undetermined
Z-0812-2014 | 67164 | No | CAH | Breathing circuit bacterial filter. | Anesthesiology | General Hospital | 510(k) | Vital Signs Anesthesia Circuits |  |  | December 26, 2013 | 2013 | Vital Signs Colorado Inc. | A circuit in use was occluded by a blue port cap that was inadvertently attached to the circuit as a result of an internal inspection error and will cause occlusion of gas flow. | Process control | GE Healthcare sent an Urgent Medical Device Correction letter on December 12, 2013 to all affected customers via certified letter to follow the safety instructions to:1) Please follow all safety instructions in accordance with the instructions for use (IFU).2) During installation, visually inspect the Vital Signs Breathing Circuits for any plugged ports. If 15mm or 22mm blue port caps are attached to any ports, please remove and discard.3) Prior to use, the circuits should be pressure tested for leaks and flow tested for obstructions in accordance with the ventilator manufacturers specifications. The long-term solution is to update the instructions for theassembly operator and modify the circuit air test process by replacing the cap with a plug tethered to the workstation so the plug cannot be left on the circuit.If you have any questions or concerns regarding this notification, please contact GEMSIT Customer Service at +1-800-588-7044 (Domestic) or Vital Signs Customer Service at +1-800-932-0760 (International). Hours ofOperation: 8:00 am EST to 6:00 pm EST.Please be assured that maintaining a high level of safety and quality is our highest priority. If you have any questions, please contact us immediately per the contact information above. |  | 2332680 | 1,827,920 units US , 504,760 units OUS | Worldwide Distribution - USA (nationwide) including Puerto Rico; and Internationally to: Belgium, Canada, Chile, Mexico, Costa Rica, Kuwait, United Kingdom, Greece, Czech Republic, Latvia, Turkey, China, France, Italy, Brazil, Korea, Germany, Thailand, Egypt, Austria, Dominican Republic, India, Israel, Colombia, Japan, Netherlands, Venezula, Saudi Arabia, Argentina, Malaysia, Sweden, Lebanon, South Korea, Finland, United Arab Emirates, Taiwan, Panama, Uruguay, Singapore, Hong Kong, Ecuador, Peru, Oman, Spain, Lithuania, and South Africa. | Terminated | July 09, 2014 | 195 | 2332680 | Undetermined
Z-0818-2014 | 67057 | No | MKW | Cyclosporine test system. | Clinical Chemistry | Toxicology | 510(k) | Dimension Vista Cyclosporine Flex reagent cartridge |  |  | December 05, 2013 | 2013 | Siemens Healthcare Diagnostics, Inc. | Complaints were received regarding variability in recovery of QC and patient results with certain lots of the Dimension Vista CSA Flex reagent cartridges. | Nonconforming Material/Component | Siegmens sent an Urgent Medical Device Recall lletter dated December 2013 to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to discontinue use and discard any remaining inventory of the affected product. Customers were also instructed to complete the attached form and fax it to 302-631-8467. Customers were asked to retain the letter with their laboratory records, and forward the letter to any who have received the affected product. Customers with questions were instructed to contact Siemens Customer Care Center - Technical Solutions or their local Siemens technical support representative.For questions regarding this recall call 800-441-9250. |  | 1586 | 1586 | Worldwide Distribution - USA (nationwide) and Canada. | Terminated | August 27, 2014 | 265 | 1586 | Undetermined
Z-1117-2014 | 67377 | No | INM | Mechanical chair. | Physical Medicine | Physical Medicine | 510(K) Exempt | XL CONVALESCENT RECLINER 5291/529S with Padded armrests with guard, Dual fold down side-tables, Removable side panels, Built-in head-rest. |  |  | September 06, 2013 | 2013 | Winco Mfg., LLC | A supplier manufacturing assembly defect with the right hand recline mechanism. Use of this product may result in the loss of functionality of the three (3) position lock bar when in a reclined position. | Process control | Winco sent a Recliner Field Correction Notification letter dated September 11, 2013 to all affected consignees. The letter identified the product, the problem and action to be taken by the consignees. The consignees were instructed to complete and return the enclosed field correction response form back to Winco so they can send replacement mechanism(s) along with instructions on how to exchange the mechanism(s). Customers with questions were instructed to contact Customer Care at 800-237-3377 or email QA-RA@wincomfg.com.For questions regarding this recall call 352-854-2929, ext 130. |  | 369 | 369 units (all models) | Worldwide Distribution - USA nationwide including Alabama, Arkansas, Arizona, California, Colorado, Connecticut, Florida, Georgia, Idaho, Illinois, Indiana, Kanas, Kentucky, Louisiana, Massachusetts, Maryland, Maine, Michigan, Minnesota, Missouri, Mississippi, North Carolina, North Dakota, New Jersey, New Mexico, New York, Ohio, Oklahoma, Pennsylvania, South Carolina, South Dakota, Tennessee, Texas, Virginia, Vermont, Washington, Wisconsin, West Virginia, Puerto Rico and Internationally to the country of Canada. | Terminated | December 30, 2014 | 480 | 369 | Undetermined
Z-0817-2014 | 66474 | No | DIO | Cocaine and cocaine metabolite test system. | Toxicology | Toxicology | 510(k) | VITROS Chemistry Products COCM Reagent |  |  | October 04, 2013 | 2013 | Ortho-Clinical Diagnostics | Inaccurate quality control results. | Other | On October 4, 2013 Ortho Clinical Diagnostics issued a Product Correction Notification to customers via US Postal Service Priority Mail or FedEx regarding VITROS Chemistry Products COCM Reagent due to inaccurate quality control results. Ortho Clinical Diagnostics advised their customers to discontinue use and discard any recalled product. OCD customers with questions should be contact Customer Technical Services at 1-800-421-3311. |  | 3342 | 3,342 units | Worldwide Distribution. US nationwide (including AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, IL, IN, KY, LA, MD, MI, MO, MS, NC, NE, NJ, NY, OH, OK, PA, SC, TN, TX, UT, VA, WA, WI, Puerto Rico), Colombia, India, Japan, Mexico, Republic of Panama, Singapore, Venezuela, England, France, Germany, Italy, and Spain. | Open, Classified |  |  | 3342 | Undetermined
Z-1095-2014 | 67445 | No | MCX |  |  | Cardiovascular | PMA | CSI, Diamondback 360 Coronary Orbital Atherectomy System, Part Number 70058-02, Model Number DBEC-125 |  |  | January 27, 2014 | 2014 | Cardiovascular Systems, Inc. | Cardio vascular Systems Inc. has initiated a recall to request the immediate removal and return of Diamondback 360 Coronary Orbital Atherectomy Device (OAD).The device mistakenly contains a saline line that is used in our peripheral atherectomy device. This saline line was not part of the FDA approval for use in the coronary device. It therefore does not meet specification for this device. C | Process control | Cardiovascular Systems Inc. sent an "Urgent Medical Device Recall" letter dated January 27, 2014 to all affected customers. The letter described the affected product, recall description and actions to be taken. Customers were instructed to complete and return the attached Customer Acknowledgement Form and return the device i available. For further information they should contact CSI Customer Service at 651 Campus Drive, Saint Paul, MN 55112, Tele 877-274-0901, Fax 612-677-3355. |  | 318 | 318 | USA Nationwide Distribution in the states of AZ, CA, FL, MA, MN, MS, NC, NJ, NY, PA, TN, TX, VA. | Terminated | May 05, 2014 | 98 | 318 | Undetermined
Z-1104-2014 | 67313 | Yes | KDI | High permeability hemodialysis system. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Optiflux F200A High Flux Dialyzers; Product Number: 0500320A. |  |  | February 04, 2014 | 2014 | Fresenius Medical Care Holdings, Inc. | Dialysate Port Leak During Priming of the Dialyzers | Under Investigation by firm | Fresenius Medical issued Notice to Clinics and Clinics with Patients Letters Dated - 1st Mailing - February 4, 2014 and Patients Letters Dated - February 5, 2014 - these letters contained the correct information.Corrected letters to Clinics and Clinics with Patients issued 2/14/14 via certified mail with return receipt and fax-back form . Users are advised the leak may be stopped by tightening the dialysate port cap more securely.Questions associated with this product notification to your Customer Service Team at 1-800-323-5188. |  | 5244 | 5,244 units | Worldwide distribution: USA (nationwide) and the country of: Mexico. | Open, Classified |  |  | 7741662 | Undetermined
Z-1094-2014 | 67375 | No | MNB |  |  | Obstetrics/Gynecology | PMA | NovaSure Impedance Controlled Endometrial Ablation System Radiofrequency Controller, Model Number: 71978-001 |  |  | January 24, 2014 | 2014 | Hologic, Inc | NovaSure Radiofrequency Control Units may not meet a requirement of the electrical standard, IEC 60601-1, | Component design/selection | Hologic sent an Important Medical Device Information letter to all affected consignees on January 24, 2014 via certified letter. The letter contains instructions for scheduling the return and replacement of the affected unit, provides contact information, and includes a postage-paid return card for the consignee to complete and return to Hologic. |  | 93 | 93 units | Worldwide Distribution - USA Nationwide and the countries Canada, France, Germany, and Sweden | Terminated | July 17, 2014 | 174 | 93 | Undetermined
Z-1092-2014 | 67067 | No | KWT | Shoulder joint metal/polymer non-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | 00434811113, REF I4348-111-13 |  |  | January 10, 2014 | 2014 | Zimmer, Inc. | This single manufacturing lot has the potential for the bond strength between the Trabecular Metal material and substrate to be below the minimum required specification. Testing found bond strength approximately 80% of the minimum required by specification in one of the three zones of the Trabecular Metal. The other two zones exceeded the minimum specification. | Nonconforming Material/Component | Zimmer sent an Urgent Medical Device Recall dated January 10, 2014 to distributors via electronic mail. The hospital/surgeon that has implanted this manufacturing lot will be identified and notified by the distributor. The letter identified the affected product, problem and actions to be taken. Distributors were instructed to locate and remove the affected product in their territory, as well as identifying and notifying the hospital/surgeon that has implanted this manufacturing lot as applicable. The distributor will return on-hand affected products to Zimmer and ensure all of their products are accounted for using the return form provided in the letter. Distributors and hospital consignees are to complete and return the inventory return form and acknowledgement form indicating which hospital/surgeon has implanted the product and to which they have provided notice to. Email a copy of the completed return form to CorporateQuality.PostMarket@Zimmer.comCommunications outside of the United States will occur approximately two weeks after the United States communications.For questions call 1-877-946-2761 |  | 10 | 10 units | Distribution USA in the state of Illinois. | Terminated | September 11, 2014 | 244 | 10 | Undetermined
Z-1107-2014 | 67328 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | Philips HeartStart XL+ Defibrillator/Monitor |  |  | January 24, 2014 | 2014 | Philips Medical Systems, Inc. | HeartStart XL+ battery charge time to 100% capacity at 35ÂżC (95ÂżF) is slightly longer than the specified 3 hour duration as stated in the labeling | Component design/selection | Philips Healthcare issued a Customer Information Letter on January 24, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customers with an IFU Addendum titled: HeartStart XL+ Instructions for Use Addendum. Philips has revised the labeling in the addendum to specify the battery charge time with device turned off and AC power connected. The revised specification is: With temperature at 25ÂżÂżC (77ÂżÂżF), less than 3 hours to 100% capacity; less than 2 hours to 80% capacity.Please contact your local Philips Representative at 1-800-722-9377. |  | 7157 | 7157 units | Worldwide Distribution - USA (nationwide) including Puerto Rico and the countries of ALGERIA ANGOLA ARGENTINA AUSTRALIA AUSTRIA AZERBAIJAN BAHRAIN BELGIUM BOLIVIA BULGARIA CANADA CHILE CHINA COLOMBIA COSTA RICA CROATIA (local name: Hrvatska) CYPRUS CZECH REPUBLIC DENMARK DOMINICAN REPUBLIC EGYPT EL SALVADOR ESTONIA ETHIOPIA FINLAND FRANCE GABON GERMANY GREECE GUATEMALA HONG KONG HUNGARY ICELAND INDIA INDONESIA IRAN (ISLAMIC REPUBLIC OF) IRAQ IRELAND ISRAEL ITALY JORDAN KENYA KOREA, REPUBLIC OF KUWAIT LATVIA LEBANON LIBERIA LIBYAN ARAB JAMAHIRIYA LITHUANIA MACAU MALAYSIA MALTA MAURITIUS MOROCCO MYANMAR (Burma) NAMIBIA NETHERLANDS NEW ZEALAND NORWAY OMAN PALESTINE PERU PHILIPPINES POLAND PORTUGAL,QATAR ROMANIA RUSSIAN FEDERATION SAUDI ARABIA SENEGAL SERBIA SEYCHELLES SINGAPORE SLOVAKIA (Slovak Republic) SLOVENIA SOUTH AFRICA SPAIN SRI LANKA SWEDEN SWITZERLAND SYRIAN ARAB REPUBLIC TAIWAN, PROVINCE OF CHINA TANZANIA, UNITED REPUBLIC OF THAILAND TRINIDAD AND TOBAGO TUNISIA TURKEY UNITED ARAB EMIRATES UNITED KINGDOM and VIET NAM. | Open, Classified |  |  | 7157 | Undetermined
Z-1097-2014 | 67250 | [6 empty] | The SYMBIONIC LEG is an integrated prosthetic leg for transfemoral and knee disarticulation amputees. |  |  | October 15, 2012 | 2012 | Ossur Americas, Inc. | OSSUR is recalling the Symbionic Leg because it is not recognizing the low battery warning signal from the Symbionic Leg before it powers down. | Other | Ossur Americas contacted all affected customers via phone on October 15, 2012 informing them that a software update for this product is required, along with the recommended instructions to return the product. A loaner prosthetic will be provided to the patient while the software update is being implemented. |  | 144 | 144 units | US Nationwide distribution. | Terminated | February 28, 2014 | 501 | 144 | Undetermined
Z-1109-2014 | 67483 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | Philips HeartStart XL, Defibrillator/Monitor |  |  | February 14, 2014 | 2014 | Philips Medical Systems, Inc. | Philips HeartStart XL may experience a power board failure causing failure to defibrillate | Component design/selection | Philips Healthcare issued an Urgent Medical Device Correction notification/Field Safety Notice dated February 2014 to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers were informed Philips Healthcare Representative will perform a hardware upgrade free of charge. Customers were instructed while awaiting the replacement power board for your device, you may continue to use the HeartStart XL. Identify a readily available backup defibrillator to use in the event the affected HeartStart XL fails to deliver therapy. If you need any further information or support concerning this issue, please contact your local Philips representative at 1-800-722-9377 |  | 5234 | 5234 units | Worldwide Distribution and USA Nationwide and the countries of AUSTRALIA, BANGLADESH, BELGIUM, BRAZIL, CANADA, CHINA, COLOMBIA, ETHIOPIA, FINLAND, FRANCE, GEORGIA, GERMANY, HUNGARY, INDIA, INDONESIA, IRELAND, ITALY, JAPAN, KAZAKHSTAN, KOREA, REPUBLIC OF, MALAYSIA, MEXICO, MONACO, NEPAL, NETHERLANDS, PORTUGAL, SINGAPORE, SOUTH AFRICA, SPAIN, SRI LANKA, SWITZERLAND, TAIWAN, PROVINCE OF CHINA, THAILAND, UNITED KINGDOM, UZBEKISTAN, VENEZUELA, and VIET NAM, | Open, Classified |  |  | 5234 | Undetermined
Z-0795-2014 | 65184 | No | JTN | Antimicrobial susceptibility test disc. | Microbiology | Microbiology | 510(k) | Oxoid Antimicrobial Susceptibility Test Discs, Ertapenem 10 mcg., CT1761B, packaged in cartridges of 50 discs, 5 cartridges/pack. The firm name on the label is Oxoid Ltd., Basingstoke, Hampshire, England. |  |  | April 09, 2013 | 2013 | Remel Inc | Cartridges may contain individual discs which were not impregnated with antibiotic. | Process design | ThermoFisher Scientific sent a Medical Device Recall letter dated March 22, 2013, to affected customers. Letters dated April 9, 2013, were issued for the remaining three products. Customers were instructed to notify all personnel who need to be notified and review their inventory and discard any affected product. Customers were also instructed to return the attaced Medical Device Recall Response acknowleding the dispoasl of the affected product. Customers with questions were instructed to contact the Technical Services Department at 800-255-6730 or 913-888-0939 (International).For questions regarding this recall call 913-895-4185. |  | 30 | 30 packs | Nationwide Distribution: US including Puerto Rico. | Terminated | January 21, 2014 | 287 | 129 | Undetermined
Z-0796-2014 | 67047 | No | IZL | Mobile x-ray system. | Radiology | Radiology | 510(k) | DRX-Evolution with FF WS |  |  | November 25, 2013 | 2013 | Carestream Health, Inc. | Carestream Health Inc. has issued a recall for the Carestream DRX-Evolution (FF WS and CSJ WS) and the Kodak DirectView DR 7500 System due to the potential for unexpected device movement. | Nonconforming Material/Component | The firm, Carestream Health Inc, sent an "URGENT: Medical Device Recall" notice dated November 20, 2013 to their customers. The notice described the product, problem and actions to be taken. The customers were instructed that the Emergency Stop (E-Stop) button be used on the motorized FF WS and CSH WS Buckys for any observed uncontrolled movement as indicated in the user guide. If this movement is observed, please notify Carestream Health, and Carestream will schedule a site visit to resolve the problem. In addition, Carestream Health Inc. will apply a label(s) indicating the location of the E-stop button for customers using the FF WS. If you have any questions or concerns, please contact the Carestream Customer Care Center in the U.S. at 1-800-328-2910. |  | 675 | 675 units Total (145 units domestically &amp; 530 units internationally) | Worldwide Distribution: US (nationwide) including states of: CA, CO, FL, GA, IL,IN, MD, MI, MN, NC, ND, NJ, NY, OH, OK, PA, SC, SD, TN, TX, VA, WA, and WI; and internationally to: Australia Belgium, Canada, China, Finland, France, Germany, Hong Kong, Italy, Netherlands, New Zealand, Poland, Portugal, Spain, Sweden, Switzerland, Taiwan, Turkey, United Arab Emirates, and United Kingdom. | Open, Classified |  |  | 675 | Undetermined
Z-0712-2014 | 67139 | [6 empty] | Boston Scientific Extractor Pro XL Retrieval Balloon Catheter 12-15 mm |  |  | December 18, 2013 | 2013 | Boston Scientific Corporation | Extractor Pro XL Retrieval Balloon Catheters may have the incorrect skive hole location or balloon size | Process control | Boston Scientific issued an Urgent Medical Device Recall Removal letter dated December 18, 2013, to all affected customers via Federal Express Priority mail.3. The notification instructs the account to check their inventory, remove any affected product from their inventory, and return it to Boston Scientific. The notification also requests the completion of a Reply Verification Tracking Form as a method of documenting the presence or absence of affected product from their inventory.For questions regarding this recall call 508-382-9555. |  | 34 | 34 units | Worldwide Distribution - USA (Nationwide Distribution and Puerto Rico) and Internationally. | Terminated | August 19, 2014 | 244 | 111 | Undetermined
Z-1201-2014 | 67516 | No | FIP | Water purification system for hemodialysis. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Mar Cor Purification CWP Reverse Osmosis System. Models 101, 102, 104, and 106. |  |  | June 27, 2012 | 2012 | Mar Cor Purification | There exists a potential situation with the crimp connection at the pump head assembly. The potential exists that the reverse osmosis pump head wiring could build resistance resulting in resistive heating of the surrounding components. If this occurs, the pump head assembly may experience melting of the plastic components and potentially create smoke/fire and shut down of the unit. | Component design/selection | Mar Cor Technicians started visiting the consignees to make the correction, in June 2012 Mar Cor called the consignee prior to the visit. An Urgent Medical Device correction letter was sent to customers beginning 2/27/2014. the letter identified the affected units, described the issue, stated that a correction has been, or will be made to the devices and asked for a response form be sent back to the firm. |  | 130 | 130 | nationwide. No distribution outside the US. | Terminated | August 22, 2014 | 786 | 130 | Undetermined
Z-1256-2014 | 67503 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Natural-KneeÂż II Constrained Knee System |  |  | February 25, 2014 | 2014 | Zimmer, Inc. | A complaint was received for a discolored All Poly Tibia. Root cause analysis found that the discoloration was the result of free radical oxidation. Theoxidation process was advanced prematurely as a result of the machined component being stored in a transparent container exposed to artificiallight for six years prior to inert packaging and sterilization of the device. | Storage | Beginning 2/25/2014 all distributors were notified via electronic mail. Hospital risk managers and distributors, surgeons with affected inventory will also be notified via courier or direct mail. The Urgent Lot Specific Recall Notice identified the issue and responsibilities. Distributors were asked to locate and remove the affected product in their territory, as well as identifying hospitals and surgeons who have previously implanted the affected product. Distributors were asked to return on-hand affected products to Zimmer and ensure all of their products are accounted for using the form provided in the letter and email it back to corporatequality.postmarket@zimmer.com. Hospital risk managers and surgeons were asked to locate affected product and return it to their Zimmer sales representative. Questions or concerns should be addressed to customer call center at 1-877-946-2761 between 8:00 am and 5:00pm EST Communications outside of the United States will occur approximately two weeks after the United States communications.100 % effective checks will be done. |  | 3638 | 3638 units | Worldwide Distribution-USA (nationwide) including DC and the states of AK, AL, AR, AZ, CA, CO, CT, FL, IL, IN, KS, LA, ME, MI, MO, MS, NC, ND, NE, NJ, NY, OH, OK, OR, PA, TN, TX, UT, VA, WA, and WI, and the countries of Canada, Australia, India, Austria, Japan, Czech Republic, Seoul Korea, France, Venezuela, Germany, Luxembourg, Spain, and Switzerland. | Terminated | March 24, 2015 | 392 | 3638 | Undetermined
Z-1232-2014 | 67612 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | Philips HeartStart Home and OnSite (HS1) automated external defibrillators (AEDs). |  |  | February 06, 2014 | 2014 | Philips Medical Systems North America Inc. | Philips is recalling HeartStart Home and OnSite (HS1) automated external defibrillators (AEDs) because one or both contact pins may be contaminated by residue from the soldering process. This contamination could prevent an adequate connection between the pads cartridge and the AED. If poor contact between the pads cartridge and AED occurs, the device may be unable to deliver therapy in an emergency. | Vendor change control | Philips began sending out the Urgent Medical Device Recall HeartStart Home and OnSite (HS1) letter, dated 2014 FEB 10, to consignees on 2/17/14. The letter identified the product, the problem, and the action to be taken by the customer. Customers are advised to contact Philips for a replacement AED at 1-800-263-3342 and select option 5. Live technical support is available Monday through Friday, 7:00AM to 5:00PM PST. Customers can also leave a message 24 hours a day, seven days a week. DISTRIBUTORS are instructed to provide Philips with the information provided in the letter. Customers who have an affected device will be supplied with a replacement AED from Philips, free of charge. If customers/distributors need any further information or support concerning this issue, please visit www.philips.com/HeartStartHS1PinAction. |  | 5075 | 5075 units total (2724 units in the US and 2351 units outside the US) | Worldwide Distribution - USA Nationwide | Terminated | March 23, 2017 | 1141 | 5075 | Undetermined
Z-0730-2014 | 66990 | No | MBB | Polymethylmethacrylate (PMMA) bone cement. | Orthopedic | Orthopedic | 510(k) | PalacosÂż R+G Radioplaque Bone Cement 1 x 40 g Single (with Gentamicin), 40.8 g Methyl Acrylate Copolymer, 20 ml Methyl Methacrylate Monomer 0.5 g Gentamicin, Distributed by Zimmer Dover, OH 44522 |  |  | November 04, 2013 | 2013 | Zimmer, Inc. | The patient label associated with the product may be incorrect in that the lot number on the patient label may be different than the lot number on the outer carton label. | Other | Zimmer sent an Urgent Device Removal letter dated November 18, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to review the notification, idetntify the product and remove from use. If the affected product has been utilized, follow institutional protocol to determine if it is appropriate to place a copy of the notification letter into patient records for the affected product. Complete the attached Response Form and return it via fax to 800-871-7270. Customers were instructed to return a copy of the completed response form along with their returned product to ensure proper credit. Customers with questions were instructed to call 866-608-3941.For questions regarding this recall call 800-613-6131. |  | 1534 | 1,534 units | Nationwide Distribution including AZ, CA, CO, CT, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, ME, MI, MN, MO, NC, NJ, NV, NY, OH, OK, OR, PA, TN, TX, UT, VA, VT, WA, WI, WV &amp; WY. | Terminated | October 08, 2014 | 338 | 1534 | Undetermined
Z-1225-2014 | 67622 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | Philips HeartStart FR3 Automated External Defibrillator (AED), models 861388 and 861389. |  |  | February 06, 2014 | 2014 | Philips Medical Systems North America Inc. | Philip's is recalling HeartStart FR3 Automated External Defibrillator (AED) because the Printed Circuit Board (PCB), which contains the circuitry that supplies energy to the device, could become separated from the electrical contact points that hold the board in place. In the event the PCB becomes separated from the electrical contact points, the AED may not be ready to deliver a shock in the event of an emergency. | Process control | Philips sent an Urgent Medical Device Recall Letter dated February 10, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. If you are have an affected device, please contact Philips for a replacement AED at 1-800-263-3342 and select option 5. Live technical support is available Monday through Friday, 7:00AM to 5:00PM PST. You can also leave a message24 hours a day, seven days a week; a customer support representative will return your call promptly on the next business day.Customers who have an affected device will be supplied with a replacement AED device of the same model as the recalled device from Philips, free of charge. Please keep your AED in service until you receive a replacement if there are no alternative defibrillators readily available.Philips has asked Distributors to provide them with the following information within five calendar days: -The name, address, telephone number and, if available, email address of all customers who received affected devices from you or from any of your sub-distributors. -The serial numbers of any affected devices remaining in their possession, or in the possession of their sub-distributors.-Philips directed distributors with affected devices in inventory to not sell these products, but to note them as in their possession as part of their response to Philips. -Philips communicated in this letter that they are responsible for returning to Philips any units in their inventory or in the inventory of their sub-distributors, and Philips will provide instructions for returning the affected devices. |  | 251 | 251 units total (128 in USA and 123 untils outside USA) | Worldwide Distribution - US Distribution including the states of AL, CA, CO, WA DC, FL, GA, HI, KS, MD, ME, MI, MS, NC, NM, NY, PA, TX, VA, WA and WI., and the countries of Australia, Austria, Canada, France, Germany, India, Italy, Ireland, Japan, Malaysia, Netherlands, Norway, Switzerland, Taiwan and United Kingdom. | Terminated | June 29, 2017 | 1239 | 251 | Undetermined
Z-1217-2014 | 67486 | No | MNI | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | Gibralt Spine System Polyaxial Screw |  |  | January 06, 2014 | 2014 | Exactech, Inc. | The products have been identified to contain springs which have been manufactured with type-304 stainless steel rather than the specified titanium alloy. | Mixed-up of materials/components | Exactech sent an "URGENT PRODUCT RECALL NOTICE" dated January 6, 2014, to all affected customers. The letter identified the product and action needed to taken by the customers. Customers were instructed to: 1) Immediately cease distribution. 2) Extend the information to accounts in possession of the product. 3) Verify if they have any of the subject Gibralt Screws from the specified lots. 4) Complete and fax back the attached form. 5) Return affected products. For further questions, please call ( 352) 327-4847. |  | 242 | 242 | USA Distribution including the states of FL, NY, OR, IL, PA, MO and NJ. | Terminated | August 04, 2014 | 210 | 242 | Undetermined
Z-1212-2014 | 67472 | No | NBW | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | FreeStyle Flash Blood Glucose Monitors System. |  |  | February 19, 2014 | 2014 | Abbott Diabetes Care, Inc. | Abbott Diabetes Care has identified through internal testing and investigation that all non-applied voltage legacy meters have the potential to produce out of range control solution results and/or erroneously low blood glucose results when used in conjunction with Free Style test strips lots within expiry. | Material/Component Contamination | The firm, Abbott Diabetes Care, sent an "IMPORTANT: URGENT PRODUCT RECALL"- letter dated February 14, 2014 to all customers for whom they have a record of receiving 0mV meters on February 19, 2014 via trackable method. Abbott also issued Press Release on March 14, 2014. The letter describe the product, problem and actions to be taken. The customers were instructed to discontinue use of your FreeStyle Flash or FreeStyle Blood Glucose Meter and call Abbott Diabetes Care Customer Service at 1-888-345-5364 for replacement at no charge; upon receiving replacement meter; immediately dispose of your old meter in accordance with your local and state disposal regulations; and please read your test strip insert carefully to ensure that you are using the correct test strip type for your Blood Glucose Meter. The firm, OmniPod, sent a ACD letter dated February 7, 2014 and included Abbotts February 14, 2014 letter to its customers. The letter described the product, problem, and actions to be taken. The customers were informed that you can safely continue to use the OmniPod System for your insulin delivery needs. The customers were instructed to read all of the materials included in the Abbott communication carefully and make sure you understand how it affects you; to immediately discontinue use of FreeStyle test strips with the Freestyle meter built into the OmniPod system; call Abbott Diabetes Care Customer Service immediately at 1-877-584-5159 for replacement strips; while waiting on replacement strips, use an unaffected meter such as the FreeStyle Freedom blood glucose meter; you can always use an alternative blood glucose meter.The notification will be posted on the Abbott Diabetes Care website at 222.abbottdiabetescare.com.If you have any further questions, please call Abbott Diabetes Care Customer Service at 1-877-584-5159. For questions or concerns with the continued use of your OmniPod Insulin Management System contact Insulet Product Support at 1-800-591-34 |  | 25203 | 25,203 units | Worldwide Distribution: US (nationwide) and countries of: Austria, Canada, Germany, Israel, Netherlands, Norway, Sweden, Switzerland, United Kingdom. | Terminated | August 11, 2015 | 538 | 25208 | Undetermined
Z-0729-2014 | 67146 | No | HBE | Powered simple cranial drills, burrs, trephines, and their accessories. | Neurology | Neurology | 510(k) | Anspach Carbide Cutting Burrs &amp; Drums: |  |  | December 17, 2013 | 2013 | The Anspach Effort, Inc. | Some carbide cutting burr fractured during use. | Process design | The firm, Anspach, sent an "URGENT: MEDICAL DEVICE REMOVAL" letter dated December 19, 2013 to their costumers. The letter described the product, problem and actions to be taken. The customers were instructed to complete the attached reply form confirming your receipt of this letter and return the completed form by Fax to: 1-800-327-6661 or Email: customer.support@synthes.com. Reference RMA 200059216 when returning any devices. If you distribute(d any of the products to other services or facilities, please forward this information as appropriate . Should you have any queries please do not hesitate to contact Anspach Customer Support at 800-327-6887 or e-mail at customer.support@synthes.com. |  | 28562 | 28,562 products for the entire recall | Worldwide distribution: US (Nationwide) including state of Hawaii and Internationally to: Australia, Poland, Canada, South Africa, Israel, and Switzerland. | Open, Classified |  |  | 28562 | Undetermined
Z-1214-2014 | 67529 | No | DJR | Methadone test system. | Toxicology | Toxicology | 510(k) | Product Brand Name(s): Alere TriageÂż TOX+MTD Drug Screen, Catalog Number 94400, Lots W56013B and W56035B. |  |  | February 05, 2014 | 2014 | Alere San Diego, Inc. | Alere San Diego, Inc. is initiating a recall of the Alere TriageÂż TOX Drug Screen, Catalog Number 94400, Lots W56013B and W56035B because Alere's investigation has shown that these lots have a frequency of false negative results for PCP (Phencyclidine) that is higher than the frequency shown in the package insert. If you have questions about previously reported results. | Process control | Alere San Diego, Inc. sent out recall notification letters to customers via direct mail or email. The letter, dated February 3, 2014, titled "URGENT MEDICAL DEVICE RECALL", provided customers the following information: product description with codes, reason for recall, impact, contact information and customer/distributor required actions. A Customer/Distributor Verification Form was provided. Customers were instructed to discontinue use and discard the affected lots. Alternately, if you elect to continue to use the product customers must verify if it will not be used to report results for PCP. Customers were instructed to return the Customer Verification Form by fax to 858-805-8457 with the amount of product they have for credit. The information should be shared with their laboratory staff and the notification should be retained as part of their laboratory Quality System documentation. If the product has been forwarded customers were instructed to provide a copy of the letter to them. Customers with questions were instructed to contact Alere San Diego, Inc. at 877-308-8287 or email Triage.Support@alere.com.For questions regarding this recall call 858-805-2000, ext 3015. |  | 20130 | 305 kits (239 kits of W56013B, 66 kits of W56035B) | Worldwide Distribution - USA (nationwide) and Internationally to Taiwan. | Terminated | October 27, 2014 | 264 | 20130 | Undetermined
Z-1213-2014 | 67517 | No | FIP | Water purification system for hemodialysis. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Mar Cor Purification Millenium HX Portable Reverse Osmosis System. Model 3025975. |  |  | June 26, 2013 | 2013 | Mar Cor Purification | High inlet water pressure beyond specification and the solenoid valve does not meet application which could result in a delay in treatment. | Component design/selection | Mar Cor sent an Urgent Medical Device Correction letter dated March 5, 2014 to all affected customers. The letter identifies the affected devices, what the issue is and tells customers to contact Mar Cor to schedule correction. A response form was asked to be completed and returned. For questions contact Scott Nichols, Project Coordinator/Service Trainer at 1-800-633-3080. |  | 232 | 232 units | USA Nationwide Distribution including Puerto Rico. | Terminated | January 22, 2015 | 575 | 232 | Undetermined
Z-1209-2014 | 67572 | No | HBT | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Reusable Clip Applier Manufactured by AtriCure, Inc., 6217 Centre Park Drive, West Chester, OH 45069. |  |  | January 23, 2014 | 2014 | Atricure Inc | The firm was notified by a customer of a cracked spring in the handle of the RCD1 device. | Nonconforming Material/Component | On 1/24/2014 the firm sent Urgent Recall Notification letters to their customers. |  | 46 | 46 units | Nationwide Distribution-including the states of AL, CA, CO, CT, DE, FL, KY, MD, MI, MN, MO, NC, NY, OH, OR, PA, TN, UT, VA &amp; WA. | Terminated | May 13, 2014 | 110 | 46 | Undetermined
Z-0746-2014 | 67043 | No | FLF | Ethylene oxide gas sterilizer. | General Hospital | General Hospital | 510(k) | STERRAD 100S Sterilization System, Product Code: 10101 |  |  | December 16, 2013 | 2013 | Advanced Sterilization Products | Advanced Sterilization Products (ASP) is recalling the STERRAD NX, STERRAD 50, and STERRAD 100S Sterilization Systems because some healthcare workers detected odors or smells possibly emanating from STERRAD Systems. | Under Investigation by firm | Advanced Sterilization Products (ASP) sent an Urgent Medical Device Field Safety Notification dated December 16, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to complete and return the customer response forms. Customers with questions were also instructed to contact ASP Professional Services directly at (888) 783-7723 and select option number 2.For questions regarding this recall call 949-453-6400. |  | 2366 | 2366 units | Nationwide Distribution | Terminated | March 27, 2014 | 101 | 8183 | Undetermined
Z-1145-2014 | 67520 | No | LYB |  |  | Dental | 510(k) | Denlase: Diode Laser Therapy System and Penlase: |  |  | July 15, 2013 | 2013 | China Daheng Group Inc. | Failure to provide calibration procedures to purchasers, as required by 21 CFR 1040.11(a)(2). For example, the user manual did not contain calibration procedures for Denlase or Penlase Lasers | Radiation Control for Health and Safety Act | China Daheng Group, lnc.(CDH}, ensures tha they will execute seriously their plans for the repair of their electronics products Denlase: Diode Laser Therapy System and Penlase: Dental Laser Therapy System.China Daheng Group, lnc.(CDH}, will send their updated 238 pes of Diode Laser Therapy System operation manual and 200 pes of Dental soft Tissue Laser brochure to their Distributor. They will pay all of their expense including manufacture cost and freight cost.China Daheng Group, lnc.(CDH}, will remedy these defects through changing the manuals and brochures for their distributor, purchasers and any subsequent transferee. They will make new version manuals and brochures. Then they will send them to their customers and distributor after finishing inspection by their quality department according to the procedure Denlase-Y0029A for manual and brochure inspection.CDRH approves the CAP subject to the following condition:.Notification of all dealers and purchasers is to be made within 15 working days of receipt of this letter in the manner specified in 21 CFR 1003.21 and 1003.22. This office and the Food and Drug Administration (FDA) district office coordinator noted below are to be included in the notification. |  | 238 | 238 | CA, KS, IL, PA, NJ, WI, WY, KY, FL and NYIndia, Indonesia, Turkey, Denmark, and France, | Terminated | April 17, 2015 | 641 | 238 | Undetermined
Z-1196-2014 | 67578 | No | KNH | Contraceptive tubal occlusion device (TOD) and introducer. | Obstetrics/Gynecology | Obstetrics/Gynecology | PMA | Gyrus ACMI Disposable Falope-Ring Band Applicator Kits. REF 006889-901 and REF 005280-901 (with 8 mm Disposable Trocan). Sterile EO. Single Use Only. |  |  | February 19, 2014 | 2014 | Gyrus Medical, Inc | Due to an anomaly in the packaging process for specific lots of Disposable Falope-Ring Band Kit, it is possible that a defective seal could be present that may allow a breach of the package's sterile barrier and may compromise the sterility of the product. The breach may or may not be easily seen. | Process design | Olympus sent an Urgent Medical Device Recall letter dated February 19, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Olympus requires you to take the following action:I. Immediately cease any further use of any affected product you have, remove it from your inventory and quarantine it until it is shipped back to us.2. Call your Olympus customer service representative at 1-888-524-7266 to obtain a Returned Goods Authorization so that you may return the product with no charge to you. Olympus will issue a credit t or replacement to your facility y for any returned product.3. Please note on the enclosed questionnaire that you have received this information.4. Fax the completed questionnaire to 484-896-7128 regardless of whether you have any affected inventory at your facility.ln addition, if you may have further distributed this product, please identify your customers, notify them at once of this product recall, and appropriately document your notification process. Your notification to your customers may be enhanced by including a copy of this recall notification letter. |  | 856 | 856 units | US Distribution. | Terminated | April 14, 2015 | 419 | 856 | Undetermined
Z-0734-2014 | 67099 | No | JTN | Antimicrobial susceptibility test disc. | Microbiology | Microbiology | 510(k) | Oxoid Antimicrobial Susceptibility Testing Disc, MEM10, REF: CT0774B, 10 mcg meropenem, Packed in packs containing 5 cartridges per pack, 50 discs per cartridge. Firm name on the label is Oxoid, Ltd., Basingstoke, United Kingdom. |  |  | November 21, 2013 | 2013 | Remel Inc | Individual discs in the lot may not be sufficiently impregnated with antibiotic, resulting in smaller than expected zone sizes. | Unknown/Undetermined by firm | Oxoid and Remel part of Thermo Fisher Scientific sent an Urgent Medical Device Recall letter dated November 20, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customers.The letter asked customers to destroy the affected product, if any was found on hand. The letter also asked if the customer further distributed the product that they inform their downstream customers. A response form was included with the letter and the letter instructed customers to complete and return the response form to indicate they had received the notification and that they had complied with it.If you have any questions please contact our Technical Service Department at 800-255-6730 (US) or 913-888-0939 (International). |  | 225 | 45/5/50 - discs/cartridges/packs | Nationwide Distribution including the states of AL, AZ, CA, CO, FL, HI, IA, IL, KY, LA, MD, MI, MN, MO, MS, NC, NY, OH, SC, TN, TX, VA, WI, and WV | Terminated | December 11, 2014 | 385 | 225 | Undetermined
Z-1193-2014 | 67439 | No | JIT | Calibrator. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | VITROS Immunodiagnostic Products 25-OH Vitamin D Total Calibrators; REF/CATALOG # 684 4055, IVD --- COMMON/USUAL NAME: VITROS Vitamin D Calibrators -- Ortho-Clinical Diagnostics High Wycombe HP12 4DP, UK --- For use in the calibration of the VITROS ECi/ECiQ Immunodiagnostic Systems, the VITROS 3600 Immunodiagnostic Systems, and the VITROS 5600 Integrated System for the quantitative measurement of total 25-OH Vitamin D in human serum. |  |  | January 15, 2014 | 2014 | Ortho-Clinical Diagnostics | Ortho Clinical Diagnostics has determined that the VITROS Immunodiagnostic Products 25-OH Vitamin D Total Calibrators, Lots 0091 through 0235, used with VITROS Immunodiagnostic Products 25-OH Vitamin D Total Reagent Packs, may generate negatively biased results. As a result, users must immediately discontinue using all current lots of VITROS 25-OH Vitamin D Total Calibrators and Reagent Packs. | Component design/selection | On 1/22/2014, Customer Letters were sent via FedEx overnight courier and/or US Postal Service Priority Mail (for PO Boxes only) to US consignees to inform them of the issue and advise them to discontinue use of the affected lot(s) and to inform their Laboratory Medical Director if they suspect that a previously reported result may have been affected. On 1/22/2014, Distributor Letters were sent via US Postal Service Priority Mail and/or FedEx overnight courier to US distributor consignees and instructed them to notify their consignees of the issue and actions. |  | 1554 | 1554 units domestically | Nationwide | Terminated | April 03, 2017 | 1174 | 1554 | Undetermined
Z-1143-2014 | 67550 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Trilogy Ventilators |  |  | February 11, 2014 | 2014 | Phillips Respironics, Inc. | Internal testing revealed a potentially defective component on the power management board that could affect the function of certain Trilogy Ventilators resulting in loss of therapy to patients. | Component design/selection | Philips sent an Urgent - Medical Device Recall letter on February 11, 2014, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to review the information contained in the Notice with all members of their staff who need to be aware of the contents of the communication. For further information or support concerning the issue, customers were instructed to contact Philips Respironics at 877-387-3311.For questions regarding this recall call 724-733-0200. |  | 572 | 572 | Worldwide Distribution - USA (nationwide) and the countries of France, Great Britian, China, India, Italy, Korea, Kuwait, Nehterlands, and Singapore. | Terminated | April 17, 2015 | 430 | 572 | Undetermined
Z-1190-2014 | 66837 | No | JKA | Blood specimen collection device. | Clinical Chemistry | Clinical Chemistry | 510(k) | G1762: 3 mL Line Draw, Luer lock syringe (heparin concentration: 23.5 I.U. per mL) with Filter-ProÂż device, syringe tip cap, and 5 mL Luer lock syringe |  |  | November 22, 2013 | 2013 | Smiths Medical ASD, Inc. | issue with the plunger tip of the PortexÂż Arterial Blood Sampling Line Draw Syringes | Component change control | Smiths Medical sent an Urgent Field Safety Notice dated November 22, 2013, to all affected consignees via FedEx beginning on November 23, 2013. Distributors were instructed to notify their customers. Customers were instructed to inspect their inventory for the affected product and quarantine the affected syringes. Complete and return the attached Urgent Field Safety Notice Confirmation Form by fax to 603-358-1017 or by email to LineDraw@smiths-medical.com. Upon receipt of the completed form, a customer service representative will contact them to arrange for exchange of their unused affecte devices for credit or replacement. Customers were instructed to forward the notice to all personnel who need to be aware within their organization and to any organization where the potentially affected devices have been transferred. Customers were instructed to report any issues with the product to Smiths Medical's Global Complaint Department at 1-866-216-8806 or globalcomplaints@smiths-medical.com.For questions regarding this recall call 603-352-3812. |  |  |  | Worldwide Distribution - USA including AL, AZ, CA, CO, CT, DE, FL, GA, HI, IA, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, NC, NE, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, TN, TX, UT, VA, WA, WI, WV. Internationally to Australia, Austria, Belgium, Belize, Brazil, Canada, Czech Republic, Denmark, England, Finland, France, Germany, Greece, Hungary, Ireland, Israel, Italy, Japan, Netherlands, Norway, Poland, Scotland, Spain, and Sweden. | Open, Classified |  |  |  | Undetermined
Z-1179-2014 | 67521 | No | JJW | Quality control material (assayed and unassayed). | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Cannabinoids DAU Calibrator Control Level 1 |  |  | February 06, 2014 | 2014 | Lin-Zhi International Inc | Product catalog #s 0007c Cannabinoid 37.5 ng/mL Control Level 1 and 0008c Cannabinoid62.5 Control Level 2 of the lots mentioned may give decreased concentration readings towardsthe end of their shelf life due to use of dated but unexpired secondary stock used to create theproducts. | Nonconforming Material/Component | Lin-Zhi International, sent an Recall letter dated February 6, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.The Recall letter identified the product by catalog number and product description as well as lot numbers. The reason for recall is explained and that the issue was confirmed by Lin Zhi Quality Control. The 2 affected lots were made with unexpired but dated secondary stock that is the cause of the problem.Lin Zhi requests that customers inform laboratories and customers to properly dispose of any of the affected lots that are still in use or unopened. Lin Zhi should be notified of the amounts for credit. Questions should be directed to 408-732-3883. |  | 216 | 216 items | US Distribution including the states of SC, NC, LA, FL, TX,, AL and GA and Internationally to Germany. | Terminated | March 10, 2015 | 397 | 346 | Undetermined
Z-0770-2014 | 67129 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Weck, Electrosurgical Coagulation Suction Tube 6 (15.2cm), 11 French, with Stylet, Teleflex Medical, Research Triangle, NC 2770. |  |  | December 19, 2013 | 2013 | Teleflex Medical | Sterility cannot be guaranteed. | Nonconforming Material/Component | Consignees were notified via letter on 12/19/2013. They were instructed to, immediately discontinue use and quarantine any products with the catalog numbers listed above. Consignees are to return all affected product to Teleflex Medical per the instructions on the Urgent Recall Notice. If they have no affected stock, they were to complete the enclosed Recall Acknowledgment Form and fax it to Teleflex Medical Customer Service. Distributors were instructed to communicate this recall to any of their customers who have received product. |  | 17450 | 17,450 ea. | Nationwide Distribution. | Open, Classified |  |  | 17450 | Undetermined
Z-1170-2014 | 67183 | No | LRO | Surgical drape and drape accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | Enforcement Discretion | Contra-Angle Torque Driver Kit For Certain Internal Connection |  |  | July 22, 2013 | 2013 | Biomet 3i, LLC | During recent revised sterilization validation testing, the included surgical trays did not meet the Sterility Assurance Level required. | Unknown/Undetermined by firm | Biomet 3i sent an Urgent Medical Device Field Correction Notice dated July 2013 to all affected customer via email, facsimile, or postal mail. The letter identifed the affected product, problem and actions to be taken. Customers were instructed to complete the attached response form and fax to 561-514-6316. For questions contact Biomet 3i customer service at 1-800-342-5454. |  | 250 | 250 | Worldwide Distribution - USA Nationwide in the states of AL, AZ, CA, CO, CT, DC, DE, FL, GA, IL, KS, KY, LA, MA, MD, MI, MN, MS, MT, NC, NE, NH, NJ, NM, NY, OH, OK, OR, PA, RI, SC, TN, TX, VA, WA, WI, WV, including Hawaii and Puerto Rico and the countries of : Argentina, Austria, Canada, China, Costa Rica, Germany, Dominican Republic, Spain, France, Great Britain, Hong Kong, Hungary, Israel, Japan, Korea, Mexico, Poland, Portugal, Paraguay, Sweden, Singapore, Taiwan, and Uruguay. | Terminated | April 21, 2015 | 638 | 1163 | Undetermined
Z-1176-2014 | 67566 | No | MTD | Electrode recording catheter or electrode recording probe. | Cardiovascular | Cardiovascular | 510(k) | PENTARAY NAV High-Density Mapping Catheter and PENTARAY NAV ECO Catheter, Catalog No. D128201, D128202, D128204, D128205, D128207, D128208, D128210, D128211. |  |  | February 20, 2014 | 2014 | Biosense Webster, Inc. | Biosense Webster is recalling the PentaRay Nav Catheter because it has the potential to have partial tip separation from the catheter shaft.Biosense Webster is expanding the recall on 5/1/14 to include all lots of the PentaRay Nav Catheter and the PentaRay Nav ECO Catheter because they can lead to an insufficient bond of the distal tip to the catheter shaft. | Process design | The firm, Biosense Webster, sent an "URGENT FIELD SAFETY NOTICE MEDICAL DEVICE-VOLUNTARY FIELD RECALL (REMOVAL)" letter dated February 18, 2014 to all customers who purchased the PentaRay Nav Catheter. The letter describes the product, the problems identified and the actions to be taken. The customers were instructed to read the "Description of the Problem" section carefully; immediately identify and set aside all affected lots; maintain a copy of this letter with the affected PentaRay NAV Catheters until all units are returned to Biosense Webster; complete and return the Voluntary Field Removal Certification form via mail to: Biosense Webster, Inc., Attn: Field Action Coordinator, 15715 Arrow Highway, Irwindale, CA 91706, Fax to: 909-839-8514 or email to: rkazmi@its.jnj.com; arrange for return of all affected units you may have in your inventory; pass on this notice to anyone in your facility that needs to be informed, and if any of the affected product have been forwarded to another facility, contact that facility and arrange for the return.Customers with questions related to the information provided in the customer notification, are instructed to contact their Biosense Webster sales representative or call (866)473-7823, Monday through Friday from 7am to 8pm EST. Customers with questions related to the Acknowledgement form and its return, are instructed to contact the Field Action Coordinator at (909) 839-8468.Biosense Webster expanded the recall on 5/1/14 by sending a customer notification letter to all customers on 5/2/14 to include all product lots of the Pentaray Nav Catheter (Catalog No: D128201, D128202, D128204,D128205, D128207, D128208, D128210, D128211). The letter provides important information about the affected products and instructions on how the customers can return the affected products to Biosense Webster. |  | 15672 | 15,672 units total (12,804 units in US) | US Nationwide Distribution. | Terminated | October 24, 2014 | 246 | 15672 | Undetermined
Z-0775-2014 | 67065 | No | MVU | Analyte specific reagents. | Pathology | Hematology | 510(K) Exempt | Estrogen Receptor AB-11 (Clone ID5) |  |  | December 06, 2013 | 2013 | Lab Vision Corporation | Reports that the product is not performing as expected. In house testing showed no staining on breast cancer tissues. | Under Investigation by firm | Lab Vision Corporation part of Thermo Fisher Scientific sent an Urgent Medical Device Recall letter dated December 6, 2013, to all affected customers. Revised letters will be sent on January 8, 2014. Changes include that the reagent is an Analyte Specific Reagent (ASR) and not an in vitro diagnostic device in the US.They also clarify that a false positive was reported by a customer and a false negative was confirmed during in house testing. Letters also correct the expiry date from 12/2013 to 03/2015. Data sheet was also updated to eliminate information stating that the presence of ER in breast tumors indicates an increased likelihood of response to estrogen therapy and also that the positive control must be determined by the investigator.If you have any questions, please contact Technical Service at 1-800-522-7270, option 2, option 7. |  | 5 | 5 | Worldwide Distribution: US (nationwide) including the states of TX and MA and internationally to: Turkey. | Terminated | February 25, 2014 | 81 | 5 | Undetermined
Z-1172-2014 | 67508 | No | JOZ | Automated platelet aggregation system. | Hematology | Hematology | 510(k) | VerifyNow IIb/IIIa Test, Catalog No. 85310, 10-Test Kit and Catalog No. 85011, 25-Test kit. |  |  | February 13, 2014 | 2014 | Accumetrics Inc | Accumetrics is recalling the VerifyNow IIb/IIIa test because it may result in the reporting of an erroneous low platelet aggregation unit (PAU) result. An erroneous low PAU result may cause a low percent inhibition calculation or low baseline PAU. | Nonconforming Material/Component | Accumetrics sent an Urgent Medical Device Recall letter dated February 13, 2014 to all affected customers. The recall letter informs the customers of the problems identified and the actions to be taken. Cusotomers are instructed to complete the attached Customer Account Tracking Form and return it within 10 days by fax, email or mail. Customers with questions are instructed call Accumetrics Customer Support at (800) 643-1640 option 2 or email at support@accumetrics.com. |  | 576 | 576 units total (524 units in US) | USA Nationwide Distribution | Terminated | August 23, 2014 | 191 | 576 | Undetermined
Z-0776-2014 | 66298 | No | HTE | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Arena-L Trial Handle; Part: 11-81002 |  |  | September 24, 2013 | 2013 | SpineFrontier, Inc. | Handles and heads may not be compatible. | Other | The firm, SpineFrontier, sent an "URGENT Arena-L Trials Voluntary Recall Notification" letter dated September 24, 2013 sent to its consignee/customer. The letter described the product, problem and actions to be taken. The consignee/customer were instructed to confirm receipt of the notice and contact Customer Service to initiate the immediate return of the listed parts. In addition, consignee/customer were instructed to complete a recall ACKNOWLEDGEMENT FORM and return it to SpineFrontier via fax to: Attention: Customer Service at 978-232-3991 or email to: customerservice@spinefrontier.com.Please feel free to contact us at 978-232-3990 Monday through Friday 8:30AM to 5:30PM EST with any questions. |  | 23 | 23 devices in total | US Distribution: VA only. | Completed |  |  | 23 | Undetermined
Z-0779-2014 | 67179 | No | ERL | Ear, nose, and throat electric or pneumatic surgical drill. | Ear Nose &amp; Throat | Ear Nose &amp; Throat | 510(k) | Anspach MIA 16 (16cm Minimally Invasive Attachment), Rx Only |  |  | May 06, 2013 | 2013 | The Anspach Effort, Inc. | Product did not reach expected sterility assurance level | Unknown/Undetermined by firm | The firm, Aspach, sent an "URGENT: MEDICAL DEVICE LABELING CORRECTION" letter dated May 13, 2013 to its consignees/customers. The letter described the product, problem and actions to be taken. The consignees/ customers were instructed that Sterrad 100S system should not be used on the MIA 16 attachment (the information only applies to only Sterrad sterilization of this devices and not to any other attachments.), and complete and return the Return Receipt Requested Customer Reply form via fax to: 1-800-327-6661 or email: ProductSupportTeam@synthes.com. If you distribute any of the products to other services or facilities, please forward this information as appropriate.Should you have any queries, please do not hesitated to contact Anspach Product Support at (800) 327-6887. Hours of Operation are 8:30am-5:30pm Eastern Standard Time or email: ProductSupportTeam@synthes.com. |  | 183 | 183 | Nationwide Distribution: FL, IL, CA. IN, MI, CA, NJ, TX, NE, NC, OR, PA, TN, DC, UT, CO, MO, KY, MD, WV, NY, and MN. | Terminated | April 18, 2016 | 1078 | 183 | Undetermined
Z-1155-2014 | 67473 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | Philips HeartStart MRx Monitor/Defibrillator |  |  | February 18, 2014 | 2014 | Philips Medical Systems, Inc. | MRx may display a Red X in the Ready For Use (RFU) indicator during automated tests informing the user that a failure has been detected that may prevent the delivery of a shock and pacing | Component design/selection | Philips Healthcare issued and Urgent Medical Device Correction notification/Field Safety Notice on 02/18/14 to inform customers of the issue, Philips is asking customers to follow the Action to be Taken by Customer/User section of the Urgent Medical Device Correction notification/Field Safety Notice:You can continue to use your MRx prior to receiving the hardware upgrade, provided that if you observe that the MRx is displaying a Red X in the Ready For Use (RFU) indicator\during automated tests, a failure has been detected that may prevent the delivery of a shock or pacing due to this issue. If this is observed you should turn the Therapy Knob to Monitor. An INOP message describing the failure is displayed. If needed, run an Operational Check for further information. If the condition persists, take the device out of use and call for service.Identify a readily available backup defibrillator to use in the event the affected HeartStart MRx fails to deliver therapy.Philips will perform a hardware upgrade (replace therapy boards) free of charge to be installed by a Philips Field Service Engineer. Contact your local Philips representative or call us at 1-800-722-9377. |  | 3541 | 3541 units | NationwideForeign:AUSTRIABAHRAINBELGIUMBRAZILCANADACHINAEGYPTESTONIAFINLANDFRANCEGERMANYGHANAICELANDINDIAINDONESIAIRELANDISRAELITALYJAPANJORDANKENYAKOREA, REPUBLIC OFLATVIALEBANONLIBYAN ARAB JAMAHIRIYALITHUANIAMALAYSIAMEXICOMOZAMBIQUEMYANMAR (Burma)NETHERLANDSNETHERLANDS ANTILLESNEW CALEDONIANEW ZEALANDNORWAYPAKISTANPERUPOLANDPORTUGALQATARRUSSIAN FEDERATIONSAUDI ARABIASINGAPORESOUTH AFRICASPAINSWEDENSWITZERLANDTAIWAN, PROVINCE OF CHINATHAILANDTURKMENISTANUNITED ARAB EMIRATESUNITED KINGDOMVIET NAM | Open, Classified |  |  | 3541 | Undetermined
Z-0785-2014 | 67130 | No | IZJ | Cine or spot fluorographic x-ray camera. | Radiology | Radiology | 510(k) | ADAC ARGUS Gamma Camera Systems &amp; ADAC CIRRUS Gamma Camera Systems. Distributed by Philips Medical Systems, Cleveland, OH |  |  | December 04, 2013 | 2013 | Philips Medical Systems (Cleveland) Inc | The recalling firm determined that visual inspections of the fork joints behind the detector that are conducted during six month preventive maintenance (PM) may not detect small cracks or deficiencies in the fork weldment. | Nonconforming Material/Component | On October 30, 2013 &amp; December 4, 2013, the firm sent Urgent Medical Device Recall letters to their customers. |  | 255 | 255 units | Worldwide Distribution: US (nationwide) including states of: AL, AR, CA, CO, CT, DC, FL, GA, HI, IA, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MT, NC, ND, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, TN, TN, TX, VA, VT, WA, WI &amp; WV; and Internationally to: Austria, Belgium, Canada, Denmark, Germany, Great Britain, Italy, Netherlands, Norway, Pakistan, Russia, Singapore &amp; South Korea. | Terminated | November 06, 2015 | 702 | 255 | Undetermined
Z-1150-2014 | 67353 | No | IZL | Mobile x-ray system. | Radiology | Radiology | 510(k) | Mobile X-ray System, Model: MUX-100, MUX-100H |  |  | November 15, 2013 | 2013 | Shimadzu Medical Systems | Shimadzu Corporation is recalling the Shimadzu Mobile X-ray Sytems (MUX-100 and MUX-100H) because the connecting metal plate which connects the main chassis and the column of MUX-100 may crack. | Nonconforming Material/Component | A Field Safety Notice and an Urgent Voluntary Medical Device Recall Notice were sent to customers to inform them that SHIMADZU CORPORATION has identified a potential issue in certain manufacturing lots of the Shimadzu Mobile X-ray Systems. Customers are informed that as a result of their investigation, it has been learned that the connecting work of very few column used in specific lots for the MUX series was not done correctly, and it may cause a crack. Therefore, if the machine is continuously used in that condition, there is the possibility that the column will start to lean. The notices informs the customers of the actions to be taken. Customers are instructed to contact Shimadzu Medical Systems National Technical Support at (800) 228-1429. |  | 56 | 56 units | USA Nationwide Distribution | Terminated | June 25, 2015 | 587 | 56 | Undetermined
Z-1153-2014 | 67440 | No | HQF | Ophthalmic laser. | Ophthalmic | Ophthalmic | 510(k) | Treatment packs used on the VisuMax Laser Keratome. |  |  | February 03, 2014 | 2014 | Carl Zeiss Meditec AG | Carl Zeiss Treatment pack, lot number M 130010, size M, may actually contain size S. | Mixed-up of materials/components | The firm, Zeiss, sent an "Important Field Safety Advisory Note" letter dated February 3, 2014 to its customer. Letter indicates that a phone call preceded the letter on January 23, 2014. The letter also describes the product, problem and actions to be taken. The customers were instructed to quarantine all treatment packs size "M", batch number M 130010; return the material to Carl Zeiss Meditec, Inc., 5160 Hacienda Dr, Dublin, CA 94568 and Zeiss will issue a credit for all material received, upon receipt of the returned materia, and to complete and return the Conformation Form with the envelop provided.If you have any questions regarding this field corrective action, please call 925-557-4832 or email at paul.shahan@zeiss.com. |  | 11 | 1 package containing 10 individual treatment packs | US Distribution: NV only. | Terminated | March 07, 2014 | 32 | 11 | Undetermined
Z-1151-2014 | 67641 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Hand Control 12FR Suction Coagulator with 10 ft cord and holster, Sterile, Disposable, Model Number 0043-25, packaged TyvekÂż polyester peel pouch, 25 pouches per box. |  |  | February 27, 2014 | 2014 | Megadyne Medical Products, Inc. | Megadyne Medical Products, Inc. is voluntarily recalling Hand Control 12Fr Suction Coagulator with 10ft cord and 132405 Holster, Sterile, Disposable, Lot 132405 due to the potential for a small number of these suction coagulators to self-activate immediately on connection to the electrosurgical generatorand to remain activated until it is disconnected. | Use error | Megadyne Medical Product sent an Urgent Medical Device Recall letter dated February 27, 2014, to all affected customers with follow-up phone calls and emails. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed toexamine their inventory immediately to determine if they have any of the affected product. If so, customers were instructed to remove the product from service and contact Megadyne Customer Service at 1-801-576-9669 or 1-800-747-6110 for return authorization and replacement or credit. If customers have distributed any the affected product, they were instructed to immediately contact their distribution points to have the product returned to Megadyne. Customers with questions were instructed to contact Customer Service at 1-800-747-6110 or 1-801-576-9669 or email Regulatory @Megadyne.com.For questions regarding this recall call 801-576-9669. |  | 494 | 494 | USA including FL, MN, LA, NJ, NM, PA, TX and Internationally to Canada. | Terminated | July 09, 2014 | 132 | 494 | Undetermined
Z-0809-2014 | 67182 | No | DXZ | Intracardiac patch or pledget made of polypropylene, polyethylene terephthalate, or polytetrafluoroethylene. | Cardiovascular | Cardiovascular | 510(k) | Pulmonary Hemi-Artery SG |  |  | November 22, 2013 | 2013 | CryoLife, Inc. | Serological testing for the donor was performed with a blood sample that may have been hemodiluted due to the administration of normal saline to the donor prior to death. | Other | CryoLife sent an Urgent Human Tissue Recall letter dated November 22, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to contact CryoLife to arrange for the return of the affected product. If the affected product had been implanted or discarded, customers should provide information regarding the status of the affected product on the attached self-addressed postcard. Customers with questions were instructedto contact the Field Assurance Department at 800-438-8285.For questions regarding this recall call 770-419-3355. |  | 1 | One | US distribution in CA. | Terminated | February 20, 2014 | 90 | 1 | Undetermined
Z-0978-2014 | 65309 | No | IXI | Radiation therapy beam-shaping block. | Radiology | Radiology | 510(k) | DMLC IV-ERGO |  |  | May 01, 2013 | 2013 | Elekta, Inc. | It is possible for an intermittent error to occur in the gantry sensor of the DMLC IV Systems | Component design/selection | Elekta sent an Important Field Safety Notice letter dated May 2013 to affected customers. The letter identified the scope, problem, clinical impact, solution, and safety reference. For question call +44 (0)1293 544422 |  | 35 | 35 units | US Nationwide Distribution in the states of CT, FL, GA, IA, MO, NY, ND, OK, RI, TX, VA, WA | Terminated | January 29, 2015 | 638 | 35 | Undetermined
Z-1072-2014 | 67004 | Yes | BTL | Powered emergency ventilator. | Anesthesiology | Anesthesiology | 510(k) | GE Healthcare - GiraffeÂż and PandaÂż iRes Infant Warmer |  |  | October 01, 2013 | 2013 | GE Healthcare, LLC | Oxygen and air wall inlet fittings on the back panel of the resuscitation systems were reversed during assembly. | Component design/selection | GE Healthcare sent an Urgent Medical Device Correction letter dated November 18, 2013, with a return receipt to all affected customers. The letter identified the products, the problem, and safety instructions to be taken by the customer along with the product correction. It is recommended that all potential users in your facility are made aware of this safety notification and the recommended actions. Safety issues described the description and affected product details, safety instructions, photos showing the back of the resuscitation system where the Air and Oxygen fittings are located and an appendix of the Affected Product Serial Number Definition and Configurations of Fittings. For questions or concerns regarding this notification, please call the phone number listed below: United States: 1-800-345-2700. |  | 172 | 12, 373 units total, 8,807 = US, 3,394 = OUS, 172 = unknown | Worldwide Distribution - USA Nationwide including the states of AK, AL, AR, AZ, CA, CO, CT, DE, FL,GA. HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, PR, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, and WY; and the countries of Albania, Austria, Australia, Bahrain, Belarus, Belgium, Canada, China, Costa Rica, Cyprus, Czech Republic, Denmark, Egypt, Estonia, Finland, France, Greece, Guatemala, Germany, Honduras, Hong Kong, Hungry, India, Indonesia, Ireland, Italy, Kazakhstan, Korea, Lithuania, Malaysia, Mexico, Netherlands, Norway, Peru, Poland, Romania, Russia, Saudi Arabia, South Africia, Spain, Sweden, Switzerland Turkey, United Arab, United Kingdom, and Venezuela. | Terminated | July 22, 2016 | 1025 | 172 | Undetermined
Z-0936-2014 | 67249 | component, external, limb, ankle/foot | ISH | External limb prosthetic component. | Physical Medicine | Physical Medicine | 510(K) Exempt | The Symbionic Leg is an integrated prosthetic leg for transfemoral and knee disarticulation amputees. |  |  | February 12, 2012 | 2012 | Ossur Americas, Inc. | OSSUR initiated this recall of the Symbionic Leg because it may have a possible defect in a small batch of electronics. | Nonconforming Material/Component | OSSUR contacted all affected customers via phone on February 12, 2013, informing them that a voluntary recall for this product has been implemented, along with the recommended instructions to return the product. A new replacement prosthetic will be provided to the customer.OSSUR has inspected and made appropriate modification to all inventory. Âżssur will also further strengthen the design and our manufacturing process to ensure an even more robust product.For further questions please call (949) 303-3239. |  | 17 | 17 units | US Distribution including the states of AL, DC, GA, NY, NJ, OK and PA. | Terminated | February 06, 2014 | 725 | 17 | Undetermined
Z-0849-2014 | 67055 | No | HSB | Intramedullary fixation rod. | Orthopedic | Orthopedic | 510(k) | Teflon Tube, sterile |  |  | November 07, 2013 | 2013 | Stryker Howmedica Osteonics Corp. | Stryker became aware during laboratory testing that there is a potential that the seal integrity of the outer pouch (sterile barrier) may be compromised for certain lots of the product Teflon Tube, sterile. | Process design | The firm, Stryker, sent an "URGENT MEDICAL DEVICE REMOVAL" letter dated December 17, 2013 to its customers via e-mail. Preliminary notification letters and Product Accountability Forms were sent via FedEx on November 15, 2013 with return receipt to the Stryker locations. Notification and Product Accountability Forms were sent 12/18/2013 via Fed Ex with return receipt. Updated letters were sent via Fed Ex on 12/19/2013 with return receipt. The letter described the product, problem and actions to be taken. The customers were instructed to quarantine affected devices in local inventory as well as retrieve any devices located in hospitals; inform users of the Medical Device Removal and pass this to all those individuals who need to be aware within your organization; return all affected products available to your location to: Stryker Osteosynthesis, c/o Colleen O'Meara, Stryker Orthopaedics, 325 Corporate Drive, Mahwah, NJ, 07430, REF: PFA#2013-168 or contact Stryker customer service and refer to PFA #2013-166 for returning the product to us; complete and return the enclosed Business Reply Form and fax a copy to: 1-865-251-3635 or email a copy to Recall Coordinator at aminah.crawford@stryker.com, and keep a copy of the completed and executed Business Reply form for your records.If you have any further questions, contact Manager, Regulatory Compliance at 1-201-972-2100 or email: colleen.omeara@stryker.com. |  | 34 | 34 units | US Distribution in states of: GA, IL, KY, MI, MO, NJ, TN, UT and VA. | Terminated | March 31, 2015 | 509 | 34 | Undetermined
Z-0846-2014 | 67213 | No | OCC | Respiratory viral panel multiplex nucleic acid assay. | Microbiology | Microbiology | 510(k) | Simplexa Flu A/B &amp; RSV Direct assay, Model MOL2650. |  |  | January 10, 2014 | 2014 | Focus Diagnostics Inc | Focus Diagnostics is initiating an urgent safety notice correction for Simplexa Flu A/B &amp; RSV Direct assay (MOL2650) because Focus Diagnostics has received some customer complaints of Simplexa Flu A/B &amp; RSV lots due to sporadic false positive RSV signals, which may result in a higher RSV false positive rate. | Pending | Focus Diagnostics sent an Urgent Safety Notice Correction letter dated January 10, 2014, via email and a hard copy letter by mail. The letter informed customers of the correction, product description with codes, product intended use, issue, risk to health, background, recommendation, actions by customers, and contact information. Customers were instructed to check their inventory for the affected product, if customers detect any positive RSV results, confirm the positive result by an alternate method. Customers were asked to complete the "URGENT: SAFETY NOTICE Acknowledgement and Receipt Form Response Required". Customers with questions were instructed to contact their Technical Services department at 800-838-4548, select option 3.For questions regarding this recall call 714-220-1900. |  | 1052 | 1,052 kits | Worldwide Distribution - US (Nationwide) including states of: NY, MA, NJ, OH, CA, HI, CO, MN, FL, PA, IA, WI, AZ, TX, TN, IL, SD, MO, NM, AR, WV, MI and Internationally to: Sweden, Israel, Canada, and Slovenia. . | Terminated | October 15, 2014 | 278 | 1052 | Undetermined
Z-0850-2014 | 66832 | No | GXI | Radiofrequency lesion probe. | Neurology | Neurology | 510(k) | Kimberly-Clark RadiOpaque Radiofrequency Cannula, Gauge 22, Length 100 mm, Active Tip 10 mm, Distributed in the US by Kimberly-Clark,Product code PMF22-100-10 |  |  | October 29, 2013 | 2013 | Kimberly-Clark Corporation | Product may contain a cannula with an active tip length of only 5mm rather than 10mm as indicated on the label. | Employee error | The firm, Kimberly-Clark Corporation, sent an "URGENT: Voluntary Medical Device Recall Notification" letter dated October 29, 2013 to all direct customers and distributors via federal express priority overnight. The letter described the product, problem and actions to be taken. The customers were instructed to evaluate your Facility's inventory of Kimberly-Clark RadiOpaque Radiofrequency Cannula to determine if any of the two impacted product lots remain within your facility. If product is present, discontinue use of the impacted product and place in a quarantined location; complete and return the attached Product Recall Response Form-Customers via fax to: 1-920-225-4182 within five (5) business days of receipt of this letter (complete and return form even if either of the impacted product lots is NOT present in your facility). Once response form is received, a Kimberly-Clark representative will contact you regarding return and replacement of the product. If you require further assistance, please call Kimberly-Clark Customer Service at 800-KC-HELPS (800-524-3577). |  | 550 | 55 cases (10 units per case = 550 units) | US Distribution in states of: AZ, CA, DE, FL, LA, MD, MA, NJ, NY, OH, PA and TX. | Terminated | September 03, 2014 | 309 | 550 | Undetermined
Z-0967-2014 | 67017 | No | HAW | Stereotaxic instrument. | Neurology | Neurology | 510(k) | BrainLAB's VectorVision spine is intended for use as an intraoperative image-guided localization system for minimally invasive surgery. It links a freehand probe, tracked by a passive marker sensor system to virtual computer image space on a patient's preoperative image data that is processed by a VectorVision workstation. |  |  | November 15, 2013 | 2013 | Brainlab AG | Brainlab has determined that pins of thee Instrument Star Unit ML (Calibration with ICM4) with specific serial numbers might have been made of an incorrect material that is not suitable for the necessary disinfection and sterilization of the product. If corrosion occurs and is not detected by the user, and the device is used during surgery, corroded particles of the pin could, directly or indirect | Nonconforming Material/Component | BrainLab send a A FIELD SAFETY NOTICE / PRODUCT NOTIFICATION letter dated November 15, 2013, to all affected customers. The letter requested that customers verify that they are in possession of one or more of the affected devices, remove the affected devices from use, and that Brainlab will provide replacement parts for the customers. Customers with questions can contact Brainlab by telephone at 1-800-597-5911 or via e-mail at us.support@brainlab.com.For questions regarding this recall call 800-597-5911. |  | 25 | 25 units | Worldwide Distribution - USA including CA, MN, NC, NY, OH, TX and DC. Internationally to Australia, Chile, Germany, Hungary, Italy, Panama, Qatar, Russia, Thailand, United Arab Emirates. | Open, Classified |  |  | 25 | Undetermined
Z-0964-2014 | 67234 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Stryker Orthopaedics Navigation Compatible Accolade Broach Handle |  |  | November 20, 2013 | 2013 | Stryker Howmedica Osteonics Corp. | Stryker received reports from the field of the impaction plate dissociating/fracturing from the main body of the broach handle. | Process design | Stryker Orthopaedics sent notification letters and acknowledgement forms via Fed Ex as follows:Branches-11/20/2013Hospital Risk Manager - 11/21/2013 |  | 211 | 211 units | IL MA, MN, ND, OR, &amp; PA | Open, Classified |  |  | 211 | Undetermined
Z-0855-2014 | 66969 | No | KWP | Spinal interlaminal fixation orthosis. | Orthopedic | Orthopedic | 510(k) | Orthofix Firebird Spinal Fixation System Instrumentation Modular Screw Driver, |  |  | November 26, 2013 | 2013 | Orthofix, Inc. | Orthofix received 6 complaints which resulted in a reportable events due to extended surgical times greater than 30 minutes for the Modular Screw Driver (PN 52-1332). The complaints alleged that the screw driver's collet would malfunction resulting in the surgeon being unable to use the Modular Screw Driver to effectively place Modular Screws, which may result in a delay of surgery. | Nonconforming Material/Component | Orthofix sent an Urgent Medical Device Recall Notification letter dated November 26, 2013 by certified mail to distributors. The letter identified the affected product, problem and actions to be taken. The letter instructed distributors to immediately cease any further distribution, return affected products to Orthofix at no cost and complete the attached reply form. For questions contact your local Orthofix representative. |  | 443 | 443 | Worldwide Distribution - US Nationwide including (Puerto Rico) and the countries of: Spain, Germany, Australia, Peru, Mexico, Colombia, South Africa, and Israel. | Terminated | December 09, 2014 | 378 | 443 | Undetermined
Z-0945-2014 | 65667 | No | KKX | Surgical drape and drape accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | MICROTEK(TM) MEDICAL BUR HOLE LATEX FREE PROBE COVER W/ GEL, featuring IsoSilk(R), 13cm x 244 cm, 5" x 96", REF PC3688EU, STERILE EO, QTY 20, Microtek Medical, Inc., Columbus, MS 39702 |  |  | July 02, 2013 | 2013 | Ecolab Inc | Certain lots of these Bur Hole Probe Covers did not undergo endotoxin (LAL) testing to the proper levels. | Process design | The recall was conducted to the consumer/user level. The firm notified their direct consignees of the problem and the recall by written letter mailed on July 2, 2013. The notice requested the consignees to quarantine any recalled product on hand, and complete and return the response form by July 31, 2013. The consignees who further distributed the recalled product were asked to notify their consignees. |  | 12546 | 12,546 individual units | Worldwide Distribution-USA (nationwide) and the countries of Austria, Germany, Finland, Ireland, France, Spain, Great Britain, Belgium, Netherlands, Switzerland, Jordan, United Kingdom, Saudi Arabia, Hungary, Latvia, Malta, Norway, Poland, Portugal, and Canada. | Terminated | May 14, 2014 | 316 | 29258 | Undetermined
Z-0968-2014 | 66498 | micro mixer | JRB | General purpose laboratory equipment labeled or promoted for a specific medical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | VITROS Chemistry Products Mixing Cup Arrays utilized on VITROS 250 &amp; 350 Chemistry Systems |  |  | October 10, 2013 | 2013 | Ortho-Clinical Diagnostics | Ortho Clinical Diagnostics issued a product correction notification for VITROS Chemistry Mixing Cup Arrays utilized on VITROS 250 and VITROS 350 Chemistry Systems. Increase in Software error due to affected Mixing Cup Arrays. | Nonconforming Material/Component | Ortho Clinical Diagnostics sent an Important Product Correction Notification dated October 10, 2013 to all affected customer. The letter identifed the affected product, problem and actions to be taken. Customers were instructed to complete and return the attached Confirmation of Receipt form. For questions contact the Customer Technical Services at 1-800-421-3311. |  | 306240 | 6380 units (each unit contains 48 arrays with 10 microcups) | Worldwide Distribution - US Nationwide and the countries of: Australia, Brazil, Canada, Chile, China, Colombia, India, Japan, Mexico, PR, Republic of Panama, Singapore, Venezuela, England, France, Germany, Italy, Spain | Open, Classified |  |  | 306240 | Undetermined
Z-0848-2014 | 67100 | No | BTI | Portable air compressor. | Anesthesiology | Anesthesiology | 510(K) Exempt | Compressor Mini |  |  | December 05, 2013 | 2013 | Maquet Cardiovascular Us Sales, Llc | Maquet Compressor Mini may deliver compressed air at a temperature higher than specified. | Material/Component Contamination | Maquet sent a product recall urgent-medical device field correction letter/customer fax response via Fed Ex on December 5, 2013, to their customers.Customers are advised of the issues with the device and provided instructions. The correction will include an update with replacement of the affected connector and if needed replacement of the thermoelectric cooler if it damaged. These operations will be performed by an authorized Maquet representative. Any questions can be directed to customer service at 1-888-627-8383 option 2 followed by option 3 Monday through Friday between the hours of 8 am and 6 pm EST.Customers should completed and return the enclosed confirmation form or scan a copy to field actions@maquet.com or fax (973) 807-9290. |  | 237 | 237 units | Worldwide Distribution US, Canada and Japan. | Terminated | February 05, 2015 | 427 | 237 | Undetermined
Z-0852-2014 | 67117 | No | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | The ACUTE Innovations Modular RibLoc System (RibLoc U plus) is comprised titanium plates and screws that are used to stabilize rib fractures, fusions and osteotomies during the normal bone healing process. |  |  | October 30, 2013 | 2013 | ACUTE Innovations, LLC | The ACUTE Innovations Ribloc U plus Rib Fracture Plating System is recalled because it has the potential to malfunction during installation in a surgery. | Under Investigation by firm | The firm sent the Urgent Notice: Device Recall letter, dated 30 Oct. 2013, to their consignees.The recall notification letter requested consignees to return the product to Acute and a return authorization number provided.Consignees should contact Acute Innovations Customer Service at 866-623-4137 for any questions regarding this recall notification. |  | 176 | 176 units | US Distribution in states of Arizona, Georgia, Illinois, Indiana, Utah, and Wisconsin. | Terminated | March 13, 2014 | 134 | 176 | Undetermined
Z-0939-2014 | 67246 | No | MQV | Resorbable calcium salt bone void filler device. | Orthopedic | Orthopedic | 510(k) | ***REF 606-01-05***OPTEFORM Allograft Paste 5 cc, bone void filler of bony defects in dental and orthopedic applications.***STERILE Using Radiation. ***Rx Only***Distributor: EXACTECH 2320 NW 66th Court, Gainesville, FL 32653 USA***Manufacturer: RTI Biologicals, Inc. 11621 Research Circle, Alachua, FL 32615 USA. |  |  | December 06, 2013 | 2013 | RTI Surgical, Inc. | RTI Surgical, Inc. of Alachua, FL is recalling their OPTEFORM Allograft Paste, 5cc as it may contain a 10 cc fluid dispenser. The 10cc fluid dispenser does not include the correct fluid fill line reference for reconstitution of 5cc past allografts. | Other | RTI Surgical issued a: URGENT: Medical Device Recall Notification dated December 6, 2013 to their customer. The customer was instructed to quarantine the product immediately and return to RTI Surgical, Inc. for final disposition. For questions or concerns, call 386-418-8888. |  | 14 | 14 units | Distributed to Florida. | Terminated | September 15, 2014 | 283 | 14 | Undetermined
Z-0938-2014 | 67350 | No | NKB | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | Malibu Polyaxial Screwdriver, 8.0 Plus, Model Number 91-2109; |  |  | January 22, 2014 | 2014 | Integra LifeSciences Corp. | Integra LifeSciences has identified through an investigation of complaints that there may be the potential for the shaftof the Malibu Polyaxial Drivers to break during use. None of the complaints reported a patientinjury. However, some complaints did indicate the surgery was slightly prolonged while the additional driver providedin the kit was retrieved and used. | Manufacturing material removal | Integra LifeSciences initiated this recall by sending recall notification letter to US consignees, delivered by traceable courier service. Non US consignees notified of the recall by either traceable e-mail or facsimile. The letter dated, January 23, 2014, titled "URGENT: VOLUNTARY MEDICAL DEVICE RECALL IntegraÂż Malibu Polyaxial DriversCatalogue Numbers: 91-2109 &amp; 91- 2110", informed customers of the recall and provided product description with codes, picture of the product, advised customers to: Inspect their inventory of Malibu Polyaxial Drivers.- If you have product(s) with a lot number listed on the following page, STOP USING THEM IMMEDIATLY.- If you do not have Drivers with a lot number listed on the following page, you can continue to use them.Customers were also instructed that regardless if they have Drivers with the affected lots, to complete the attached form and return it to Integra as indicated. Customers were instructed to keep a copy of the form for their records., and contact information. Customers with questions were instructed to contact Customers Service at 1-866-942-8698.For questions regarding this recall call 609-275-0500. |  | 40 | 40 | Worldwide Distribution - USA including IA, FL, LA, IL, CA, TX, AZ, NC, MD, NM, VA, and Internationally to South Africa. | Terminated | December 17, 2014 | 329 | 40 | Undetermined
Z-0774-2014 | 67060 | No | KPR | Stationary x-ray system. | Radiology | Radiology | 510(k) | Carestream DRX Evolution System (Standard Q) |  |  | August 20, 2013 | 2013 | Carestream Health Inc. | Carestream received a report of an exposure being initiated without a patient image captured. The patient was imaged 2 times before an image was captured. The field engineer who visited the site reported the system allowed an exposure without the assignment of a bucky or image receptor. | Radiation Control for Health and Safety Act | The Center for Devices and Radiological Health (CDRH) understands that the CAP (Corrective Action Plan) involving 4 units will consist of the following:a. Distribution of a letter to all affected customers informing them of the issue, how to prevent the issue including a reminder of instructions for proper x-ray system setup and operation found in the User Manual. The letter will also include actions to be taken by Carestream to correct the issue. This communication will be marked and distributed pursuant to the requirements of 21 CFR 1003.21.b. The next stage of correction will be the installation of a software patch to any potentially affected customer sites. The software patch has been developed to prevent the system from allowing an exposure when the bucky and detector are not selected. This patch has been verified and proven successful prior to rollout in manufacturing. Once receiving FDA approval on the corrective action plan, the correction will be implemented within the US to all affected devices at identified customer sites.c. The completion of the Corrective Action Plan, once approved, is expected to take approximately 3 months.d. Carestream will provide FDA with a report demonstrating effective closure of the activity.CDRH approves the draft of the customer notification letter and CAP subject to the following condition:Notification of all dealers and purchasers is to be made within 15 working days of receipt of this letter in the manner specified in 21 CFR 1003.21 and 1003.22. This office and the Food and Drug Administration (FDA) district office coordinator noted below are to be included in the notification.For further questions please call 585-627-6473. |  | 4 | 4 units | Nationwide Distribution including TX, IL, and GA. | Open, Classified |  |  | 4 | Undetermined
Z-0935-2014 | 67315 | No | DZE | Endosseous dental implant. | Dental | Dental | 510(k) | LOCATORÂż Implant Attachment System. |  |  | January 16, 2014 | 2014 | Zest Anchors LLC | Zest Anchors, LLC is voluntarily recalling specific lots of LOCATORÂż Implant Attachment Systems because it is currently provided with a label that indicates that the final torque applied to the abutment should be 30 N-cm. It was recently noted by Zest that in order to provide an optimal performance by increasing the factor of safety between the failure load of the abutment screw and the recommend | Unknown/Undetermined by firm | Zest Anchors, LLC initiated this recall by sending customers recall notification letter on 01/16/2014 via FedEx overnight.The customer notification letter, titled, "IMPORTANT: MEDICAL DEVICE VOLUNTARY RECALL NOTICE", informed customers of the recall, reason for voluntary recall, product information with part numbers and lot numbers, actions to be taken, and contact information. The customer notification letter was accompanied by a response form titled "MEDICAL DEVICE RECALL RETURN RESPONSE Acknowledgement &amp; Receipt Form". |  | 3130 | 3,130 | Distributed in Ohio and Switzerland. | Terminated | May 15, 2014 | 119 | 3130 | Undetermined
Z-0931-2014 | 67326 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Trifurcated Extension Set, model no. ME 1224. |  |  | January 16, 2014 | 2014 | CareFusion 303, Inc. | CareFusion is recalling the Trifurcated Extension Set, models ME 1224 and MP 9230, because they may have a defect in the trifurcated component that could result in a component leak. | Nonconforming Material/Component | A customer notification letter and a distribution letter are dated 1/22/14 were sent to customers to inform them that CareFusion is recalling the Trifurcated Extension Set, models ME 1224 and MP 9230 may have a defect in the trifurcated component that could result in the component to leak. The letters inform the customers of the problems identified and the actions to be taken. Customers are instructed to complete and return to CareFusion the enclose, pre-addressed and postage paid, recall response card. Customers are instructed that once CareFusion receives the recall response card, the CareFusion Support Center will provide instructions for return of the affected lot number and replacement of the affected lot. Customers are instructed to contact CareFusion Support Center at (888) 562-6018, Customer advocacy at (888) 812-3266, customerfeedback@carefusion.com, or technical support at (888) 812-3229. |  | 9550 | 9,550 units | Nationwide in US and Canada | Terminated | August 06, 2014 | 202 | 9550 | Undetermined
Z-1087-2014 | 66774 | No | JGS | Sodium test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | VITROS Chemistry Products Na+ Slides, manufactured by Ortho Clinical Diagnostics 513 Technology Blvd Rochester, NY 14652 |  |  | October 17, 2013 | 2013 | Ortho-Clinical Diagnostics | Ortho Clinical Diagnostics has recalled VITROS Chemistry Products Na+ Slides and the Calibrator Kit 2 due to false positive serum and plasma sample results. | Under Investigation by firm | Ortho Clinical Diagnostics sent an Urgent Product Correction Notification dated October 17, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customers. Ortho Clinical Diagnostics recalled VITROS Chemistry Products Na+ Slides and the Calibrator Kit 2 due to false positive serum and plasma sample. Please forward a copy of the notification letter if the products were further distributed. Complete and return the attached Confirmation of Receipt form by October 25, 2013. Customers with questions please call Customer Technical Service at 1-800-421-3311. |  | 548663 | 548,663 units | Worldwide Distribution - USA (nationwide) including Puerto Rico and the states of AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NY, OH, OK, OR, PA, SC, SD, TN, TX, UT, VA, WA, WI, WV and WY., and the countries of Australia, Brazil, Canada, Chile, China, Colombia, India, Japan, Mexico, Republic of Panama, Singapore, Venezuela, England, Germany, France, Italy and Spain. | Open, Classified |  |  | 548663 | Undetermined
Z-0930-2014 | 67021 | No | DIP | Diphenylhydantoin test system. | Toxicology | Toxicology | 510(k) | ADVIA CentaurÂż Systems Phenytoin Calibrator N ( PHTN) 6-pack |  |  | December 18, 2013 | 2013 | Siemens Healthcare Diagnostics, Inc | ADVIA CentaurÂż Systems Phenytoin Calibrator N, lots CN58 and higher, change in recovery observed with the ADVIA Centaur Systems Phenytoin assay compared to recoveries using previous Calibrator N lots. | Under Investigation by firm | Siemens Healthcare Diagnostics Customer Notification notice was sent to affected Siemens Healthcare Diagnostic customers in the United States on December 18, 2013.. The notifications inform customers that Siemens Healthcare Diagnostics received customer feedback regarding the performance of the ADVIA CentaraÂż Systems Phenytoin assay, indicating that the assay was recovering higher relative to the performance of other assays in proficiency studies. Siemens internal investigation confirmed the positive bias for Calibrator N lot CN 57 in the ADVIA Centaur Systems Phenytoin assay. This lot has expired and is no longer in use. Expected performance has been restored starting with Calibrator N lot CN58. Siemens is notifying customers that have received Calibrator N lots CN58 and higher that there will be a change in recovery observed with the ADVIA Centaur Systems Phenytoin assay compared to recoveries using previous Calibrator N lots. Siemens is providing data to inform customers of the expected difference between Calibrator lots CN58 and higher versus prior Calibrator N lots.Customers were instructed to retain the letter with their laboratory records, and forward to those who may have received the affected product.Customers with questions were instructed to contact their Siemens Customers Care Center or their local Siemens technical support representative.For questions regarding this recall call 508-668-5000. |  | 1294 | 1294 (144 units US) | Worldwide Distribution - USA (nationwide) and Internationally to Argentina, Australia,Brazil,Canada, Columbia, China, Ecuador, Egypt, India, Japan, Korea, Mexico, Saudi Arabia,Singapore, Thailand, and Uruguay. | Open, Classified |  |  | 1294 | Undetermined
Z-0927-2014 | 66946 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | CARESCAPE Monitor B850, a multi-parameter high acuity patient monitor intended for use in multiple areas within a professional healthcare facility. |  |  | December 20, 2013 | 2013 | GE Healthcare, LLC | There is a potential for communication loss associated with the F5-01 Frame when connected to CARESCAPE Patient Data Module (PDM) in the CARESCAPE Monitor B850. The PDM patient parameter acquisition module can sometimes spontaneously lose the communication with the host patient monitor when attached to the F5-01 Frame due to an issue with the PDM Ethernet communication hardware. The communicatio | Component change control | GE Healthcare sent an "Urgent Medical Device Correction" letter dated December 20, 2013 to all affected customers. The letter was addressed to Healthcare Administrator / Risk Manager, Chief of Nursing and Director of Biomedical Engineering. The letter addressed the Safety Issue, Safety Instructions, Affected Product Details, Product Correction and Contact Information (Call Technical Support at 1-800-558-7044 or your local Service Representative). |  | 12 | 12, 945 units (6,239 US, 6,706 OUS) | Worldwide Distribution - US Nationwide including Puerto Rico and the countries of: Australia, Austria, Azerbaijan, Bahrain, Belgium, BRAZIL, Canada, CHILE, COSTA RICA, Cyprus, Czech Republic, Denmark, ECUADOR, Egypt, Estonia, Finland, France, Germany, Greece, GUAM, India, INDONESIA, Iraq, Ireland, Israel, Italy, Japan, Kazakhstan, Korea, Republic of, Kuwait, Lithuania, Malaysia, Malta, Mexico, NEPAL, Netherlands, New Zealand, Norway, Pakistan, PANAMA, PARAGUAY, PERU, Poland, Portugal, Qatar, Romania, Russian Federation, Saudi Arabia, Singapore, Slovenia, South Africa, Spain, Sweden, Switzerland, Taiwan, Tanzania, United Republic of, Turkey, United Arab Emirates, United Kingdom, URUGUAY. | Open, Classified |  |  | 12 | Undetermined
Z-0887-2014 | 67147 | No | HQB | Ophthalmic photocoagulator. | Ophthalmic | Ophthalmic | 510(k) | 27ga ONE-STEP (tm) Endo Illuminator, Synergetics, inc., length 7.5 ft, sterile / EO, 56.02.27P. Packed 12 units per box. Manufactured for Synergetics, Inc., O'Fallon, MO. |  |  | December 09, 2013 | 2013 | Synergetics Inc | Faulty seals on the outer pouch of various, double-pouched, single-use, sterile, ophthalmic devices could result in seal failure. | Use error | Domestic customers were sent recall notices by FedEx on December 9, 2013. Foreign customers were sent recall notices by e-mail between December 6 and December 13, 2013. The recall notice instructed customers to return the recalled product. |  | 96 | 96 units | Worldwide Distribution-USA (nationwide) and the countries of Argentina, Austria, Brazil, Canada, China, Egypt, France, Germany, Greece, Hong Kong, Italy, Japan, Mexico, Morocco, the Netherlands, Pakistan, South Korea, Spain, Sweden, Switzerland, and the United Kingdom. | Terminated | March 25, 2015 | 471 | 2855 | Undetermined
Z-0924-2014 | 67133 | No | JJY | Quality control material (assayed and unassayed). | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | ABX PENTRA Urine Control L/H. |  |  | December 18, 2013 | 2013 | Horiba Instruments, Inc dba Horiba Medical | HORIBA Medical is informing all ABX PENTRA Urine Control L/H customers running the Creatinine assayon ABX PENTRA Chemistry Analyzers that the Target Range for Creatinine has been updated. The ABX PENTRA Urine Control L/H lot specific annexes have been updated toinclude revised acceptable Target Ranges for Creatinine.No other analyte valuesare affected. | Pending | HORIBA Medical sent an Urgent Safety Notice dated December 18, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to update their Quality Control procedures to include the new Creatinine TargetRanges for the ABX PENTRA Urine Control L/H Reference #: A11A01674 and Lot #s: 1207001, 1303601, and 1300701. Customers were instructed to complete and return the enclosed response form to verify their understanding of the Field Safety Notice. Customers with questions were instructed to contact their local HORIBA Medical representative.For questions regarding this recall call 949-453-0500. |  | 531 | 531 | Worldwide Distribution - USA (nationwide) and Canada. | Terminated | October 09, 2014 | 295 | 531 | Undetermined
Z-0919-2014 | 67275 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | IMPAX RIS QDOC 5.8 |  |  | January 08, 2014 | 2014 | AGFA Corp. | Patient name displayed (printed) on the Patient Report was the wrong patient name. | Unknown/Undetermined by firm | On January 8, 2014, an "URGENT FIELD SAFETY NOTICE" (UFSN) letter was emailed to consignees. The letter described the safety alert and mitigation, including directions to consignees as to how the disable Microsoft Word Background Printing. Acknowledgment via FAX-Back or email, that the information was received and understood was requested from the consignees. |  | 14 | 14 | CA, FL, GA, ME, OH, OR, SC, SD, TX and Canada | Terminated | February 28, 2014 | 51 | 14 | Undetermined
Z-0921-2014 | 67172 | No | MEB | Infusion pump. | General Hospital | General Hospital | 510(k) | B. Braun AccuFlo Elastomeric Infusion Pump 200ML/H 100 ML |  |  | December 13, 2013 | 2013 | B. Braun Medical, Inc. | Internal testing of two lots of the AccuFlo CT-2000-100 Elastomeric Infusion Pump identified flow rates which are outside the specified infusion time and may result in overly rapid infusion of medications. | Process design | B. Braun Medical sent an Urgent Medical Device Recall letter dated December 13, 2013 to affected customer. The letter identified the product the problem and the action needed to be taken by the customer. The letter was sent to Progressive Medical Inc., the Importer and B. Braun's sole customer.WHAT ACTIONS SHOULD I TAKE?1. Further use and/or distribution of the affected product should be discontinued immediately. Please identify and quarantine any impacted materials in your possession.2. Review this recall notification in its entirety and ensure that all users within your organization and any consignees your organization may have are informed of this recall. If you are a distributor, please forward this recall notice to any customers who are in receipt of the impacted product.3. Determine your current inventory of the affected lots within your facility. Do not destroy any affected product.4. Utilizing the attached "Recall Acknowledgement" form record the number of full cases and individual units (for any loose units or partial cases) in your possession. If you have no inventory, please indicate that you have zero (0) inventory on the "Recall Acknowledgement" form.5. Return the completed "Recall Acknowledgement" form to B. Braun Medical Inc. Quality Assurance who will administer the recall on behalf of B. Braun Melsungen AG. The form must be e-mailedto PA_QualityAssurance.BBMUS.Service@bbraun.com or faxed to (610) 849-1197 within two (2) weeks of receipt even it the inventory in your possession is zero.6. Instructions for Credit, Replacement, Et Return:a.If you have purchased the product from a third-party and have any full cases, partial cases, or unused individual pieces in your possession please contact the seller to obtain instructions for handling these units.b.If you have purchased the product directly from B. Braun Medical Inc. and have any full cases, partial cases, or unused individual pieces of these affected products, ple |  | 30240 | 30, 240 units | US Distribution one importer/wholesale distributor in MO. | Terminated | September 03, 2014 | 264 | 30240 | Undetermined
Z-0870-2014 | 67245 | No | MDB | Microbial growth monitor. | Microbiology | Microbiology | 510(k) | BD BACTEC FX 40 instrument, catalog number 442296, manufactured by BD Diagnostic Systems, 7 Loveton Circle, Sparks, MD 21152. |  |  | December 09, 2013 | 2013 | Becton Dickinson &amp; Co. | Improperly functioning component of a diagnostic medical device may cause false negative results in clinical specimens. | Nonconforming Material/Component | BD Diagnostics initiated a field correction by phone on December 9, 2013, to affected customers and followed the calls with letters, e-mails, or fax. BD Diagnostics sent a secondary Urgent Field Corrective Action letter dated November 2013 to all affected customers. The recalling firm Secondary notifications were serial number specific and explained that some BD BACTEC FX instruments contain racks that may demonstrate variable indicator light intensities. Specifically, the Anonymous vial indicator lights which should appear yellow may instead appear more green in color. BD noted that affected instruments continue to function properly and will not report an incorrect result. The issue is limited to the yellow indicator lights only. Red and green indicator lights appear as intended. BD is developing software to correct this issue and anticipates that the new software will be available to begin instrument updates by February, 2014. The new software will not affect instrument if the indicator lights are functioning as intended. Until the software can be updated on your instrument, BD provided workflow instructions to ensure that the indicator light issue does not occur.For questions regarding this recall call 410-316-4054. |  | 2 | 2 units | Worldwide Distribution - USA (nationwide) and Internationally to IN, AU, CN, AR, BE, PE, KR, BR, CO, SG, TH, BN, HK, JP, and CA. | Terminated | August 20, 2014 | 254 | 200 | Undetermined
Z-0914-2014 | 67225 | Yes | BTM | Manual emergency ventilator. | Anesthesiology | Anesthesiology | 510(k) | Stand-Alone Resuscitation Unit, Bag and Mask, part number M1226444, labeled in part ***GE Healthcare 8880 Gorman Road Laurel, MD 20723***www.gehealthcare.com***. |  |  | December 12, 2013 | 2013 | GE Healthcare | Medical device adjustment knob may be loose and affect the accuracy of blended Air/Oxygen gas mixture to patients. | Process control | GE Healthcare notified end users with a letter flagged as Urgent Medical Device Correction letter on 12/12/2013. The notification advised that GE Healthcare has become aware of a potential safety issue associated with the loosening of the blender knob shaft on all Resuscitation units. The firm advised that accounts ensure that all potential users in their facility are made aware of this safety notification and the recommended actions described in the notification:1. Perform blender accuracy check as outlined in Section 2.5-2.5.1of the Service Manual to verify that the O2 concentration set at21% is within 16-26% O2 and the concentration at 100% settingis within the recommended range (95-105% O2).2. If the blender accuracy test results do not meet specifications,remove the unit from patient use and contact a GE HealthcareService Representative at the number listed below.3. If the blender accuracy check passes in Step 1 above, you maycontinue to use your resuscitation system. The following precautionsare recommended to ensure blender accuracy is maintained betweensystem checks.Âż Use an independent oxygen analyzer, as determined by your hospitalpolicy, to check blended Air/O2 accuracy.Âż When delivering Oxygen, use a pulse oximeter to monitor your patient.Customers were requested to contact a GE Healthcare Service Representative at 1-800-345-2700 to arrange for correction if the device is found to be defective. |  | 15 | 15 units | Worldwide Distribution-USA (nationwide) including Puerto Rico and the countries of Uruguay, Venezuela, Albania, Algeria, Australia, Austria, Bahrain, Belarus, Belgium, Bolivia, Canada, Chile, China, Costa Rica, Cyprus, Czech Republic, Denmark, Egypt, Estonia, Finland, France, Germany, Greece, Guatemala, Honduras, Hong Kong, Hungary, India, Indonesia, Iraq, Ireland, Israel, Italy, Kazakhstan, Korea, Kuwait, Lithuania, Macedonia, Malaysia, Mexico, Morocco, Netherlands, New Zealand, Nigeria, Norway, Peru, Philippines, Poland, Portugal, Romania, Russia, Saudi Arabia, Singapore, Slovakia, South Africa, Spain, Sweden, Switzerland, Thailand, Tunisia, Turkey, United Arab Emirates, and United Kingdom. | Terminated | June 01, 2015 | 536 | 12201 | Undetermined
Z-0922-2014 | 67378 | No | DZC | Orthodontic appliance and accessories. | Dental | Dental | 510(K) Exempt | SuperElastic Regular force NiTi lower or upper Natural arches. |  |  | December 13, 2013 | 2013 | Highland Metals, Inc. | Labeling issue: Product labeled as "Upper" may contain "Lower" archwires with a red midline mark. Product labeled ad "Lower" may contain "Upper" archwires with a black midline mark. | Employee error | Highland Metals sent an Urgent Field Safety Notice with a response form to all affected consignees on December 13, 2013, via email.. The letter identified the product, the problem, and the action to be taken by the consignee. Consignees were instructed to pass the Field Safety Notice to all those who needed to be aware within their organization or to any organization where the potentially affected devices have been transferred. Consignees were instructed to complete the Acknowledgement of Receipt and Action form and fax to 1-408-271-2962 or email to quality@highlandmetals.com. For questions regarding this recall call 408-271-2955. |  | 9 | 9 packages of each lot number. | Worldwide Distribution - USA including IL, KY, MI, NY, CA, FL, LA and Internationally to Switzerland., Denmark, Greece, Australia, England, and Canada. | Terminated | February 04, 2014 | 53 | 46 | Undetermined
Z-0906-2014 | 67215 | No | LPH | Hip joint metal/polymer/metal semi-constrained porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | ***REF 186-01-08***InteGrip ACETABULAR AUGMENT***SMALL 8mm. Use with 48/50mm Shell.***EXACTECH, Gainesville, FL 32653-1630. |  |  | September 11, 2013 | 2013 | Exactech, Inc. | Exactech is recalling the InteGrip Acetabular Augments due to an out of range condition for an in-vitro biological evaluation standard. | Employee error | Exactech sent an "Important Product Recall Notice" dated September 11, 2013, to all affected customers. The letter identified the product and the action needed to be taken by the customers. Customers were instructed to cease distribution or use of the products. Extend this information to your accounts that may have this product in their possession. Verify if you have any of the subject InteGrip Acetabular Augments (catalog numbers 186-01-08 to 186-05-13) in the specified lots. Complete and fax back the attached form. Further questions, please call 1-800-392-2832. |  | 235 | 235 | USA Distribution including the states of : FL, VA, NY, OH, ME, TX, CO, and GA. | Terminated | August 04, 2014 | 327 | 235 | Undetermined
Z-0866-2014 | 67203 | No | HXX | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | SET SCREW RETAINING DRIVER, REF 9010000820, QTY: 1 EA, Medtronic Sofamor Danek, USA, Inc. |  |  | January 03, 2014 | 2014 | Medtronic Sofamor Danek USA Inc | The retaining tabs component of the Set Screw Retaining Drive may be oversized, which will not allow proper mating between the Set Screw Retaining Drive and the set screw. | Reprocessing Controls | The recall notices were mailed directly to Risk Managers and Sales Representatives, with Recall Questionnaires and Confirmation forms included, on 01/03/2014. |  | 7 | 7 units | US Distribution in the states of CA, NC, TX, IN, NJ, and MI. | Terminated | May 16, 2014 | 133 | 7 | Undetermined
Z-0918-2014 | 67362 | Yes | NPT |  |  | Cardiovascular | PMA | Crimper Model 9100CR26 |  |  | January 20, 2014 | 2014 | Edwards Lifesciences, LLC | Edwards Lifesciences is recalling two lot numbers of the 26mm transcatheter heart valve crimper because they may not open sufficiently when the handle isrotated in the full open position. As a result, user may have difficulty fitting a 26mm Sapien transcatheter heart valve into the crimper orifice when preparing the valve for crimping. | Nonconforming Material/Component | Edwards Lifesciences issued an Urgent Product Notice to all customers. Letters were sent via Fed Ex on 01/20/14. The letter outlines steps to be taken to enable a valve to fit into the open orifice, and allow the valve to be crimped according to normal instructions. Customers were instructed to forward the Urgent Product Notification letter to necessary personnel within their organization. Customers were also instructed to acknowledge they have reviewed and understand the notification letter by signing and dating the form included with the letter. Customers with questions were instructed to contact their Edwards Clinical Specialist. |  | 422 | 422 | Nationwide Distribution | Terminated | March 24, 2014 | 63 | 422 | Undetermined
Z-0925-2014 | 67346 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | SmartSite Needleless Connector, Model No. 2000E |  |  | January 22, 2014 | 2014 | CareFusion 303, Inc. | CareFusion is recalling the SmartSite Needlefree Connector, model number 2000E, because of connection issues. The affected SmartSite connector lots may unintentionally disconnect from a female luer, may be difficult to disconnect from a female luer, or may fail to disconnect from a female luer once attached. Leakage may also be observed if the connector disconnects from the female luer during in | Under Investigation by firm | The firm, CareFusion, sent an "URGENT: Medical Device Recall Notification" letter dated January 22, 2014 to customers. The letter described the product, problem and actions to be taken. Customers were instructed to discontinue use of the listed lot numbers, contact CareFusion for a replacement of the affected lots, and if you have inventory of the listed products immediately complete and return to CareFusion the enclosed, pre-addressed and postage paid, Recall Response Card. All recalled product should be returned directly to the distributor from whom it was purchased. If you have any question, contact your distributor directly or call the CareFusion Support Center at numbers below.For recall related questions, customers are instructed to contact CareFusion Support center at (888) 562-6018. For adverse event reports, customers are instructed to contact Customer advocacy at (888) 812-3266 or email customerfeedback@carefusion.com. For technical questions regarding the Alaris System, customers are instructed to contact technical support at (888) 812-3229. |  | 3500000 | 3,500,000 units (2,136,200 units in US; 1,404,700 units foreign) | Worldwide Distribution: US (nationwide) including Puerto Rico; and internationally to: Australia, Canada and New Zealand. | Terminated | August 06, 2015 | 561 | 3500000 | Undetermined
Z-0867-2014 | 66055 | No | CDP | Total triiodothyronine test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | T3F, Free T3 Assay, REF/Catalog Number LKF31, Siemens Material Number (SMN) 10381626; an IVD Immunoassay kit for use with the IMMULITE/IMMULITE 1000 Analyzers --- Shipping or unit package: 100, 200, 500 and 600 test kits --- Origin: UK Siemens Healthcare Diagnostics Products Ltd. Llanberis, Gwynedd, LL55 4EL UK. |  |  | August 27, 2013 | 2013 | Siemens Healthcare Diagnostics | Siemens Healthcare Diagnostics confirmed customer complaints regarding an increase in the number of euthyroid patients (those with normal function of the thyroid) demonstrating values above the recommended normal range as published in the Instructions For Use for the IMMULITE/IMMULITE 1000 and/or IMMULITE 2000/IMMULITE 2000 XPi for several lots of Free T3 kits. A positive bias in quality control | Nonconforming Material/Component | Urgent Medical Device Recall Letters (dated 8/27/13) and Field Recall Effectiveness Check Forms were sent to the affected customers (starting on 8/27/13) informing them to discontinue use of and discard the IMMULITE Free T3 reagent kit lots. Customers were instructed to contact their local Siemens representative for assistance. Customers are requested to complete and return the Effectiveness Check Form included with the recall letter within thirty (30) days. |  | 2201 | 2201 kits in total | Worldwide Distribution-USA (nationwide) and the countries of Albania, Algeria, Argentina, Australia, Austria, Bangladesh, Belgium, Bulgaria, Bosnia-Herzegovina, Brazil, Belarus, Canary Islands, China, Colombia, Costa Rica, Croatia, Cyprus, Czech Republic, Dominican Republic, Ecuador, Egypt, El Salvador, Estonia, France, Germany, Great Britain, Greece, Guatemala, Honduras, Hungary, Iceland, India, Indonesia, Ireland, Israel, Iran, Italy, Jordan, Republic of South Korea, Kuwait, Kazakhstan, Latvia, Lebanon, Libya, Lithuania, Macedonia, Mexico, Malaysia, Malta, Nepal, Netherlands, Norway, Panama, Peru, Philippines, Pakistan, Paraguay, Poland, Portugal, Qatar, Romania, Russian Federation, Saudi Arabia, Serbia, Singapore, Slovakia, Spain, Sri Lanka, South Africa, Switzerland, Turkey, Thailand, Taiwan, Uruguay, Usbekistan, Venezuela, Vietnam, and United Arab Emirates. | Open, Classified |  |  | 2201 | Undetermined
Z-0974-2014 | 65753 | No | KKX | Surgical drape and drape accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | 5 x 93 US Probe Cover CMS-6049, Sterile, Single Patient Use. Custom Medical Specialties, Inc., Pine Level, NC, Federal (USA) law restricts this device to the sale by or on the order of a physician. |  |  | July 12, 2013 | 2013 | Custom Medical Specialties, Inc. | No endotoxin (LAL) testing to proper level. | Under Investigation by firm | Microtek - Ecolab sent a Urgent Medical Device Recall Letter dated July 12, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Our records indicate that you purchased this product CMS-6049 and this recall is dating back as far was 2008. Attached is a list affected lot number, however, all lots of the CMS-6049 are being recalled. Please inspect your inventory and return any product to us.To prevent any potential patient harm, discontinue the use of the probe cover.In order to advise the Food and Drug Administration about the effectiveness of this recall, you are requested to complete and return the enclosed questionnaire promptly fax or email it no later than July 29, 2013. It is necessary to report to us even if all products have been used.Any product still in your inventory is to be returned to us using a Return Authorization Number and our shipping number which will be given to you by calling (919) 202-8462Ext. 205. |  | 2610 | 2,610 | US Distribution including the states of CT, FL, IL, NJ, NY, PA, TN, VA and WI. | Terminated | April 08, 2014 | 270 | 2610 | Undetermined
Z-0983-2014 | 67387 | No | MWI | Cardiac monitor (including cardiotachometer and rate alarm). | Cardiovascular | Cardiovascular | 510(k) | Invivo Corporation Expression Information Portal (Model IP5) is intended to be used as a secondary display control unit, where the device monitors the vital signs of a patient undergoing an MRI procedure from the MR control room, induction room, or recovery room and then outputs the measured patient data to a hospital information system (HIS). |  |  | December 04, 2013 | 2013 | Invivo Corporation | Invivo Corporation located at 12151 Research Parkway, Orlando, FL 32826 is recalling the Expression Information Portal (Model IP5). Upon power up, the IP5 may fail to produce audio. | Component design/selection | Philips Healthcare sent an "URGENT - MEDICAL DEVICE RECALL" letter dated December 4, 2013, to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Customers were instructed to refer to the "Action to be Taken by Customer/User section on the attached pages. A Philips representative will contact the customer regarding their affected devices. All affected devices will have a firmware update installed to correct the problem free of charge. If you need any further information or support concerning this problem, please contact your local Philips representative: 877-468-4861 option 1, then option 2.Philips apologizes for any inconveniences caused by this problem. |  | 299 | 299 | Worldwide Distribution - USA (nationwide) and the countries of Australia, Austria, Denmark, France, Germany, Netherlands, Poland, Singapore, Sweden, Switzerland and United Kingdom. | Open, Classified |  |  | 299 | Undetermined
Z-0833-2014 | 67171 | No | KCT | Sterilization wrap. | General Hospital | General Hospital | 510(k) | Symmetry and CODMAN brands, QUAD-LOCK Sterilization Container. Sterilization container systems manufactured and distributed under the following names: |  |  | December 05, 2013 | 2013 | Symmetry Medical/SSI | The Quad-Lock Sterilization Container System may not maintain an effective sterile barrier. | Process control | The firm notified their consignees of the recall by letter and email on 12/02/2013 instructing customers not to use the containers until further notice from Symmetry Surgical. |  | 2991 | 2,991 units | Worldwide Distribution - US (Nationwide) and PR; and Internationally to: Italy, Australia, Saudi Arabia, UK, China, Canada, Germany, New Zealand, Colombia, Korea, and Chile. | Terminated | April 13, 2015 | 494 | 2991 | Undetermined
Z-0984-2014 | 66958 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Access 2 Immunoassay System Analyzer. |  |  | November 26, 2013 | 2013 | Beckman Coulter Inc. | The Access 2 Immunoassay System Analyzer and the UniCel DxI 600 Access may have been manufactured with non-conforming bushings. Non-conforming bushings have a remote probability of generating incorrect results. | Nonconforming Material/Component | Beckman Coulter initiated this recall by sending recall notifications to customers. An Urgent Product Correction letter sent to affected customers. The recall letter dated November 22, 2013 informed customer of the recall, products with description and codes, issue, impact, action, resolution, and contact information. |  | 3 | 3 units | Worldwide Distribution: US Distribution in South Dakota and the countries of Germany, Lebanon, and China. | Terminated | July 27, 2014 | 243 | 4 | Undetermined
Z-1089-2014 | 66660 | No | KWQ | Spinal intervertebral body fixation orthosis. | Orthopedic | Orthopedic | 510(k) | Azure Anterior Cervical Plate System, 2-Level Plate, 36 mm, Implant Grade Titanium and Nitinol, Orthofix Inc., RX Only |  |  | October 29, 2013 | 2013 | Orthofix, Inc | There is a possibility that the locking mechanism of the Azure Anterior Cervical Plate may fracture during or after implantation. | Component design/selection | Orthofix sent a Urgent Medical Device Recall Notification dated October 29, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Immediate Action:1. Immediately cease any further distribution or use of the AZURE" plates.2. Notify your surgeons and staff customers that Azure" plates cannot be used and are to be returned to Orthofix. at no cost to you, see Instructions below.3. Please complete the attached reply form indicating that you have received this letter.Product Return Instructions: To return the Azure plates, please follow these Instructions. 1. Check you inventory of Azura Implant Trays..2. Complete the attached form with the quantity you are returning.3. Contact your local Orthofix representative or Orthofix customer servicerepresentative at 888--298-5700 for return instructions and shipment information and replacement with a different Orthofix Anterior Cervical Plate System.. |  | 4172 | 4,172 units | Worldwide Distribution: USA (nationwide) and the country of Germany. | Terminated | December 09, 2014 | 406 | 4172 | Undetermined
Z-0815-2014 | 66991 | No | FOZ | Intravascular catheter. | General Hospital | General Hospital | 510(k) | Titanium Single Lumen Low-Profile port, with Pre-Attached open-Ended Silicone 6.6Fr Catheter. Product Code 0602180 |  |  | December 06, 2013 | 2013 | Bard Access Systems | Bard Access Systems is recalling various Low Profile Port with 6.6 Fr Single Lumen Broviac Catheter because a small number of Silicone catheters in the may be undersized to 5.3 Fr instead of 6.6 Fr as stated in the labeling for the device. | Process control | BARD sent an Urgent Product Recall letter dated October 2013 to affected customers via certified letter. The letter identified the affected product, problem and actions to be taken. Customers were asked to check all inventory locations, remove and return any unused affected prodcut to Bard Access Systems through return goods authorization number RAP80498 for a credit or replacement, and to complete the Reply Form and Inventory Reconcilation Form. For questions contact BAS Customer Service at 1-800-290-1689. |  | 200 | 200 | Worldwide Distribution - US nationwide, Belgium, and Canada. | Terminated | August 05, 2014 | 242 | 1558 | Undetermined
Z-0811-2014 | 64966 | No | HQF | Ophthalmic laser. | Ophthalmic | Ophthalmic | 510(k) | 1) The FEMTO LDV Z2 Femtosecond Surgical Laser is an ophthalmic surgical laser indicated for use in the creation of corneal incisions in patients undergoing LASIK surgery or other treatment requiring lamellar resection of the cornea at constant depth relative to the corneal surface. |  |  | April 16, 2013 | 2013 | Ziemer Usa Inc | A defect in the controller board can lead to the condenser to the video monitor becoming too warm. This could potentially cause the video monitor to shut off. | Component design/selection | A recall notification letter dated April 16, 2013 was mailed to direct accounts. The letter notified the direct accounts of the problem and advised them that a technician would be arriving at their facility during the week of June 2, 2013 to replace the defective controller board. The technician will also perform preventative maintenance on the device at the same time. Direct accounts are advised to contact Ziemer USA, Inc. at 618-462-9301 if they have any questions about this issue. |  | 5 | 5 units | Worldwide Distribution: US (nationwide) including states of: NJ and IL; and Internationally to: Finland, Japan, and Germany. | Open, Classified |  |  | 5 | Undetermined
Z-0826-2014 | 67122 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Port Kit with Access Infusion Set under the following labels: 1) 8F/SL/Plastic Port/F, Product Number H965450140, Catalog Number 45-014; |  |  | November 25, 2013 | 2013 | Navilyst Medical, Inc. | Navilyst Medical has recalled a Port Kit with Access Infusion Set due to a possible breach of sterilization barrier. | Under Investigation by firm | Navilyst Medical, Inc., sent an Urgent Medical Device Recall - Immediate Action Required letter dated November 25, 2013 to their customers via Federal Express. The letter identified the affected product, problem and actions to be taken. Customers are instructed to segregate and return all affected devices to Navilyst Medical, Inc. Distribution Center Attn: QA Returns 10 Glens Falls Technical Park Glens Falls, NY 12801 and complete and return the Reply Verification Tracking Form provided in the recall notification via Fax to: 1-800-782-1357. For questions call 1-800-833-9973 between 8:30 am and 7:00 pm (EST). |  | 920 | 920 Total kits | US Nationwide Distribution in the states of CA, CT, DC, MA, NJ, NY, PA, MD, VA, WV, NC, GA, FL, TN, OH, MN, MT, IL, MO, NE, OK, TX, NM, and WA. | Open, Classified |  |  | 920 | Undetermined
Z-1076-2014 | 67235 | No | JIX | Calibrator. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Dimension Vista Enzyme 1 Calibrator (ENZ 1 CAL) lots 3FD034 and 3HD012. |  |  | January 06, 2014 | 2014 | Siemens Healthcare Diagnostics, Inc. | Complaints were received regarding Quality Control (QC) shifts with Pseudocholinesterase (PCHE) and imprecision during PCHE calibration with the Dimension Vista Enzyme 1 Calibrator (ENZ 1 CAL). | Nonconforming Material/Component | A recall notification letter, dated January 2014, was sent to Users. |  | 3496 | 3496 | Nationwide Distribution. | Terminated | May 06, 2015 | 485 | 3496 | Undetermined
Z-1075-2014 | 67066 | No | CHL | Blood gases (PCO2, PO2) and blood pH test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | pCO2 D788 Membranes. Distributed by Radiometer America, Inc., Westlake, OH. |  |  | July 05, 2013 | 2013 | Radiometer America Inc | RADIOMETER has become aware that some D788 pCO2 membranes can cause biased measurement results on patient results as well as QC results. | Under Investigation by firm | On 7/18/2013, the firm sent Product Notification letters to their customers. |  | 38890 | 38,890 units (7,860 units imported into the US) | Worldwide Distribution-USA (nationwide) including DC and the states of AK, AL, AR, AZ, CA, CO, CT, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV &amp; WY, and the countries of Bermuda and Canada. | Terminated | September 19, 2015 | 806 | 38890 | Undetermined
Z-0824-2014 | 67209 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | ACUSON SC2000 Ultrasound System; |  |  | December 22, 2013 | 2013 | Siemens Medical Solutions USA, Inc. | The user interface assembly on the ACUSON SC2000 may become loose with the potential for the entire module to fall off of the ultrasound system. | Nonconforming Material/Component | Siemens sent an Customer Safety Advisory Notification letter on December 22, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customers. Customers were instructed to:.If you observe looseness in the attachment of the user interface module to the column on the ultrasound system, immediately stop using your system and contact your Siemens service representative for repair.Your local Customer Service Engineer will contact you to make the necessary repairs on your system. You will be contacted to schedule an appointment. If you have any questions, please contact your local service support person for information regarding timelines and status.Until the necessary repairs have been completed on your system, please share this information with all personnel within your organization will need to be aware of thls issue. As always, patient safety issues are a high priority, and we thus alert customers of product concerns which may occur undervarlous conditions. To date, no patient injury has been reported.This problem was discovered as part of our ongoing quality process. ÂżÂż ÂżÂżWe sincerely regret any inconvenience this condition may cause in your daily operations. For further questions please call (650) 694-5398. |  | 843 | 843 active units | Worldwide Distribution - US (nationwide) and the countries of Albania, Australia, Austria, Bangladesh, Belgium, Bosnia Herzeg, Brazil, Brunei, Canada, Columbia, Denmark, Ecuador, Egypt, Finland, France, Germany, Iceland, India, Israel, Italy, Japan, Kuwait, Malaysia, Mauritius, Mexico, Netherlands, Norway, Oman, P.R. China, Philippines, Poland, Qatar, Rep of Yemen, Republic Korea, Romania, Russian Fed. San Marino, Saudi Arabia, Singapore, Slovakia, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, U.A.E. United Kingdom and Venezuela. | Terminated | May 06, 2015 | 500 | 843 | Undetermined
Z-1077-2014 | 67159 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | PFC SIGMA Cruciate Retaining (CR); Cemented Femoral devices; (Size 5/RT) |  |  | December 09, 2013 | 2013 | DePuy Orthopaedics, Inc. | Knee implant femoral devices were found to have anomalous microstructure. Porosity and the presence of blocky carbides, within the cobalt chrome part, can decrease the overall fatigue strength. This decrease in strength could cause implants to fracture under what would normally be acceptable loads. | Process design | The recall is extended to the DePuy Distributor, Hospital, and Surgeon levels. The two affected distributors were contacted by telephone. Written communication to the two hospitals and one surgeon who received the devices will be delivered by DePuy Orthopaedics, Inc. via email or regular mail. The one affected surgeon will also be contacted by telephone by the Medical Safety Officer and a metallurgist, who will discuss the surgeon communication with the surgeon.The sales representatives aid customers in the affected device returns, as needed. The devices will be returned through the normal DePuy Returns process, to attention of Returns and marking H13-30C on the outsideof the box. All (21) remaining devices have been verified as being returned and in quarantine.Effectiveness will be determined by product reconciliation, receipt of all hospital reconciliation forms, and receipt of international declaration from the International Affiliate. DePuy will follow-up with the affected hospitals until all hospital reconciliation forms are returned, and DePuy will follow-up with the International Affiliate until all international actions are complete and the international declaration is completed and return. |  | 2 | 2 | Worldwide Distribution-USA including the states of IA, CA, KY, and FL, and the countries of Finland, Germany, Sweden, and Czech Republic. | Terminated | June 10, 2014 | 183 | 2 | Undetermined
Z-0963-2014 | 67266 | Yes | BTR | Tracheal tube. | Anesthesiology | Anesthesiology | 510(k) | ISIS HVT Tracheal Tube Cuffed with Subglottic Secretion Suction Port and Preloaded Stylet, 9.0, Rx only, |  |  | January 06, 2014 | 2014 | Teleflex Medical | The tracheal tube can kink during patient use. | Other | Teleflex Medical sent an Urgent Medical Device Recall Notification letter dated January 6, 2014, to all affected consignees. The letter identified the product, the problem, and the action to be taken by the consignee. Consignees were instructed to immediately discontinue use and quarantine any affected product. Return the product by completing the enclosed Recall Ackknowledgment Form and fax it to 1-866-804-9881, Attn: Customer Service to receive a Return Goods Authorization (RGA) number and instructions for the return of the affected product to Teleflex Medical. Consignees with no affected stock, should also complete the enclosed Recall Acknowledgment Form and fax it to 1-866-804-9881. Consignees with questions were instructed to call 1-866-246-6990.For questions regarding this recall call 919-572-8014. |  | 52489525 | 190,871 each (148, 275 US) | Worldwide Distribution - USA (nationwide) Puerto Rico and Internationally to Canada, Barbados, Buenos Aires, Bolivia, El Salvador, Mexico, Columbia, Chile, Costa Rica, Dominican Republic, Ecuador, Barbados, Belgium, South Korea, Thailand, New Zealand, Netherlands. | Terminated | June 17, 2016 | 893 | 52489525 | Undetermined
Z-0830-2014 | 67176 | No | GFF | Surgical instrument motors and accessories/attachments. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Anspach REF: TURQ-434-4 Twist Drill for 1.5mm x 4mm screw. Rx only |  |  | March 31, 2011 | 2011 | The Anspach Effort, Inc. | Cutter was longer than was stated on the label | Pending | Anspach sent an Urgent Medical Device Product Removal letter dated March, 32, 2013, to all affeced customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to inventory their batches and return the products for replacement. Customers needing assistance were instructed to contact Anspach at 800-327-6887.For questions regarding this recall, call 800-327-6887. |  | 29 | 29 cutting burs | Worldwide Distribution - USA including CA, CO, IN, KY, MD and TX and Internationally to Japan, Taiwan, United Kingdom, Sweden, and Germany. | Open, Classified |  |  | 628 | Undetermined
Z-1066-2014 | 67404 | No | MBL | Hip joint metal/polymer/metal semi-constrained porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Reflection Interfit Shell, 52 MM OD, 3 HOLE, REF 71336052, QTY 1, STERILE R, Smith &amp; Nephew, Inc. |  |  | January 23, 2014 | 2014 | Smith &amp; Nephew Inc | Inner spherical radius of the shell is undersized. | Under Investigation by firm | Smith &amp; Nephew Inc. sent an Urgent - Product Recall 1st Notification letter dated January 23, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customers.Actions for Hospital Representatives:1. Please inspect your inventory and locate any unused devices from the above listed product and batch numbers, and quarantine them immediately.2. Complete the last two columns in the Inventory Return Certification Form on the following page, indicating the quantities that need to be returned, and include your phone and fax number in the spaces provided. If you do not have product to return, please place an x in the column No product to return .3. Please contact Smith &amp; Nephews Safety Affairs Department via email at FieldActions@smith-nephew.com or fax +1-901-566-7975 to obtain a return authorization (RA) number.4. Return any affected product to the address listed on the Inventory Return Certification Form. Please indicate the RA number on your return shipment.Actions for Smith &amp; Nephew Inventory Specialists1. Please inspect your inventory and locate any unused devices from the above listed product and batch numbers, cease distribution and use of those products, and quarantine them immediately.2. Carry out a physical count of all affected product in your territory and record this data on the Inventory Return Certification Form on the following page.3. If you do not have product to return, please place an x in the column No Product to Return , and then fill in the Acknowledgement of Responsibility at the bottom of the form.4. Fax a copy of the completed Inventory Return Certification Form to: +1-901-566-79755. If you have affected products, return the product, along with the completed Inventory Return Certification Form, to the address indicated |  | 10 | 10 units | Worldwide Distribution: US Distribution to NJ and FL; and the countries of: Colombia and Costa Rica. | Open, Classified |  |  | 10 | Undetermined
Z-1069-2014 | 67193 | No | GCJ | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | VirtuoSaphÂż Plus Endoscopic Vessel Harvesting System, Part number: VSP550, Sterile***Rx Only*** |  |  | December 12, 2013 | 2013 | Terumo Cardiovascular Systems Corporation | Sterility of medical devices intended for use in surgical procedures may be compromised. | Process control | Terumo CVS notified afected end users by phone on December 12, 2013 and advised that they had become aware that the sterile packaging barrier was breached in specific lots of VirtuoSaphÂż Plus Endoscopic Vessel Harvesting Systems. Customers were requested to return all unused affected product and verify that all users at their institution have been made aware of the risks associated with using the affected devices. For questions call 734-663-4145. |  | 26 | 26 units | USA Nationwide Distribution in the state of LA, NJ, and NY. | Terminated | July 28, 2015 | 593 | 26 | Undetermined
Z-1054-2014 | 66080 | No | LRO | Surgical drape and drape accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | Enforcement Discretion | STERILE SAMPLE BUILD; REF: SAMP0147; Packaged for Medline Industries, Inc., Mundelein, IL 60060; 1-800-Medline; www.medline.com |  |  | August 14, 2013 | 2013 | Medline Industries, Inc. | Guidewire .035x150 3MMJ TCFC, item number 88241, has a potential for the coating to flake off of the wire. The guidewire is used in a number of surgical convenience kits assembled and marketed by Medline Industries, Inc. | Nonconforming Material/Component | The firm, Medline Industries, Inc., sent a "RECALL IMMEDIATE ACTION REQUIRED" letter dated August 14, 2013, to its direct accounts/customers. The letter described the product, problem and actions to be taken. The customers were instructed to immediately check their stocks for inventory of the affected products, quarantine any affected products, return the product using the enclosed shipping slips for credit-If additional stickers are needed, please call 866-359-1704, and complete and return the enclosed FAX-BACK VERIFICATION FORM listing the quantity of affected product(s) on hand via fax to Medline at 866-767-1290.If you have any questions, please contact 866-359-1704. |  | 5 | 5 units | Nationwide Distribution: US to states of: AL, CA, FL, IL, LA, OH and VA. | Open, Classified |  |  | 2853 | Undetermined
Z-1057-2014 | 67261 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Synthes Driving Cap/Threaded |  |  | January 07, 2014 | 2014 | Synthes, Inc. | The Synthes Driving Cap/Threaded can bind in the IM Nail Radiolucent Insertion Handles when assembled becoming difficult to separate. | Process design | Synthes sent an Urgent Notice - Medical Device Recall Letter dated January 7 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customers.Action to be taken: Our records indicate that your facility has the product(s) subject to this Recall. Synthes asks that you review your inventory and immediately remove the affected lots from stock.Please take the following actions:If you DO have any of the identified devices, please take the following steps:"Call DePuy Synthes at 1-800-479-6329 to obtain a Return Authorization Number."Complete the Verification Section at the end of this letter by checking the appropriate box indicating affected product has been located. Also, please indicate the number of devices found and note the Return Authorization Number. Please include your name, title, telephone number and signature in the spaces provided."Return the Verification Form (page-3 of this letter) with the product to:"Credit/Returns, Synthes, 1101 Synthes Avenue, Monument, CO 80132.If you DO NOT have the identified product, please take the following steps:"Complete the attached Verification Section at the end of this letter by checking the appropriate box indication that no affected product has been located. :Please include your name, title, telephone number and signature in the spaces provided this return documentation acknowledges your receipt of medical device removal information."Return the documents to DePuy Synthes by:" Fax: 877-271-1473 or" Scan/email: FieldAction@synthes.comIf you have any questions,please call 610-719-5450. Thank you for your attention to this issue. |  | 143 | 143 | USA (nationwide) Distribution. | Open, Classified |  |  | 143 | Undetermined
Z-0980-2014 | 66777 | No | NGT | Intravascular catheter. | General Hospital | General Hospital | 510(k) | Merit Medical's Custom Procedural Trays or Kits containing 1% Lidocaine HCl Injection, USP, 10 mg/mL, Lot 25-090-DK. Custom Procedural Tray or kit is a combination of legally marketed medical devices placed into one container for the convenience of the user and used within the limits of the product's intended uses. |  |  | September 30, 2013 | 2013 | Merit Medical Systems, Inc. | Custom Procedural Trays/Kits contain 1% Lidocaine HCl injection which were recalled because one confirmed customer report where visible particulate were found in the primary product container in the form of dark red/black particles. The particular matter may dislodged through the tissue and could possibly cause a pneumothorax or hemothorax. | Other | Merit Medical sent via FedEx an "URGENT PRODUCT RECALL NOTICE" dated September 30, 2013 to all affected customers. The letter identified the affected product, description of problem and recommended user actions to be taken. Consignees were given two options: 1) Based upon immediate patient need, they can keep the affected trays and discard the Lidocaine at the point of use (a) Merit will provide the customer with warning labels, along with replacement Lidocaine for each affected tray based on their current inventory or (b) The label will instruct the user to dispose of the Lidocaine at the point of use and replace it with the new lot. 2) They may return all affected trays remaining in inventory to Merit for credit or replacement. Consignees were requested to complete and sign the Customer Response Form and immediately fax a copy of the completed Customer Response Form to Joyce Braxton at 1-801-316-4852 and mail the completed original Customer Response Form using the provided postage-paid, self-addressed envelope to Merit. If you choose to return the trays, please return the trays as instructed on the Customer Response Form. If customers choose to keep the trays, Merit can send them labels and Lidocaine upon receipt of the Customer Response Form. If you have any questions, please call 1-804-416-1048 and/or 1-801-316-4822.For questions regarding this recall call 804-416-1030. |  | 20872 | 20, 872 trays | Nationwide Distribution including AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, IA, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, NC, ND, NE, NH, NJ, NY, OH, OK, PA, SC, TN, TX, UT, VA, WA, WI, WV. | Terminated | May 13, 2014 | 225 | 20872 | Undetermined
Z-1062-2014 | 67283 | No | HQL | Intraocular lens. | Ophthalmic | Ophthalmic | PMA | Bausch &amp; Lomb AKREOS AO Micro Incision Lens |  |  | October 16, 2013 | 2013 | Bausch &amp; Lomb Surgical, Inc. | Lens was manufactured with incorrect raw material. | Nonconforming Material/Component | The firm, Bausch + Lomb, telephoned and sent an "URGENT - MEDICAL DEVICE RECALL" letter dated October 17, 2013 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to: 1) Determine the disposition of the lenses; 2) Complete and provide the enclosed acknowledgement form to the sales respresentatives collecting the lenses, and 3) Return all unused products. If you have any questions, please contact Bausch + Lomb at (800) 338-2020. |  | 336 | 336 IOLs (283 IOLs in the US, 53 IOLs outside the US) | Worldwide Distribution: US (nationwide) and Internationally to: Great Britain, France, Spain, Portugal, Sweden, Russia and Guadeloupe. | Terminated | September 17, 2014 | 336 | 336 | Undetermined
Z-0995-2014 | 67400 | No | DWF | Cardiopulmonary bypass vascular catheter, cannula, or tubing. | Cardiovascular | Cardiovascular | 510(k) | Medtronic Select 3D Arterial Cannulae. Vented, models 78618 (18 Fr.) and 78620 (20 Fr.). Non-Vented Cannulae, models 78718 (18 Fr.) and 78720 (20 Fr.). Sterilized using Ethylene Oxide. Non-pyrogenic. |  |  | January 30, 2014 | 2014 | Medtronic Inc. Cardiac Rhythm Disease Management | Medtronic is recalling all models of its Select 3D and Select CAP Arterial Cannula product families due to reports of the cannula body splitting near the suture collar. | Under Investigation by firm | The firm, Medtronic, sent an "Urgent Medical Device Recall" letter dated January, 2014 via UPS on January 30, 2014 to their customers. The letter described the product, problem and actions to be taken. The customers were instructed to immediately identify and quarantine all listed product in your inventory; return all listed product to Medtronic and contact Customer Service at 800 854-3570 and reference RADAR #167826 to initiate a return and credit of product; complete and return the attached Customer Confirmation Certificate via fax to Medtronic at 651-367-0612 to the attention of Customer Focused Quality, and share this notification with others in your organization as appropriate.Select 3D and Select CAP Arterial Cannula are not available for purchase at this time. Please contact your Medtronic sales representative or Medtronic LifeLine Technical Services at 1 877-526-7890 or (+1 763) 526-7890 to assist with alternative cannula selection and any additional questions you may have related to this notification. |  | 19483 | 19483 total (7669 US, 11814 OUS) | Worldwide distribution: US (Nationwide) including Puerto Rico, and Internationally to: Algeria, Argentina, Austria, Belgium, Bosnia And Herzegovina, Brazil, Canada, Cyprus, El Salvador, France, Georgia, Germany, Greece, Hungary, India, Ireland, Italy, Japan, Jordan, Kuwait, Netherlands, Oman, Paraguay, Poland, Portugal, Russian Federation, Saudi Arabia, Slovakia, Spain, Sweden, Switzerland, Thailand, Turkey, United Kingdom, Venezuela, and Yemen. | Terminated | October 06, 2014 | 249 | 19483 | Undetermined
Z-0841-2014 | 66931 | [6 empty] | ZYM B Reagent (REF 70493) |  |  | November 08, 2013 | 2013 | Biomerieux France Chemin De L'Or | bioMÂżrieux identified a visual defect and activity issue on the ZYM B reagent (REF 70493); leading to a false negative results of some API biochemical tests using the ZYM B reagent. | Under Investigation by firm | bioMerieux sent an Urgent Product Removal Notice letter dated November 15, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to ensure the letter and guidance were distributed and received by all appropriate personnel in their organization, customers were instructed to discontinue use and destroy ZYM (B 70493), API NH (10400) API Listeria (10300) with ZYM B ampules having a visual defect.Customers still in possession of some kits of ZYM B (70493-Lots: 1) without visual defect that is used with API 20 Strep (20600), API Coryne (20900), API Staph (20500) and API ZYM (25200), have the option to:a) Use the kits of ZYM B (70493) still in your inventory with the following mandatory requirement for the API Strep and API Coryne: perform a daily quality control with the ATCCÂż strain described in the paragraph Reading and Interpretation of the package insert of each API strip prior to using the ZYM B ampules.b) Use the kits of ZYM B (70493) still in your inventory without any requirement with API Staph and API ZYM, orc) Discard all kits of ZYM B (70493) still in their inventory.If you are still in possession of some kits of API NH (10400-Lots: See Annex 1) without visual defect of ZYM B, you have the option to:a) Use the kits of API NH (10400) still in your inventory with the mandatory requirement: perform a daily quality control with the ATCCÂż strain 49917 Haemophilus paraphrophilus described in the paragraph Use of the reagents of the package insert of API NH prior to using the ZYM B ampules or,b) Discard all the kits of API NH (10400) still in your inventory.5. If you are still in possession of some kits of API Listeria (10300-Lots: See Annex 1) without visual defect of ZYM B, we ask you to use the API Listeria (10300) still in your inventory with the mandatory requirement: perform a daily quality control with the ATCCÂż strain 33090 Listeria innocua as |  | 39439 | 39439 kits (includes ZYM B reagents, API Listeria * API NH kits) | Worldwide Distribution - USA (nationwide) and Internationally to U.S. Nationwide; Algerie, Argentina, Armenia, Australia, Austria, Azerbaidjian, Bahrain, Bangladesh, Belarus, Belgium, Benin, Bolivia, Brazil, Brunei, Bulgaria, Burkina Faso, Cambodia, Cameroon, Canada, Chad, Chile, China, Colombia, Democratic Republic of Congo, Costa Rica, Cote d'Ivorie, Germany, Czech Republic, Dominican Republic, Dutch Antilles, Dutch Antilles, Ecuador, Egypt, Republica De El Salvador, Estonia, Germany, Ethiopia, France, Gabon, Georgia, Germany, Greece, Guatemala, Guinea, Herzegovina, Honduras, Hungary, India, Germany, Indonesia, Iraq, Israel, Italy, Jamaica, Japan, Iraq, Kazakhstan, Kenya, Korea, Kosovo, Kuwait, Latvia, Lebanon, Libya, Lithuania, Macedonia, Malaysia, Malaysia, Malaysia, Mali, Malta, Mexico, Romania, Mongolia, Morocco, Mozambique, Singapore, Netherlands, Managua, Niamey Niger, Nigeria, Nigeria, Oman, Pakistan, Panama, Paraguay, Peru, Philippines, Philippines, Poland, Portugal, Puerto Rico, Qatar, Romania, Russia, Uganda, Kingdom of Saudi Arabia, Serbia, Singapore, Slovakia, South Africa, Spain, Sudan, Sweden, Switzerland, Syria, Tanzania, Italy, Thailand, Trinidad &amp; Tobago, Tunisia, Turkey, Uganda, Ukraine, UAE, United Kingdom, Uruguay, Venezuela, Vietnam, Vietnam, Yemen and Zimbabwe | Open, Classified |  |  | 39439 | Undetermined
Z-1466-2014 | 67983 | No | DWF | Cardiopulmonary bypass vascular catheter, cannula, or tubing. | Cardiovascular | Cardiovascular | 510(k) | Edwards Lifesciences Fem-Flex II Arterial and femoral Cannulae, arterial cannula sizes 16, 18 and 20 French, models: FEMII016A, FEMII016AS, FEMII018A, FEMII018AS, FEMII020A, and FEMII020AS. |  |  | April 09, 2014 | 2014 | Edwards Lifesciences, LLC | Potential for separation of dilator tips. | Nonconforming Material/Component | Customers were notified via letter on 4/4/14 to return affected products. |  | 13115 | 13,115 | Nationwide, including foreign, govt/VA/Military consignees. | Open, Classified |  |  | 42000 | Undetermined
Z-1058-2014 | 67090 | No | DZE | Endosseous dental implant. | Dental | Dental | 510(k) | External Implant RBT, D3 Threadform 4.0 X 10.5mm, REF 40105D3, STERILE R, Rx Only, BioHorizons(R), 2300 Riverchase Center, Birmingham, AL 35244 |  |  | December 10, 2013 | 2013 | BioHorizons Implant Systems Inc | The body length of the External Implant 40105D3 Lot 1302762 labeled as 10.5mm in length is in fact 12mm in length. | Mixed-up of materials/components | The firm, BioHorizons, sent an "Urgent Medical Device Recall" letter dated December 10, 2013 to its Consignees. The letter described the product, problem and actions to be taken. The customers were instructed to return unused affected implants and receive replacements free of charge; complete and return the enclosed Customer Reply Form via fax to (205) 484-2195 or scan to PDF and e-mail to implants@biohorizons.com; If you have affected implants, upon receipt of your completed Customer Reply Form, you will be contacted by BioHorizons Customer Care and provided a Return Authorization number. In addition, if you have already placed any affected implants without incident, no special patient action needs to be taken.If you have any questions regarding this information, please contact our Customer Care Department at (888) 246-8338, Monday through Friday, 7:30am to 6:00pm Central Time. |  | 119 | 119 units | Worldwide distribution: US (nationwide) to states of: LA, CA, NC, IL, PA, and NY; and Internationally to: Seoul, Korea; New Delhi, India; Taipei City, Taiwan; and Madrid, Spain. | Terminated | March 28, 2014 | 108 | 119 | Undetermined
Z-1063-2014 | 67446 | apparatus, traction, non-powered | HST | Nonpowered orthopedic traction apparatus and accessories. | Orthopedic | Neurology | 510(K) Exempt | Gelzone Shoulder Sleeve, Catalog Number SLV634, Lot Number 852766. |  |  | January 29, 2014 | 2014 | Implantech Associates Inc | Implantech initiated the voluntary recall of Gelzone Shoulder Sleever, SLV-634, Lot Number 852766 because it may not satisfy customer requirements for the robustness of the Velcro closure. | Nonconforming Material/Component | Implantech sent a Recall Notification letters dated January 29, 2014 via mail and fax to the affected customers. The recall notification letter informed customers of the voluntary recall, product description with code information, reason for recall, contact information, and actions to be taken. The letter also contains a response section to be completed and returned via fax. For questions call 1-800-733-0833 x235. |  | 300 | 300 | Worldwide Distribution - USA (nationwide) in the states of: CA, AZ, IL, NC, CO and the country of: South Africa. | Terminated | March 12, 2015 | 407 | 300 | Undetermined
Z-1045-2014 | 67505 | No | DQK | Programmable diagnostic computer. | Cardiovascular | Cardiovascular | 510(k) | CartoUnivu Module, Model: Carto 3 Software versions 3.2.2 and 3.2.3; Catalog No. KT-5400-124 |  |  | February 12, 2014 | 2014 | Biosense Webster, Inc. | Biosense Webster has initiated a recall of the CartoUnivu Module within the Carto 3 System (V3.2.2 and V3.2.3) when used with Siemens Axiom Artis VB Fluoroscopy systems in particular procedural configurations only. If the table rotation is not returned to zero point, there is a misalignment between the Carto 3 System map display and the fluoroscopic capture. | Other | Biosense Webster sent an Urgent Field Safety Notice dated February 12, 2014, to all affected customers to inform them that Biosense Webster has become aware of an issue with a particular configuration of the CartoUnivu Module withn the Carto 3 System (v3.2.2 and v3.2.3) when used with Siemens Axiom Artis VB Fluoroscopy systems in particular procedural configurations only. The letter informs the customers of the problems identified and the actions to be taken. Customers with questions are instructed to contact their Biosense Webster sales representative or cal (866) 473-7823, Monday through Friday from 7am to 8pm EST. Customers with questions about the acknowledgement form and its return, are instructed to call (909) 839-8468.For questions regarding this recall call 909-839-8500. |  | 8 | 8 units total (5 units in US) | Worldwide Distribution - USA including CA, KY, NJ, MA, and UT and Internationally to Belgium, Austria, and Germany. | Terminated | May 23, 2014 | 100 | 8 | Undetermined
Z-0990-2014 | 67154 | Yes | FKB | Hemodialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Orbit soft Infusion sets, 9 mm soft cannula, P10 (42") tubing; Single-Use Set |  |  | December 20, 2013 | 2013 | Perrigo Diabetes Care | Perrigo was notified by our supplier, Ypsomed AG, that they are recalling these batches due to being manufactured with a non-conforming female luer-lock connector. The luer-lock connectors are incompatible with standard male luer-lock connectors for use in connection to compatible infusion pumps. Perrigo became was notified of a quality complaint related to this issue on November 20, 2013 from a | Nonconforming Material/Component | Perrigo Diabetes Care sent a recall notice dated December 16, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to segregate the affected productt and request return of all undistributed/unused sets from their customer warehouses to Omni Manufacturing Services S.A. de C.V., Parque Industrial Queretaro, La Noria Lt. 13. Mz 3 No. 125, Santo Rosa Jauregui, MX-76220 Santiago de Queretaro. Customers were instructed to distribute the attached "Urgent Field Safety Notice" to all users or patients who have been shipped any products from the affected lots. Customers with questions were instructed to call Perrigo Diabetes Care at 1-888-775-1997. |  | 7924 | 7924 units | Nationwide Distribution including GA &amp; TX | Terminated | August 05, 2014 | 228 | 7924 | Undetermined
Z-0845-2014 | 67319 | No | NCX | Thermal regulating system. | Cardiovascular | Neurology | 510(k) | InnerCool RTx Endovascular System, Model Numbers: 861470 (120V), 861472 (240V). |  |  | January 13, 2014 | 2014 | Philips Healthcare | Phillips Healthcare initiated this action because the Main Control Board (MCB) may produce an inaccurate low patient temperature reading and low temperature alarm. | Nonconforming Material/Component | Philips Healthcare Recall Notice for this proactive performance correction was provided verbally to hospital biomedical engineers. Philips product support engineering scheduled and conducted the update on customer devices to the most current MCB configuration. |  | 50 | 50 | Worldwide Distribution - US (Nationwide) and Internationally to: Austria, France, Germany The Netherlands, Spain, Sweden, and Switzerland. | Terminated | February 13, 2014 | 31 | 50 | Undetermined
Z-0993-2014 | 67076 | No | IPL | Standup wheelchair. | Physical Medicine | Physical Medicine | 510(k) | Superstand stand-up wheelchair |  |  | May 16, 0013 | 13 | The Standing Company | Power Superstand Standing Wheel chair, flex shaft (the flexible wand which supports the joystick control module) may fracture if it is bent repeatedly and /or beyond 90 degrees. A newly designed Flexshaft which corrects this problem is available at no charge. | Component design/selection | The Standing Company sent an Urgent Medical Device Field Correction/Part Recall dated February 6, 2014, to all affected customers. The Urgent Recall notification identified the Flexshaft holding the control joystick on all Power Standing Wheelchairs were being replaced after investigation of complaints found the flexshaft could break because the inner spring of the shaft had failed, although the outer core remained intact. The notice explained how the failure would appear and that the break could happen without warning. The notice explained that the break would not cause the flexshaft to fall off the chair, or hit the ground and the chair could still be mobilized by having the user lean over and pick up the joystick. The notice warned customers not to over bend the flexshaft until replacement was performed. Consignees were asked to call 1-800-STANDING ( 1 800-782-6346) immediately if their wheelchair needed immediate flexshaft replacement. The notice stated that consignees would be contacted to set up a time and date to replace their current flexshaft, at no charge. Consignees were asked to complete and return the response form with customer contact information to Standing Wheelchair Company 5848 Dixie Hwy, Saginaw, MI 48601; Fax (989) 746-9185 or E-mail mail@thestandingcompany.com. Questions and Concerns should be addressed to Customer Service Manager Chris Moran, 1 800 782-6346.For questions regarding this recall call 989-746-9100. |  | 119 | approximately 119 | Nationwide Distribution including AL, CA, DE, IA, IL , IN, KS, KY, MI, MO, NJ, OH, OK, TX, WA, WI,and WV. | Terminated | June 25, 2014 | 730890 | 119 | Undetermined
Z-0991-2014 | 67492 | No | DYB | Catheter introducer. | Cardiovascular | Cardiovascular | 510(k) | MORPH ACCESSPRO Steerable Introducer; |  |  | February 07, 2014 | 2014 | Biocardia, Inc. | Tyvek packaging may be worn, causing loss of package integrity. | Process control | Sales Representative called each affected hospital on February 7, 2014, to notify them of the recall. BioCardia has initiated a recall of Morph AccessPro due to a potential compromise of the sterile barrier. Use of this product may result in patient infection. Sales representatives contacted hospital sites identified in Attachment 3 by telephone. Script provided as follows: "BioCardia has initiated a recall of Morph Access Pro due to a potential compromise of the sterile barrier. Use of this product may result in patient infection. Our records indicate that we sold you. The recall is limited to these units in this lot sold to you. We would like to confirm if you still have the unit(s). If you do, we ask that you quarantine the unit(s) and our customer service department will contact you to coordinate the return of the unit(s)."For further questions please call (650) 226-0133. |  | 11 | 11 units | US Distribution including the states of FL, MA, WY. OH, MI and KS. | Terminated | April 03, 2014 | 55 | 11 | Undetermined
Z-0992-2014 | 67396 | No | FZX | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Synthes Matrix Mandible Short Threaded Drill Guide intended for oral, maxillofacial surgery; trauma and reconstructive surgery, specifically for fractures of the subcondylar region of the mandible and fractures of the condylar basis region of the mandible. |  |  | January 27, 2014 | 2014 | Synthes, Inc. | One lot of Matrix Mandible Short Threaded Drill Guides (part No. 03.503.043) has a gray colored band instead of a green colored band as pictured in the Technique Guide. | Process control | A recall letter, dated January 27, 2014, was sent to end users and sales consultants. |  | 12 | 12 | US distribution to MI, TX, and CA. | Terminated | July 20, 2015 | 539 | 12 | Undetermined
Z-0989-2014 | 67174 | No | JXG | Central nervous system fluid shunt and components. | Neurology | Neurology | 510(k) | Aesculap proSA Adjustment Disc Size Large (L) |  |  | December 23, 2013 | 2013 | Aesculap, Inc. | The accuracy is out of specification for the Aesculap proSA Adjustment Disc Size L. | Employee error | Aesculap sent an Important Correction and Removal Notification dated December 23, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customers..Yom"Aesculap Sales Representative will contact you to schedule a time to remove the FV795T from Set FV792T. Please use FV790T in the absence of the FV795T to ensure the correct adjustment has been made. Please complete the attached Inventory Sheet which is necessary to comply with FDA regulations. When completing the inventory sheet, please fill in the quantity being returned. If you cannot locate the product, please provide an explanation as to why the inventory will not be returned (discarded, etc.).AIC (USA) appreciates your cooperation on this matter and apologizes for the inconvenience this may cause. Thank you for your patience and continued support of this product. Please call (610) 984-9265 or (610) 984-9414 with any questions. |  | 79 | 79 | Worldwide Distribution - US Distribution including the states of AZ, MD, MA, WI, MI, MN, DE, KS, and FL., and the countries of Australia, Germany, Great Britain, Norway and Japan. | Terminated | October 01, 2014 | 282 | 79 | Undetermined
Z-0987-2014 | 67236 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | SAFSITEÂż Injection Sites and IV Sets manufactured with SAFSITEÂż Injection Sites. |  |  | January 08, 2014 | 2014 | B. Braun Medical, Inc. | Complaints were received reporting leakage on certain lots of SAFSITEÂż Injection Sites and IV Sets manufactured with SAFSITEÂż Injection Sites.The SAFSITEÂż Injection Site is intended to provide needle free access on IV and Extension Sets. | Process control | A recall notification letter, dated January 7, 2014, was sent to Users and distributors. |  | 2357944 | 2,357,944 | Nationwide Distribution. | Terminated | November 03, 2014 | 299 | 2357944 | Undetermined
Z-0713-2014 | 67044 | No | CHL | Blood gases (PCO2, PO2) and blood pH test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | pO2 D999 membrane units. Distributed by Radiometer America, Inc., Westlake, OH 44145 |  |  | June 10, 2013 | 2013 | Radiometer America Inc | RADIOMETER became aware that the regarded lot pO2 units 942-042R0971 were produced with inner solution for pCO2 membrane units. This solution is a salt solution similar to the pO2 inner solution, but also contains glycerol which can be reduced at the cathode causing a suddenly rising zero current, causing the membrane unit to fail the following Cal2 calibration. | Mixed-up of materials/components | Radiometer sent a Product Notification letter dated June 18, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to check their inventory and remove any affected product, discard all affected product in their possession, and complete page 2 of the letter and return to their Radiometer representative. Radiometer will replace any quantity of affected product they have at their location. Customers with questions were instructed to contact Radiometer America Technicial Support at 1-800-736-0600, opt 4.For questions regarding this recall call 440-871-8900, ext 209. |  | 1254 | 209 boxes/6 membranes | Worldwide Distribution - US (nationwide) including states of: CO, CT, GA, MA, MO &amp; MS and country of: Canada. | Terminated | June 19, 2014 | 374 | 1254 | Undetermined
Z-0904-2014 | 66912 | No | JJY | Quality control material (assayed and unassayed). | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | ABX PENTRA N Control, Model No. A11A01653 |  |  | November 21, 2013 | 2013 | Horiba Instruments, Inc dba Horiba Medical | Horiba Medical is recalling the ABX PENTRA N Control because it has been updated to include a revised Target Value assignment for Potassium. | Under Investigation by firm | Horiba sent an customer notification letter dated November 32, 2013, to all affected customers. The letter was sent via email to all customers who purchased the ABX PENTRA 400. HORIBA Medical is informing all ABX PENTRA N Control customers running the Potassium-E assay on the ABX PENTRA 400 with I.S.E. Module that the Target Value and Target Range for Potassium has been updated. The letter informs the customers of the actions to be taken. Customers are instructed to complete and return the enclosed response form within 10 days by fax (949) 453-0600 Customers with questions are instructed to contact their local HORIBA Medical representative.For further questions please call (949) 453-0500 X 208 |  | 698 | 698 units total (80 units in US) | Worldwide Distribution - USA (nationwide) | Terminated | October 28, 2014 | 341 | 698 | Undetermined
Z-1451-2014 | 67876 | walker, mechanical | ITJ | Mechanical walker. | Physical Medicine | Physical Medicine | 510(K) Exempt | Sit/stand Rolling Walker, SKU/Item Number 66838. Intended for use by those with difficulty walking and weighing 225 lbs. or less. |  |  | March 27, 2014 | 2014 | Central Purchasing LLC DBA Harbor Freight Tools | Harbor Freight Tools intiaited a recall of Sit/Stand Rolling Walker (SKU/Item Number 66838, sold June 2009 to March 2014) because, over time, there is the possibility that a walker wheel may come off this product following repeated use without tightening the wheel bolt to the frame, presenting a potential hazard for the user. | Under Investigation by firm | Recall communication to web site customers who purchased the product, notifying them of recall and availability of full refund for product.Store posting notifying retail customers of the recall and availability of fullrefund for product (with a similar posting provided on the HFT web site). Consumers are informed of the recall, product info. with photo, item number, reason for recall, hazard, instructed to stop using these walkers immediately, and contact information provided telephone number, and e-mail address. Consumer letters also provided with a Merchandise Return Label.On 04/02/14, an updated consumer letter issued. |  | 43621 | 43,621 | Nationwide Distribution. | Terminated | August 05, 2015 | 496 | 43621 | Undetermined
Z-0532-2014 | 66064 | No | IZL | Mobile x-ray system. | Radiology | Radiology | 510(k) | Spectre Wireless Footswitch/Handswitch system Model 98/99, Model FDS, Model FDB. |  |  | August 22, 2013 | 2013 | Trans American Medical, Inc. | TransAmerican Medical Imaging is recalling certain lots of TMI Spectre Wireless Footswitch/Hand switch for GE OEC 9800/9900 and Philips FDS and FDB Cath Labs due to potential safety issue related to the use of this product producing uncommanded X-ray. | Component design/selection | TransAmerican Medical sent an Urgent Medical Device Correction letter dated August 26, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to ensure that all potential users in their facility are made aware of the safety notification and the recommended actions. If the customers system does produce an uncommanded X-ray, turn the system off immediately. Remove the Spectre Wireless Footswitch and receiver and reinstall the standard OEM footswitch. Follow proper shutdown/system reboot instructions as indicated in the system operator's manual. Customers will be contacted to make arrangements for replacement of their TMi Spectre Wireless Footswitch at no charge. Customers will be provided a new TMI Spectre Wireless Footswitch and instructions on how to return the affected footswitch. Customers with questions were instructed to call 801-796-7335, ext 203. |  | 68 | 68 | US Distribution: including states of: WA and UT. | Terminated | December 29, 2015 | 859 | 68 | Undetermined
Z-1266-2014 | 67576 | No | IMP | Wheelchair component. | Physical Medicine | Physical Medicine | 510(K) Exempt | Otto Bock Aerial and Aerial Active Wheelchair Cushions. Material numbers: SK757 476C00=SK701, 476C00=SK702, 476C00=SK703, 476C00=SK704, 476C00=SK705, 476C00=SK706, 476C00=SK707, The following material numbers were not distributed within the US: 476C00=SK751, 476C00=SK752, 476C00=SK753, 476C00=SK754, 476C00=SK755, 476C00=SK756, 476C00= SK757. |  |  | January 31, 2014 | 2014 | Bock,Otto,Orthopedic Ind,Inc | Otto Bock is recalling all Aerial and Aerial Active Wheelchair Cushions from specific material numbers. During internal testing, it was discovered that there is a small risk of flammability to the cushion after it has been washed at high temperatures. The cushions should either be returned to Otto Bock or destroyed. | Unknown/Undetermined by firm | An Urgent Voluntary Device Recall letter was sent to customers beginning 1/31/2014. The letter identified the affected material numbers, described the reason for recall, and asked that product be returned or destroyed. A response Questionnaire was asked to be completed and returned. |  | 464 | 464 units (US 113, OUS 351) | Worldwide Distribution-USA (nationwide) and the countries of Australia, Austria, Bosnia and Herzegovina, Brazil, China, Columbia, Croatia, Denmark, France, Germany, Greece, Hong Kong, Hungary, Iceland, Japan, Mexico, Netherlands, Poland, Romania, Russia, South Africa, Spain, Sweden, Switzerland, Thailand, and United Kingdom. | Open, Classified |  |  | 464 | Undetermined
Z-1302-2014 | 67712 | No | DYB | Catheter introducer. | Cardiovascular | Cardiovascular | 510(k) | Morph AccessPro Steerable Introducer; |  |  | March 03, 2014 | 2014 | Biocardia, Inc. | Firm discovered the potential for particulates in the inner lumen of Morph Access Pro Steerable Introducer, which could result in am embolic event. | Other | The firm, BioCardia, Inc., sent an "URGENT-VOLUNTARY MEDICAL DEVICE RECALL" letter dated March 3, 2014 its customers, hospital's Director of Risk Management, via USPS mail and priority mail to Cardiac Catheterization Laboratory. The letter described the product, problem and actions to be taken. The customers were instructed to immediately examine your inventory; quarantine the product; return the product to BioCardia and contact Customer Service at 650-226-0148 to coordinate product return; in the event you no longer have this device as a result of further distribution, identify the purchaser and notify them at once of this recall; and complete and return the enclosed MEDICAL DEVICE RECALL RETURN RESPONSE Acknowledgement and Receipt Form via FAX to 650-631-3731 or mail to: BioCardia, Inc., 125 Shoreway Road, Suite B, San Carlos, CA 94070 as soon as possible.If you have any questions, call the Vice President, Quality Assurance at 650-226-0143. |  | 2541 | ~2541 units - both catalog numbers. | Worldwide Distribution: US (nationwide) and country of: Israel. | Terminated | June 19, 2014 | 108 | 2541 | Undetermined
Z-0626-2014 | 65409 | No | LCP | Glycosylated hemoglobin assay. | Hematology | Hematology | 510(k) | HbA1c Flex(R) reagent cartridge (K3105A, Siemens Material Number 10470481) for the Dimension(R) Clinical Systems. |  |  | April 22, 2013 | 2013 | Siemens Healthcare Diagnostics, Inc. | Positive bias on the lots listed compared to alternate methods and a higher frequency of "Above Assay Range" flags with the lots listed. | Process control | Siemens sent an Urgent Medical Device Recall letter in April 2013 to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to immediately discard any remaining inventory of the affected product. Siemens will replace any unused inventory at no charge. Customers were asked to complete the attached form and fax it to 302-631-8467 to indicate they have received the information. Customers were asked to retain the letter with their laboratory records and forward the letter to those who may have received the affected product. Customers with questions were instructed to contact their Siemens Technical Solutions Center or their local Siemens technical support representative.For questions regarding this recall call 1-800-441-9250. |  | 11006 | 11006 Total | Worldwide Distribution - USA (nationwide) and internationally to Canada and Mexico. | Terminated | February 12, 2014 | 296 | 11006 | Undetermined
Z-1305-2014 | 67646 | No | NTE | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | FIBERNET Embolic Protection System: |  |  | March 03, 2014 | 2014 | Medtronic Vascular, Inc. | Medtronic is recalling a specific subset of lot numbers of the FIBERNET Embolic Protection System (EPS) where the yellow polyimide tube on the FIBERNET EPS was not removed from the FIBERNET EPS during the manufacturing process and subsequently may remain on the FIBERNET EPS. | Manufacturing material removal | Medtronic sent an Urgent Device Recall letter dated March 3, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Our records show that your facility has received one or more of these FIBERNET EPS, as shown in the attachment to this notification. Consequently, Medtronic requests you immediately take the following actions:1. Immediately remove and quarantine all potentially affected product that remains in your inventory.2. Return these units to Medtronic. Contact Customer Service at 1-888- 283-7868 and reference RADAR # 174028 to initiate a return and credit of product. A Medtronic representative can assist in facilitating the return and replacement of this product as necessary.3. Complete the attached Customer Confirmation Certificate and email to RS.CFQFCA@Medtronic.com, or fax to Medtronic at 651-367-0612 to the attention of Customer Focused Quality.Medtronic has taken the necessary steps to prevent any future shipment of the potentially affected product. Regulatory agencies will be notified about this recall as applicable.Please share this notification with others in your organization as appropriate. lf any FIBERNET EPS within the scope of this recall have been forwarded to another facility, please notify that facility accordingly and facilitate the retrieval of the affected product.We apologize for the inconvenience this may cause you; please be assured that patient safety and product quality remain our primary concern. Should you have any questions, please contact your Medtronic representative or Medtronic LifeLine Technical Services at 1 (877) 526-7890 or (+1 763) 526-7890. |  | 50 | 50 Total all sizes: 18 in US, 32 Internationally | Worldwide Distribution - US including the states of IL, PA, MA, VA. RI and LA., and the countries of Germany, Italy, Poland, Romania, Saudi Arabia, Kosovo, Slovakia and Spain. | Terminated | June 03, 2014 | 92 | 50 | Undetermined
Z-1320-2014 | 67844 | No | FMF | Piston syringe. | General Hospital | General Hospital | 510(k) | Merit Custom Syringe Kit, Convenience Kit, I.R. Embolization Pack, K02-01010A, Sterile EO. |  |  | March 13, 2014 | 2014 | Merit Medical Systems, Inc. | The products are labeled as sterile but were not sterilized. | Release of Material/Component prior to receiving test results | Merit Medical sent an Urgent Product Recall Notice dated March 14, 2014 to all customers and sales representatives via email. The letter identified the affected product, problem and actions to be taken. Customers were provided a Product Retrieval Form, product identification information, instruction to immediately quarantine any devices and discontinue use, ensure all personnel to whom devices were distributed are made aware of this field action, and instructions to contact their Merit representative to arrange product return and replacement. For questions call 1-801-316-4822. |  | 15 | 15 | Worldwide Distribution - USA Nationwide in the state of WI and countries of Thailand and Hong Kong. | Terminated | July 07, 2014 | 116 | 95 | Undetermined
Z-0652-2014 | 67087 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Arthrocare SPORTS MEDICINE, Knot Pusher, REF 25-3011, LOT #1049402, CON 1, 2013-10, NON-STERILE, Rx only |  |  | December 09, 2013 | 2013 | ArthroCare Medical Corporation | A sharp edge on the cannulated tip of the device may damage or cut a suture. | Nonconforming Material/Component | All consignees were provided with written notice via Urgent Recall Notification dated December 9, 2013 with delivery confirmation. In addition, the firm plans on notifying the EU, and the governments of France, Germany, Switzerland, and Spain. The firm will also notify the Swedish government because the firm's EU representative is located in Sweden. |  | 76 | 76 | Worldwide distribution: US (nationwide) including states of: CA, FL, GA, IL, IN, KY, MN, MO, NY, OH, TX, and WI. Internationally to: Canada, France, Germany, Spain, and Switzerland. | Terminated | April 08, 2015 | 485 | 76 | Undetermined
Z-0628-2014 | 67039 | No | HRY | Knee joint femorotibial metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | REF 42-411571 QTY 1 CATALOG CAUTION. SEE INSTRUCTIONS FOR USE |  |  | November 26, 2013 | 2013 | Materialise USA LLC | Shipped femur guide contained a non-conformance which could possibly impact fit of the drill insert. | Nonconforming Material/Component | Materialise USA LLC notified the Field Representative on November 26, 2013 of the recall by phone followed by E-mail notification. The notification identified the device and reason for the recall. Notification directions included instructions to find and quarantine the device and return it Materialise. A replacement device was shipped November 27, 2013 receiving the defective device was called, followed by E-mail notification the Recall, which identified the device, problem, and instructed the distributor to find and quarantine product. For questions regarding this recall call 734-662-5057. |  | 1 | 1 | US Distribution: MO only. | Terminated | January 14, 2014 | 49 | 1 | Undetermined
Z-2250-2012 | 62856 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Newport HT70 and HT70 Plus Ventilators, Model Number: HT70 and HT70 Plus. |  |  | August 03, 2012 | 2012 | Newport Medical Instruments Inc | May emit a continuous high priority alarm and the ventilator may stop ventilating, due to a component failure on the control board. | Nonconforming Material/Component | Newport Medical sent an Urgent Field Correction Notice letter dated August 3, 2012 to all affected customers, via DHL and FedEx courier services. The letter identified the affected product, reason for recall and actions to be taken. Customers were instructed to immediately stop using the affected ventilator and retrieve the replacement kit provided free of charge. In addition, the letter states that if the product was further distributed, to contact their customers and notify them of the recall. Customers were asked to complete and return the enclosed Replacement form as replacements are completed. For questions contact Newport Medical Technical Support Department at 800-451-3111, extension 282, for U.S. customers; 714-427-5811, extension 500, for international customers.". |  | 72 | 72 | Worldwide distribution: USA (nationwide) and countries of: Brazil, Chile, Egypt, France, Hong Kong, India, Japan, Norway, Philippines and South Africa. | Terminated | April 14, 2014 | 619 | 72 | Undetermined
Z-1306-2014 | 67513 | No | CHL | Blood gases (PCO2, PO2) and blood pH test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | i-STAT G3+ Cartridge |  |  | January 01, 2014 | 2014 | Abbott Point Of Care Inc. | Abbott Point of Care has determined that some individual patient results generated with thei-STAT G3+ cartridge lot N13183 have the potential to exhibit incorrectly elevated PCO2 anddepressed pH results. | Process control | APOC sent an Urgent Recall Notice dated January 2014 via Fed Ex to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to discontinue use of the affected product and provide a copy of this letter to any institution that was forwarded the affected product. All unused cartridges will be credited and are to be returned as instructed on the enclosed Business Reply Card. To ensure continuity of supply at their facility, customers were instructed to follow their standard process to place an order or call Customer Sedrvice at 1-800-323-9100, Option #2, then Option #9). Customers with questions were instructed to contact Abbott Point of Care Technical Support at 1-800-366-8020, Option 1 of their Abbott Point of Care representative.For questions regarding this recall call 800-323-9100. |  | 48 | 48, 450 US, 1400 OUS (49,850 catridges) | Worldwide Distribution - USA (nationwide) and Internationally to Canada. | Terminated | October 20, 2014 | 292 | 48 | Undetermined
Z-1439-2014 | 67857 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | Philips HeartStart XL+ Defibrillator/Monitor |  |  | March 31, 2014 | 2014 | Philips Medical Systems, Inc. | A component on the HeartStart XL+ Therapy Board could malfunction potentially affecting the ability to deliver therapy | Reprocessing Controls | Philips Healthcare issued an Urgent-Medical Device Correction letter dated March 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. This Field Safety Notice is intended to inform you about:" what the problem is and under what circumstances it can occur" the actions that should be taken by the customer/user in order to prevent risks for patients or users" the actions planned by Philips to correct the problem.A component on the HeartStart XL+ Therapy Board could malfunction potentially affecting the ability to deliver therapy.Please see the attached Field Safety Notice that provides information on how to identify affected devices and instructions on actions to be taken. Please follow the ACTION TO BE TAKEN BY CUSTOMER / USER section of the notice. To correct this issue Philips will provide a replacement therapy board to customers with affected units free of charge.This voluntary correction has been reported to the appropriate regulatory agencies.If you need any further information or support concerning this issue, please contact your local Philips representative at Philips representative contact details to be completed by the KM / country.Philips apologizes for any inconveniences caused by this problem.Further questions please call (978) 687-1501. |  | 19 | 19 units | Worldwide Distribution - US Distribution including the states of IN and WA., and the countries of : CHINA FRANCE GERMANY ITALY MALAYSIA PANAMA SAUDI ARABIA TURKEY and UNITED KINGDOM. | Terminated | October 21, 2015 | 569 | 19 | Undetermined
Z-1455-2014 | 67797 | Yes | FJK | Hemodialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Gambro Cartridge Blood Set, blood transport system for hemodialysis, Model Number(s): 101025 (003410510) - Cartridge set, STND PRM LN and 103401 (003414500) - Cartridge set, PRM_LN_INJ_PT. The Gambro Cartridge Sets are single use sterile tubing sets intended to provide extracorporeal blood transport circuit for hemodialysis treatments for Gambro Phoenix and COBE Centrisystem 3 (and 3+) Dialysis Delivery Systems. |  |  | December 09, 2013 | 2013 | Gambro Renal Products, Incorporated | Occluded heparin tubing events preventing anticoagulation dosing on the Gambro Cartridge Blood Sets. | Process control | On 12/9/13 and 3/12/14, a field safety notice sent via UPS overnight delivery informing customers of issue and if any customer wants to return product, a hold at distribution centers and manufacturer quarantine areas. Upon reconciliation of returned product a destruction disposition with certification. |  | 4597875 | 306,525 sets (20435 boxes of 15 sets) | Worldwide Distribution-USA (nationwide) and the countries of Canada, Mexico, and Colombia. | Terminated | December 22, 2015 | 743 | 4597875 | Undetermined
Z-1298-2014 | 67502 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Draeger PS500 Power Supply Unit of the Infinity ACS Workstation Critical Care (Evita Infinity V500), Evita V500 and /or Babylog VN500 Ventilators, used for ventilation during short term transport of patients within a hospital. |  |  | February 07, 2014 | 2014 | Draeger Medical Systems, Inc. | The battery capacity of the optional PS500 Power Supply Unit of the Infinity ACS Workstation Critical Care (Evita Infinity V500) did not last as long as expected. | Component design/selection | Drager sent an Urgent Medical Device Recall letter dated February 2014. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to use an affected device for patient transport only if absolutely necessary; not to rely on the battery charge status indicator, always supervise the patient and the ventilator during transport; ensure that they have a manual breathing bag available for manual ventilation as recommended in the instructions for use. If the power fail alarm occurs during transport, immediately provide manual ventilation and connect ventilator to a wall power source to resume ventilation. DragerService representative will schedule the implementation of a permanent solution as quickly as possible. The letter should be distributed to users of the Evita V500 and /or Babylog VN500 Ventilators to make them aware of this device behavior and risk mitigation measures. Customers with questions were instructed to call 1-800-543-5047. For questions regarding the operation and/or servicing of their Draer ventilator contact Technical Support at 1-800-543-5047. For questions regarding this recall call 215-721-5400. |  | 1544 | 1544 | Nationwide Distribution | Terminated | August 17, 2015 | 556 | 1544 | Undetermined
Z-0650-2014 | 66398 | No | CGA | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Nova StatStrip/StatSensor Hospital Blood Glucose Meter; Nova StatStrip Lactate Hospital Meter System; Nova StatSensor Creatinine Hospital Meter, |  |  | October 10, 2013 | 2013 | Nova Biomedical Corporation | Revised Battery Instructions for Removal and Expiration Date due to battery swelling and cell leakage | Component design/selection | On November 18, 2013, an URGENT Field Correction Letter was sent to all StatStrip Glucose Hospital meter, StatStrip Lactate Hospital Meter and StatSensor Creatinine Hospital Meter customers via email to the Point of Care Coordinator to alert users that lithium batteries that are utilized beyond their intended useful life may swell and/or crack posing a potential safety hazard. In some cases, the swelling may potentially cause the battery door on the meter to dislodge and render the meter inoperable. |  | 200000 | est. 200,000 | Nationwide Distribution. | Terminated | March 24, 2014 | 165 | 200000 | Undetermined
Z-0624-2014 | 66273 | No | HXB | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | PedFuse Bone Probe; Part: SI81008 spinal pedicle fixation. |  |  | September 11, 2013 | 2013 | SpineFrontier, Inc. | Bone Probes used to create a pathway in pedicle bone were manufactured and marked to incorrect specifications. | Process change control | On 9/13/2013, consignees were notified of the recall via email. They were asked to return the recalled product and complete and return the recall acknowledgement form via email or fax. |  | 13 | 13 devices | US Distribution: including states of: TX and VA. | Terminated | October 08, 2014 | 392 | 13 | Undetermined
Z-1329-2014 | 67736 | No | FTM | Surgical mesh. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Proceed Surgical Mesh, Product Codes PCDB1, PCDT1, and PCDJ1. |  |  | February 19, 2014 | 2014 | Ethicon, Inc. | The possibility exists that an incomplete seal on the packaging compromised the sterility of the device and introduced the potential for delamination. | Under Investigation by firm | Ethicon sent an Urgent Medical Device Recall letter dated February 19, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Customers are asked to examine their inventory and determine if they have any of the affected product. If any product has been distributed, those hospitals should be notified. The business reply form should be completed and faxed or emailed to Stericycle at 1-888-641-9732 or ethicon5798@stericycle.com. For clinical or product support contact a customer representative or call the Customer Support Center at 1-877-ETHICON to speak with a registered nurse or contact Stericycle at 1-888-345-0524 and reference event #5798. Stericycle is handling returns. |  | 223 | 223 units | US Distribution including the states of FL, NJ, MO, IL, PA, NC. TX, TN, WA, ID, UT, WV, MS. CA, AL, MA. GA, LA and OK. | Open, Classified |  |  | 223 | Undetermined
Z-0648-2014 | 67161 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | SOLAR 8000 AND TRANSPORT PRO WITH PATIENT DATA MODULE (Marketed As GE Healthcare CARESCAPE Patient Data Module) |  |  | December 16, 2013 | 2013 | GE Healthcare, LLC | GE Healthcare has recently become aware of a potential safety issue which affects all Patient Date Modules produced from April 2007 through June 2010, encompassing Fiscal Week 13, 2007 through 26, 2010 due to fluid ingress into the top cover assembly which can result in heat buildup. | Nonconforming Material/Component | Consignees were sent on 12/16/2013 a GE Healthcare "Urgent Medical Device Correction" letter dated December 5, 2013. The letter was addressed to Healthcare Administrator / Risk Manger, Chief of Nursing and Director of Biomedical Engineering. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction and Contact Information (Contact Technical Support at 1-800-558-7044 or your local Service Representative). |  | 17444 | 17,444 (11,418 units US, 6,026 units OUS). | Worldwide Distribution: US (nationwide) including states of: AL, AZ, AR, CA, CO, CT, DC, FL, GA, ID, IL, IN, IA, KS, KY, LA, ME, MD, MA, MI, MN, MS, O, NE, NV,NH, NJ, NM, NY, NC, OH, OK, OR,PA PR, RI,SC, TN, TX, UT, VT, VA, WA, WI, and WY. Internationally to: ARGENTINA, AUSTRALIA, AUSTRIA, BANGLADESH, BELGIUM, BRAZIL, CANADA, CHILE, CHINA, COLOMBIA, CZECH REPUBLIC, DENMARK, EGYPT, ESTONIA, FINLAND, FRANCE, GERMANY, GREECE, GUATEMALA, INDIA, INDONESIA, IRELAND, ISRAEL, ITALY, JAMAICA, JAPAN, JORDAN, REPUBLIC OF KOREA, KUWAIT, LEBANON,LITHUANIA, MALAYSIA, MAURITIUS, MOROCCO, NETHERLANDS, NEW ZEALAND, NORWAY, POLAND, QATAR , SAUDI ARABIA, SINGAPORE, SOUTH AFRICA, SPAIN, SWEDEN, SWITZERLAND, TAIWAN, THAILAND, TRINIDAD AND TOBAGO, TUNISIA, TURKEY, UNITED ARAB EMIRATES, and UNITED KINGDOM. | Terminated | April 19, 2016 | 855 | 17444 | Undetermined
Z-1330-2014 | 67849 | No | DYW | Orthodontic plastic bracket. | Dental | Dental | 510(k) | Spirit MB Bracket, UR Central +14T +5A.018, Part Number 494-0110, Lot Number 081351013. |  |  | March 25, 2014 | 2014 | Ormco/Sybronendo | Ormco Corporation initiated this recall of Spirit MB Brackets, Part Number 494-0110, Lot Number 081351013, because the ID dot, which designates the bracket as distal/gingival, is misplaced. The misidentification may lead an orthodontist to place the bracket on the wrong tooth, leading to movement 10 degrees in the wrong direction. | Use error | Ormco Corporation initiated this recall by sending recall letters to customers on March 25, 2014 via 1st class mail. The letter identified the affected product, problem and actions to be taken. The consignees are instructed to complete the "Return Form" and return any affected product in their inventory. The letter was written in Japanese language and accompanied by a "Return Form". |  | 150 | 150 units | Internationally Distributed in the country of Japan. | Terminated | April 18, 2014 | 24 | 150 | Undetermined
Z-0632-2014 | 67131 | No | GFF | Surgical instrument motors and accessories/attachments. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Anspach B-Blue-S 11.25 cm attachment and Anspach B-Green Standard Craniotome. |  |  | June 11, 2012 | 2012 | The Anspach Effort, Inc. | Anspach was contacted by a hospital to inform Anspach that a Blackmax motor had been used in a case where the patient was diagnosed with Creutzfeldt-Jacob Disease (CJD). | Unknown/Undetermined by firm | none |  | 2 | 1 B-Blue-S and 1 B-Green | US Distribution: Maryland only. | Terminated | April 18, 2016 | 1407 | 6 | Undetermined
Z-0654-2014 | 66323 | No | HSZ | Surgical instrument motors and accessories/attachments. | General &amp; Plastic Surgery | Orthopedic | 510(K) Exempt | ***REF BLACKMAX-N***Swivel/Angle Motor Assembly***Rx Only***The Anspach Effort, Inc. |  |  | April 25, 2013 | 2013 | The Anspach Effort, Inc. | The firm Anspach Effort of Palm Beach Gardens, FL initiated a voluntary removal of the Pneumatic Drill System, including the motor hand-piece, Auto Lube foot control, without a pressure relief valve (PRV) due to the hand piece's outer hose can rupture without warning if the air flow is blocked during use.. | Component design/selection | The firm, Anspach sent an "URGENT" Medical Devics Removal" letter dated April 24, 2013, to its customers. Customer Immediate Actions: 1. Screen your current inventory and return all Anspach Pneumatic Drill System without a pressure Relief valve (PRV) from your inventory. 2. If you have any product listed on the attached form to arrange for immediate return and replacements. 3. complete the attached Customer Reply form indicating the product being returned. 4. Upon receipt of a replacement product, equipment being returned must be received by Anspach within 10 business days of the replacement request. Should you have any queries, please contact Anspach Product Support at (800) 327-6887. |  | 2785 | 2,785 | Worldwide Distribution: US (nationwide) including Puerto Rico and states of: AL, AR, AZ, CA, CO, CT, DC, DE, GA, FL, IA, ID, IL, IN, KY, LA, MA, MD, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OR, PA, RI, SC, SD, TN, TX, UT, VA, VE, VT, WA, WI, WY, and WV. Internationally to: AT, AU, BR, BS, CA, CH, CL, CO, CZ, CR, DE, DK, DO, EG, ES, FI, FR, GB, GR, HK, HU, ID, IL, IN, IS, JP, KH, KR, LB, LT, MX, MY, NL, NO, NZ, PA, PE, PH, PL, PT, SA, SG, SY, TH, TR, TW, and ZA. | Terminated | July 20, 2015 | 816 | 4997 | Undetermined
Z-1450-2014 | 67804 | No | FSA | Non-AC-powered patient lift. | General Hospital | General Hospital | 510(K) Exempt | 400003 Medcare Lift with Scale |  |  | May 10, 2013 | 2013 | ErgoSafe Product LLC dba Prism Medical Services | MedCare Products has determined that a number of their Floor Lifts manufactured from 9/1/2010 to 11/7/2012 do not meet MedCare's high standards. The original design includes a pin that over time and years of use may stop performing as designed and may break. | Nonconforming Material/Component | Medcare sent a Notification letter dated May 10th, 2013 to affected customers. The letter described the affected product, problem, actions to be taken and provided a replacement assembly. It instructed consignees to replace the old assembly with the new one included in the package. For questions they can contact Todd Perrine in their Technical Support Department at 952-300-2101. |  | 514 | 514 | USA Nationwide Distribution in the states of: CA, CO, IL, IN, IA, KS, KY, LA, ME, MD, MI, MN, MS, NO, NE,NH,NY, ND,OH, OK, OR, PA , SD, TX, WI. | Terminated | November 20, 2014 | 559 | 514 | Undetermined
Z-1336-2014 | 67597 | No | HGX | Powered breast pump. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | Playtex Nurser Deluxe Double Electric Breast Pump, Model X06578AO, powered by a 12V DC power supply, Model SY-12120, which is included with the package. |  |  | March 18, 2014 | 2014 | Energizer Personal Care | Some of the power adapters outer casings may become loose and separate, resulting in a potential for electric shock. | Component change control | Energizer sent and Voluntary Retail Recall Notice letter dated March 18, 2014 to retail customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to stop sale and return the affected product. The letter provides instruction to customers to return the affected product. For questions or concerns, contact Energizer Personal Care at 1-888-207-1492 or visit www.playtexproducts.com.On 9/9/2014 the firm expanded the recall to one additional production lot to include: Suspect date codes for Breast Pumps: P12315-XXXX to P13205-XXXX. Suspect adaptor date codes: 1238 to 1324Revised Press issued 9/9/2014. |  | 35803 | 35,803 | USA Nationwide Distribution in the states of: AL, AZ, CA, CO, CT, DE, FL, GA, IA, IL, IN, KS, KY, MD, MI, MN, MO, NC, NJ, NV, NY, OH, ON, OR, PA, SC, TN, TX, VA, VT, WA, WICanada - one consignee, Energizer (recalling firm's) company store | Completed |  |  | 35803 | Undetermined
Z-1299-2014 | 67624 | No | LCP | Glycosylated hemoglobin assay. | Hematology | Hematology | 510(k) | D-10" Hemoglobin Testing System, |  |  | February 27, 2014 | 2014 | Bio-Rad Laboratories, Inc. | On a rare occasion the D-10 software might generate an extra result by duplicating the result form the previous sample and the Sample ID/Injection number from the next sample. | Nonconforming Material/Component | Bio-Rad Laboratories, Inc. sent an Urgent Medical Device correction letter dated February 27, 2014, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. A Customer medical Device Correction Response Form must be collected from each customer to ensure that they have received this communication. Forms should be faxed to the Bio-Rad CSD Regulatory Affairs Department at 510-741-3954 along with the completed Subsidiary Medical Device Correction Response Form.For questions regarding this recall call 570-741-6114. |  | 483 | 483 instruments in US. OUS to be determined. | Worldwide Distribution - USA (nationwide) and Internationally to Canada, Japan, China, India, SriLanka, Argentina, Paraguay, Uruguay, Chile, Peru ,Bolivia ,Mexico ,Panama, ,Costa Rica ,Ecuador, Venezuela, Colombia, Brazil, Jamaica, Austria, Switzerland, France, Germany, Greece, Portugal, Spain, Italy, UK, Belgium, Sweden Norway, Finland ,Israel/Palestine, Tunisia, Algeria, Egypt, ,Morocco, Jordan, Yemen, Iraq, Kuwait, Qatar, Oman, Lybia, Lebanon ,UAE, Saudi Arabia, Pakistan, Afghanistan, Turkey, South Africa, Chad, Kenya, Benin, Burkina Faso, Cameroon, Gabon, Senegal, Ghana, Mali, Ivory Coast, Nigeria, Djibouti, Czech Republic, Slovakia, Slovenia, Serbia, Hungary, Romania, Poland, Russia, Ukraine, Kazakhstan, Bulgaria, Belarus, Georgia, Mauritius, Maldives, Vietnam, Philippines, Korea, Thailand, Hong Kong , Indonesia, Taiwan, Malaysia, Singapore, New Zealand, and Australia. | Terminated | June 04, 2014 | 97 | 483 | Undetermined
Z-1440-2014 | 67677 | Yes | BTL | Powered emergency ventilator. | Anesthesiology | Anesthesiology | 510(k) | Draeger Oxylog 3000 plus ventilators, emergency and transport ventilator. Part number: 5704833 |  |  | March 05, 2014 | 2014 | Draeger Medical Systems, Inc. | The potentiometers (knobs) for setting the ventilation parameters may have contact interrupted and the device discontinues ventilation, gives an audible alarm and displays the error message Poti unplugged. | Process control | Drager sent an Urgent Medical Device Recall letter dated February 2014 to affected customers. The letter identified the affected device, problem and actions to be taken. The notification directed customers to perform manual ventilation in the event of a failure. Drager plans to repair the affected devices free of charge. If you have any questions regarding the Recall letter please call 1-800-543-5047 (press 1 at the prompt, then 2, then 32349). For questions regarding the operation and/or servicing of your DrÂżger ventilator please contact DrÂżgerService Technical Support at 1-800-543-5047 (press 4 at the prompt). |  | 3 | 3 | USA Nationwide Distribution in the states of WI and NY. | Terminated | October 31, 2014 | 240 | 3 | Undetermined
Z-1276-2014 | 67710 | No | GCJ | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | WeckÂż Vista", Universal Laparoscopic Port, Size 5/10mmx100mm, |  |  | March 11, 2014 | 2014 | Teleflex Medical | Complaints were received of leakage of insufflation gas through the device, which can cause difficulty or delay in the completion of laparoscopic surgical procedures. | Nonconforming Material/Component | Teleflex sent an Urgent Medical Device Recall Notification letter dated March 11, 2014 to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to return all affected product to Teleflex Medical per the instructions provided in the letter. Customers were asked to complete the enclosed Recall Acknowledgment Form and fax it to 1-866-804-9881, Attn: Customer Service. For questions contact your local sales representative or Customer Service at 1-866-246-6990. |  | 55521 | 55,521 ea. in total | Worldwide Distribution - USA Nationwide including Puerto Rica and the countries of: Australia, Brazil, China, Canada, Malaysia, Korea, Ireland Singapore and Vietnam. | Terminated | August 15, 2017 | 1253 | 55521 | Undetermined
Z-1447-2014 | 67821 | orthosis, limb brace | IQI | Limb orthosis. | Physical Medicine | Physical Medicine | 510(K) Exempt | K-Lite Knee 2 Stays with Visco, Model No. 122015. |  |  | March 14, 2014 | 2014 | Breg Inc | Breg is recalling the Economy Elastic Knee Sleeve Open and K-Lite Knee 2 Stays with Visco because they contain latex but were not labeled with the required caution label. | Process change control | Breg representatives informed customers of the recall by telephone on March 14, 2014 and then a customer notification letter was sent to customers via email onMarch 17, 2014. The letter informs the customers of the problems identified and the actions to be taken. Customers are instructed to complete the return response form. Customers with questions are instructed to contact Breg Customer Care at (800) 321-0607. |  | 2 | 2 units | USA Nationwide Worldwide in the states of: TX, KY, IN, and FL. | Terminated | April 14, 2014 | 31 | 22 | Undetermined
Z-1293-2014 | 67733 | No | GAT | Nonabsorbable poly(ethylene terephthalate) surgical suture. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Sutures, Non-absorbable, Synthetic, Polypropylene, Sterile, Rx only, |  |  | March 11, 2014 | 2014 | Teleflex Medical | The products are being recalled because they did not meet minimum needle attachment strength requirements. | Employee error | Teleflex sent an Urgent Medical Device Recall Notification letter dated March 11, 2014 to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to return all affected product to Teleflex Medical per the instructions provided in the letter. Customers were asked to complete the enclosed Recall Acknowledgment Form and fax it to 1-866-804-9881, Attn: Customer Service. For questions contact your local sales representative or Customer Service at 1-866-246-6990 |  | 32271 | Total 32,271 ea. | Worldwide Distribution - U.S. Nationwide and the countries of Canada, Germany, and Ireland | Terminated | August 15, 2017 | 1253 | 32271 | Undetermined
Z-1274-2014 | 67263 | No | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Drill Template and the Saw Guide for the Ulna Osteotomy System. |  |  | January 07, 2014 | 2014 | Synthes, Inc. | The Drill Template and the Saw Guide for the Ulna Osteotomy System may exhibit jamming, bending, or breaking of the screw connecting the Saw Guide to the Drill Template or jamming, bending, or breaking of the K-wire during fixation of the Drill Template. | Component design/selection | A recall notification letter, dated January 7, 2013, was sent to sales consultants and Users. |  | 343 | 343 | Nationwide Distribution-including the states of AR, AK, VA, IN, PA, OH, CT, TX, MI, WI, NC, CO, IL, and MA. | Terminated | September 10, 2015 | 611 | 343 | Undetermined
Z-1448-2014 | 67675 | No | HSB | Intramedullary fixation rod. | Orthopedic | Orthopedic | 510(k) | Pega Medical Inc., Fassier-Duval IM Telescopic System, male component (solid shaft), Part number: M040-SS-110; Used in Implant Catalogue number: FD-040(LS)-SS. orthopedic intramedullary rod. |  |  | February 07, 2014 | 2014 | Pega Medical Inc. | The Male Components in this lot are made of material with lower strength than manufacturing specifications. | Nonconforming Material/Component | Pega Medical sent an Urgent Field Safety Notice-Recall letter dated February 7, 2014, to their customers via email. The letter identified the affected product, the problem, and the action to be taken by the customer. Customers were instructed to identify the affected devicee, verify if it was used in a patient and place the affected device in quarantine immediately if not used. If not used, return the affected device to Pega Medical Inc. to the address provided. Pega's customer notification stated "If used, it is recommended to continue the normal patient follow up.Customers with questions should call (450) 688-5144 ext. 242 or email egarcia@pegamedical.com |  | 15 | Total of 5 units (2 units in the US and 3 units in Germany) | Worldwide Distribution - USA Nationwide in the states of Missouri, Texas, and the country of Germany. | Terminated | April 23, 2014 | 75 | 15 | Undetermined
Z-0629-2014 | 66622 | Yes | KDI | High permeability hemodialysis system. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Fresenius 2008 Series Hemodialysis Machines: |  |  | November 12, 2013 | 2013 | Fresenius Medical Care Holdings, Inc. | 2008 Series Hemodialysis Machines complaints of saline bags inappropriately filling during recirculation and setup | Under Investigation by firm | Fresenius Medical issued a Product Notification letter dated November 12, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers who encounter a saline bag that is inappropriately filling, should discontinue the setup and remove the machine from service until the drain line and drain-associated hydraulic components can be checed by a qualified service technician. Customers were also instructed to ensure that the hemodialysis machine is installed and operated in compliance with the Machine Specification section of their operator's manual, including the specification located in the Water/Drain Section. Customers with questions were instructed to call 1-800-227-2572.For questions regarding this recall call 800-662-1237. |  | 115142 | 111,504 units US and 3,638 Canada and Mexico- | Worldwide Distribution - USA (nationwide) and Internationally Mexico and Canada. | Open, Classified |  |  | 115142 | Undetermined
Z-1444-2014 | 67714 | No | LGZ | Infusion pump. | General Hospital | General Hospital | 510(k) | Smiths Medical HOTLINEÂż Fluid Warming Set L-80. |  |  | March 17, 2014 | 2014 | Smiths Medical ASD, Inc. | Injection port of the Set may leak blood/ IV fluids from the septum of the injection port or the septum may become dislodged | Nonconforming Material/Component | Smiths Medical issued Urgent Field Safety Notice Recall on 3/17/14 via mail service. Distributors will be instructed to notify their customers. The Field Safety Notice will also include an Urgent Recall Notice Confirmation Form (Confirmation Form) that consignees will be instructed to send back to Smiths Medical for carrying out the action and tracking effectiveness.Accounts are to inspect inventory and return recalled product. If you should have any questions regarding this information, please contact Shirley Duggan,Quality System Engineer at 1-781-763-9335 or by E-Mail Shirley.duggan@smiths-medical.com. |  | 870 | 870 units | Worldwide Distribution - USA Nationwide in the states of: AU, CA, CL, CN, CZ, GT, ID, IN, IT, JP, LY, NZ, QA, SG, SK, TR, TW. | Terminated | June 09, 2015 | 449 | 394080 | Undetermined
Z-1436-2014 | 67724 | No | JPK | Hematology quality control mixture. | Hematology | Hematology | 510(K) Exempt | Boule Con-Diff Tri-Level |  |  | February 24, 2013 | 2013 | Clinical Diagnostic Solutions | Elevated MCV results on the 1311-682, 1311-683, and 1311-684 lots. | Under Investigation by firm | Clinical Diagnostic Solutions sent a Notification letter dated February 26, 2014 to affected customers. The letter identified the affected product, problem and the actions to be taken. For questions contact Technical Support at 1-800-453-3328. |  | 3540 | 3,540 | Worldwide Distribution: USA Nationwide and the countries of IVORY COAST, MEXICO, and Philippines. | Terminated | August 07, 2014 | 529 | 3540 | Undetermined
Z-0717-2014 | 66030 | No | NJT | Immunohistochemistry reagents and kits. | Hematology | Pathology | 510(K) Exempt | Leica BIOSYSTEMS; BOND Polymer Refine Red Detection; Catalog No: DS9390; Leica Biosystems Newcastle Ltd; Balliol Business Park West; Benton Lane; Newcastle Upon Tyne NE12 8EW; United Kingdom; www.LeicaBiosystems.com; Bond Polymer Refine Red Detection is a biotin-free, polymeric alkaline phosphate (AP)-linker antibody conjugate system for the detection of tissue-bound mouse and rabbit IgG and some mouse IgM primary antibodies. |  |  | July 31, 2013 | 2013 | Leica Biosystems Newcastle Ltd. | The affected lots of the product may contain a contaminant of fungus that can cause a loss of staining sensitivity and/or obstructions to fluid delivery to the slide. | Process control | Leica Biosystems sent an Urgent Field Safety Correction Notice dated July 31, 2013, to all affected customers. The notice requests that direct accounts either destroy any unused reagents or confirm that the product has been used and in no longer in stock. Direct accounts are requested to sign the notice and either fax or scan and e-mail it back to Leica Biosystems Newcastle Ltd. The notice further requests that direct accounts who further distributed the recalled products forward the notification to their customers.Customers with questions were instructed to contact their Leica representataive if a replacement is necessary, for any questions customers were instructed call 847-405-5413 or microsystems.com.For questions regarding this recall call 847-405-5413. |  | 2344 | 1) 2,304 units; 2) 37 units | Worldwide Distribution - USA including AL, AR, AZ, CA, CO, CT, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, MT, NC, NH, NJ, NM, NY, OH, OK, PA, SC, TN, TX, UT, VA, WA, WI, and WV); and Internationally to Chile, Denmark, Finland, France, Germany, Ireland, Italy, Netherlands, Sweden, Switzerland, Poland, Romania, Singapore, South Korea, and the United Kingdom. | Terminated | August 14, 2017 | 1475 | 2344 | Undetermined
Z-0653-2014 | 66870 | No | DKZ | Amphetamine test system. | Toxicology | Toxicology | 510(k) | Product Brand Name(s): |  |  | June 03, 2013 | 2013 | Alere San Diego, Inc. | Alere initiated this recall of certain lots of the Oral Fluid Drug Screen Device because the sponge may become dislodged from the Saliva Collector handle prior to or during the oral fluid sample collection process. If the sponge becomes dislodged during the specimen collection process, the donor may inadvertently swallow the sponge which could present a choking risk. | Other | Firm sent customers notification letter via combination of fax, email or direct mail. The notification letter dated June 3, 2013, titled "URGENT PRODUCT SAFETY NOTICE", informed customers of the recall, reason for recall, product description with lot codes, potential hazard statement, customer required actions, and contact information. |  | 280710 | 280,710 | US Distribution: including states of: AL, CA, PA, and VA. | Terminated | March 18, 2014 | 288 | 280710 | Undetermined
Z-1460-2014 | 67780 | No | IZZ | Radiologic table. | Radiology | Radiology | 510(K) Exempt | AXIOM Vertix MD Trauma systems |  |  | March 05, 2014 | 2014 | Siemens Medical Solutions USA, Inc | There is a potential issue and possible hazard to patients when using the AXIOM Vertix MD Trauma systems. In rare cases, steel ropes inside the lift column of the system can be defective without triggering the safety lock, which can result in the U-arm dropping down unexpectedly during movement in vertical direction, potentially causing serious injury. | Component design/selection | Siemens sent a Safety Advisory Notice dated March 5, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Customers were advised as a first check it is strongly recommended for the users to check whether metallic dust or rubbed off parts of metal are visible underneath the lifting column or around the system. If this is the case, it is strongly recommended to immediately stop using the Vertix MD Trauma system and call the local Siemens service. To avoid any risk until the implementation of the modification mentioned below, it is furthermore strongly recommended to perform up/down movements of the lifting column not directly above the patient, but complete the vertical movement beside the patient and then move the system horizontally above the patient.We appreciate your understanding and cooperation with this Safety Advisory Notice and ask you to immediately instruct your personnel accordingly. Please ensure that this Safety Advisory Notice is placed in the system's instructions for use until the update has been installed. If you have sold or otherwise disposed of this equipment and it is no under your control, we kindly ask that you forward this Safety Advisory Notice to the new user of the equipment. Please also inform us about the new owner of the equipment.We apologize for any inconvenience this may cause. Further questions please call (610) 219-6300. |  | 2 | 2 | US Distribution including MO and OH. | Terminated | March 16, 2015 | 376 | 2 | Undetermined
Z-1247-2014 | 67775 | No | ERL | Ear, nose, and throat electric or pneumatic surgical drill. | Ear Nose &amp; Throat | Ear Nose &amp; Throat | 510(k) | 2.5MM Aggressive Blade |  |  | March 18, 2014 | 2014 | Stryker Instruments Div. of Stryker Corporation | Stryker Instruments is voluntarily recalling the single use 2.5MM Aggressive Blade because the line on the affected part (P/N 5290-928-000, Lot #13280CG2) was a straight line without a number scale or logo. The line required for this part number includes a number scale that goes from 1cm  7cm. | Employee error | Stryker Instruments sent an Recall Notification dated March 18, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Customers were instructed if you have further distributed this product, please forward this letter and the attached Business Reply Form (BRF)to all affected locations. Please indicate each location on the BRF. A FedEx shipping label will be emailed to each account upon receipt of the signed Business Reply Form. This label will allow the account to return the product to be destroyed.A Business Reply Form will be included in each mailing. Return of this form will serve as confirmation that the account has received the notification and executed the recall. Follow up phone calls, emails, and faxes will be sent to non-responding accounts as necessary.Product destruction will be handled at Stryker Instruments, Kalamazoo. Labeling will be defaced and the Aggressive Blades will be disposed of in the trash compactor.For questions regarding this recall please contact Stryker Instruments:Monday-Friday 8am-5pm (EST) 269-389-2921 |  | 90 | 90 units | US Distribution including the states of NV, CA, WA, TN, OR, SD, ID, IN and CO., and Internationally to Malaysia | Terminated | May 15, 2014 | 58 | 90 | Undetermined
Z-0634-2014 | 67134 | No | IOR | Mechanical wheelchair. | Physical Medicine | Physical Medicine | 510(k) | Tracer IV wheelchairs, Heavy duty manual wheel chair, 24" seat width, 450 lbs. weight capacity. Distributed by Invacare, Elyria, OH 44035 |  |  | April 16, 2012 | 2012 | Invacare Corporation | The wheels on Tracer IV wheelchairs (Tracer IV 24" seat width, 450 lbs. weight capacity manual wheel chair) were warped and wobbled when rolled. | Component design/selection | The firm, Invacare, sent a memo dated December 4, 2013 to their customer. The memo described the product, problem, and actions to be taken . On February 8, 2012 representatives from the recalling firm visited the customer to investigate the complaint. The wheelchairs were returned and destroyed.If you have any question, please call 440-329-6875 or email: ssaigal@invacare.com. |  | 227 | 227 Wheelchairs | US Distribution: Michigan only. | Terminated | January 10, 2014 | 634 | 227 | Undetermined
Z-1453-2014 | 67893 | No | DQY | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | Pipeline Embolization Device (PED). Used endovascular treatment of adults with intracranial aneurysms in the internal carotid artery. |  |  | April 01, 2014 | 2014 | Micro Therapeutics Inc, Dba Ev3 Neurovascular | Potential for the PTFE (polytetrafluoroethylene) coating to delaminate and detach from the delivery wire. | Process change control | A customer notification letter was sent on 4/1/14 to all customers who purchased the Pipeline Embolization Device and the Alligator Retrieval Device. The letter informs the customers of the problems identified and the actions to be taken. Customers are instructed to complete the customer response form and fax the completed form to Covidien at (763) 591-3519. Customers with questions are instructed to contact Dr. Mark A. Turco at (202) 310-5120.A press release was sent on 4/11/14 to Business Wire Newsroom to inform the public about the recall from Covidien. The recall is regarding certain lots of its Pipeline Embolization Device and Alligator Retrieval Device where the polytetrafluoroethylene (PTFE) coating applied to the delivery wire could delaminate adn detach from the devices. Consumers are instructed to report any adverse reactions or quality problems experienced with the use of the product to the FDA's MedWatch Adverse Event Reporting program either online, by regular mail or by fax. Consumers with may also contact Covidien Customer Service at (800) 716-6700 between 7am to 7pm (central) or email at customerserviceus@covidien.com. |  | 32 | 32 units total (23 units in US) | Worldwide Distribution-USA (nationwide) including Puerto Rico and the countries of Australia, Germany, France, UK, Turkey, Argentina, Canada, Colombia, Mexico, , Austria, Belgium, Denmark, Hungary, Italy, Netherlands, Norway, Poland, Portugal, Russia, Slovakia, Slovenia, Spain, Switzerland, and UAE. | Terminated | October 31, 2014 | 213 | 32 | Undetermined
Z-0638-2014 | 67005 | No | GDZ | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Sandel Healthcare Safety Devices Weighted Safety Scalpel #3 Disposable Safety Handle with #11 Blade |  |  | November 12, 2013 | 2013 | Ansell Sandel Medical Industries LLC | Select lots of SANDEL Weighted Safety Scalpels are packaged with the scalpel blade exposed beyond the safety shield which could compromise the integrity of the package and the sterility of the scalpel leading to injury. | Under Investigation by firm | The firm, Ansell, sent an "URGENT: VOLUNTARY RECALL OF ANSELL SANDEL WEIGHTED SAFET SCALPEL" letter dated November 11, 2013 to its customers and posted a website notification on November 11, 2013. The letter described the product, problem and actions to be taken. The customers were instructed: Do not ship SANDEL Weighted Safety Scalpel products 2210-L, 2210-N, 2211-L, 2211-N, 2215-L, 2215-N, or 2220-N to a hospital or end user with the listed lot numbers; quarantine SANDEL Weighted Safety Scalpel products in your inventory; and complete the attached response form and return it to attention: Ansell; via fax at 818-534-2510 or email sandel-cs@ansell.com. Additionally, if you have distributed this product to other accounts, please notify them of the voluntary recall and instruct them to return any outstanding stock to you as indicated in the attachment.If you have any questions or need additional information, please contact Customer Service by calling 866-764-3327 (6:00am-5:00pm PST/9:00am-8pm EST) or email sandel-cs@ansell.com |  | 96 | 96 cases (288 units) | Worldwide Distribution: US (nationwide) including states of: AL, CA, CO, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, MN, MS, NC, NE, NY, OK, PA, PR, RI, TN, TX and WI; and country of: Canada. | Terminated | April 06, 2016 | 876 | 1521 | Undetermined
Z-1321-2014 | 67861 | No | EJF | Orthodontic appliance and accessories. | Dental | Dental | 510(K) Exempt | Titanium Orthos Brackets, an orthodontic bracket. 022 +9/+9/0, Part Number 449-6212, Lot Number 121359169. A fixed attachment on a tooth which holds an archwire during orthodontic treatment. |  |  | March 25, 2014 | 2014 | Ormco/Sybronendo | Ormco Corporation inititated recall of U2R Titanium Orthos (0.22+9/+9/0), Part Number 449-6212, Lot Number 121359169, because the color ID dot was marked on the wrong tie-wing of the brackets. | Other | Recall notification letters sent to all affected consignees via 1st class mail on March 24, 2014. The consignees are instructed to complete the Return Form and have the option to return any affected product in their inventory to Ormco Corporation.The letter is dated "March 24, 2014" and titled: "URGENT: MEDICAL DEVICE RECALL". The letter informed customers of the recall, reason for recall, product description with part number and lot number, instructions to follow, and contact information. The letter was accompanied by a "Recall Return/Acknowledgement Form". |  | 150 | 150 units | US and Canada. | Terminated | October 27, 2014 | 216 | 150 | Undetermined
Z-0649-2014 | 67150 | No | OOE | Ophthalmic laser. | Ophthalmic | Ophthalmic | 510(k) | Catalys Precision Laser System (United States) - Catalys-U: |  |  | December 20, 2013 | 2013 | Optimedica Corporation | Potential for unexpected heating of the transformer mounting bolt located within the system chassis and the subsequent melting of a non-slip computer shelf mat, which may manifest an odor and visible smoke. | Under Investigation by firm | The firm, OptiMedica, issued an "ADVISORY NOTICE" dated December 20, 2013 to all customers with a Catalys Precision Laser System via Federal Express. OptiMedica also released a Field Service Bulletin (Document No: SB-00065) internally for the impacted product with instructions to implement the corrections through the OptiMedica service organization. The notice describes the product, problem and actions to be taken. The customers were instructed to complete the site header information on the attached fax-back form and return via fax to 1-408-890-4681 within the next three (3) business days or mail to OptiMedica at 1310 Moffette Park Drive, Sunnyvale, CA 94089.If you have any questions related to the operation of the Catalys Precision Laser System or the documentation requirements associated with this letter, please call OptiMedica Technical Support at 1-855-662-0911 and you will be directed to an OptiMedica representative who will be pleased to assist you. |  | 78 | 78 (39 in US, 39 internationally) | Worldwide Distribution: US (nationwide) including states of: TX, NC, OH, IA, UT, KY, ME, CA, NY, MA. NJ, IL, PA,, TN, FL, MO, GA, HI, SC, NE, MI, and GA and countries of: Dominican Republic, Germany, Switzerland, Australia, Spain, Japan, United Kingdom, Canada, France, The Netherlands, Austria, and Singapore. | Terminated | June 26, 2014 | 188 | 78 | Undetermined
Z-1314-2014 | 67647 | No | DYJ | Orthodontic appliance and accessories. | Dental | Dental | 510(K) Exempt | AOA Distalizer Lock Nut Screw, found in products listed as: Distal Jet, Spring Jet, Mesial Jet, and Uprighter Jet. Removable functional orthodontic appliances that are to modify the growth of the jaws in a prescribed growth pattern. |  |  | February 24, 2014 | 2014 | Allesee Orthodontic Appliances | The Lock Nut component of the Jet appliances may be missing the screw collar, which serves as a mechanical stop to prevent over-tightening in appliance assemby. Without this collar, an over-tightened lock nut may stop the piston assembly from moving freely, subsequently preventing intended movement of the teeth. | Nonconforming Material/Component | An advisory letter dated January 10, 2014 was sent to customers. The letter titled "ATTENTION: FIELD CORRECTIVE ACTION", informed customers the reason for recall, product description and devices affected dates of manufacturer, potential hazard, a photo comparison of a lock nut with screw collar and lock nut with no screw collar, instructions to the customers, and contact information. The letter was accompanied by an acknowledgement form titled "DISTALIZER JET LOCK NUT ACKNOWLEDGEMENT FORM".A follow up letter dated February 24, 2014 on March 7, 2014 to customers. This letter addresses the issue of the replacement screws as a follow up to the initial January 10, 2014 letter. The follow up letter, titled "ATTENTION: MEDICAL DEVICE RECALL" is dated Februray 24, 2014, refrences the January 10, 2014 letter that AOA sent, attaches the letter for reference, and informs customers that "AOA is now offering the option to obtain a Lock Nut with Screw Collar to replace within devices determined to be affected by this recall". The follow up letter informs customers of the reason for recall, the description and manufacturing dates of product affected, instructions to follow, and contact information. The customers were instructed to 1) continue to monitor the patient during the time of treatment; or 2) remove the device from the patient's mouth and return it to AOA for evaluation/replacement; or 3) request a replacement Lock Nut Screw with collar by contacting our Customer Support Department at 1-800-262-5221, and complete and return the attached acknowledgement form titled "DISTALIZER JET LOCK NUT RECALL ACKNOWLEDGEMENT FORM" via fax to 262-886-6879.If you have any questions, contact the Customer Support Department at 1-800-262-5221. |  | 411 | 411 | Worldwide Distribution: US (nationwide) and the country of : Canada. | Terminated | August 23, 2014 | 180 | 411 | Undetermined
Z-1327-2014 | 67819 | No | GCJ | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | 5/10/12 mm x 100mm Weck Vista Universal Cone Open Access, Catalog No. 405933. |  |  | March 11, 2014 | 2014 | Teleflex Medical | Complaints of leakage of insufflation gas through the device. | Nonconforming Material/Component | Letters were issued on March 11, 2014 to customers informing them of the recall. The letter asked the consignee to immediately discontinue use and quarantine any products on hand. The letter included a Recall Acknowledgement Form which is to be returned. A letter was also issued to distributors on March 11, 2014, requesting them to conduct a sub-recall. |  | 55521 | 55,521 total | Worldwide Distribution-USA (nationwide) including Puerto Rico and the states of AL, AZ, AR, CA, CO, CT, FL, GA, IL, IN, KS, KY, MD, MA, MI, MN, MS, MO, NH, NJ, NY, NC, ND, OH, OR, PA, SC, TN, UT, VA, WA, WV, WI and the countries of Australia, Brazil, Canada, China, Israel, Korea, Malaysia, Singapore, and Vietnam. | Terminated | January 28, 2016 | 688 | 55521 | Undetermined
Z-1458-2014 | 67609 | No | DZE | Endosseous dental implant. | Dental | Dental | 510(k) | BIOMET 3i Endosseous Dental Implant. |  |  | August 29, 2013 | 2013 | Biomet 3i, LLC | Due to residual machining fluid identified by discoloration (darker in appearance) on the external dental implant surface. | Process control | Biomet 3i sent an Urgent Medical Device Notice dated September 9, 2013, to all affected customers. The notice identified the product, problem and actions to be taken. Each customer was instructed to check their respective inventory for the affected product and to return any unused quantities along with the attached Recall Response Form. Biomet 3i will issue replacement product at no charge. For questions or concerns, call Biomet 3i at 1-800-342-5454. |  | 34744 | 34,744 | Worldwide Distribution - USA and countries of United Arab Emirates, Argentina, Austria, Australia, Belgium, Brazil, Canada, Switzerland, Chile, Colombia, Costa Rica, Germany, Dominican Republic, Spain, France, Great Britain (UK), Greece, Hong Kong, Ireland, Israel, India, Italy, Japan, Korea (South), Lebanon, Liechtenstein, Lithuania, Mexico, Netherlands, Poland, Portugal, Paraguay, Saudi Arabia, Sweden, Singapore, Turkey, Taiwan, Uruguay, Vietnam and South Africa. | Terminated | October 17, 2014 | 414 | 34744 | Undetermined
Z-1246-2014 | 67590 | No | JTN | Antimicrobial susceptibility test disc. | Microbiology | Microbiology | 510(k) | Oxoid Vancomycin 30 mcg, each cartridge contains 50 susceptibility discs, CT0058B. The firm name on the label is Oxoid Ltd., Basingstoke, Hampshire, England. |  |  | February 18, 2014 | 2014 | Remel Inc | Some of the discs may not be impregnated with the antibiotic. | Process control | ThermoFisher sent an Urgent Medical Device Recall letter dated February 18, 2014, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to notify any personnel who need to be notified of the potential for a false indication of resistance to Vancomycin. Customers were also instructed to review their inventory for the affected product and discard it. Customers should review results reported using this lot and consider retesting and/or seek appropriate expert advice. Return the attached Medical Device Recall Return Response acknowledging receipt of the notice and disposal of the affected product. If product was further distributed, customers should contact those entities, advise them of the situation and provide them with a copy of the letter. Customers with questions were instructed to contact the Technical Services Department at 800-255-6730 (US) and 913-888-0939 (International).For questions regarding this recal lcall 913-895-4185. |  | 200 | 40/5/50-disc cartridge packages | Nationwide Distribution | Terminated | June 25, 2014 | 127 | 200 | Undetermined
Z-1265-2014 | 67702 | No | HKI | Ophthalmic camera. | Ophthalmic | Ophthalmic | 510(k) | Halogen Lamp component in RetCam 3 System - |  |  | March 07, 2014 | 2014 | Clarity Medical Systems Inc | One lot of Halogen lamps have the potential to fail sooner than the expected life of the lamp which could result in damage limited to the internal electrical components inside the RetCam 3 computer module. | Nonconforming Material/Component | The firm, Clarity, sent a "Field Bulletin" dated March 12, 2014 to customers - risk managers and distributors by e-mail on March 12, 2014. They plan to follow this with a second letter on or about March 28, 2014 which will include a second bulletin, replacement bulbs and installation instructions. Clarity stated in the letter to its customers "In the meantime, you may continue to use your RetCAM 3."If you have any question of concerns, please contact Clarity at +1 925 621 322 or email to: service@retcam.com |  | 234 | 234 lamps | Worldwide Distribution: US (nationwide) to states of: TN, MA, CA, TX, MD,FL, MS, IL, and PA; and to countries of: France, Thailand, Egypt, Mexico, Kuwait, India, Austria,Japan, Sweden, Peru, United Kingdom, Philippines,The Netherlands, China,, UAE,Israel, Germany, Italy. | Terminated | August 11, 2014 | 157 | 234 | Undetermined
Z-0630-2014 | 67040 | No | KCY | Pneumatic tourniquet. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Stryker Color Cuff, REF 5921-034-135 |  |  | December 06, 2013 | 2013 | Stryker Instruments Div. of Stryker Corporation | The Color CuffÂż Sterile Disposable Tourniquet Cuffs with Quick-Connect Fitting may have an incorrect end-connector.A customer reported having received a luer connector instead of a quick connect fitting. An incorrect connector maylead to an inoperable system due to incompatibility with the Smart Pump. | Component change control | Distributor was notified by E-mail 11/26/2013, with follow-up UPS letters sent to Hospital Risk/ Materials Managers identifying the product, reason for the recall, and risk. The notice asked consignees to find and quarantine all implicated product, with hospitals requested to complete the Business Reply Form and FAX to 866-521-2762. Stryker will email a pre-paid shipper and credit all accounts. Question and concerns should be addressed to Kelly Jo Whipple 269-389-2921 or kellyjo.whipple@stryker.com |  | 5 | 5 boxes (50 units) | US Distribution: including states of: MI, OH, and TX. | Terminated | February 27, 2014 | 83 | 5 | Undetermined
Z-0642-2014 | 67034 | No | MEG | Piston syringe. | General Hospital | General Hospital | 510(k) | Jelco Smiths Medical Hypodermic Needle-ProÂż Fixed Needle Insulin Syringe, 1.0ml 29g x 1/2 in |  |  | July 10, 2013 | 2013 | Smiths Medical ASD, Inc. | Visible orange/ brown particulate on the needles of the 29g Hypodermic Needle-ProÂż Fixed Needle Insulin Syringes | Storage | The firm, Smiths Medical, sent an "URGENT FIELD SAFETY NOTICE" dated July 10, 2013 to its consignees/customers and distributors. The notice described the product, problem and actions to be taken. The Distributors were instructed to notify their customers. The consignees/customers were instructed to inspect your inventory for the suspect lots and remove all affected devices from use and complete and return the Urgent Field Safety Notice Confirmation Form by fax to 603-358-1017 or by email to FNS@smiths-medical.com within five (5) days of receipt of this notice. If you or your facility has distributed these affected products to other persons or facilities, please promptly forward the recipients of this Urgent Field Safety Notice. Upon receipt of the completed form, a customer service representative will contact you to arrange for exchange of your unused affected devices for credit or replacement.If you should have any questions regarding this information, please contact Smiths Medical's Customer Service Department at 800-258-5361 |  | 363015 | 363015 | Worldwide Distribution: US (nationwide) including states of: AZ, CA, FL, IL, LA, MO, MS, NH, NY, OH, PA, TN, TX, VA, WI, and WV; and countries of: Canada, Italy and Spain. | Open, Classified |  |  | 1125640 | Undetermined
Z-0491-2014 | 66744 | No | IZL | Mobile x-ray system. | Radiology | Radiology | 510(k) | Carestream DRX-Revolution Mobile X-Ray System -- Made in U.S.A. by Carestream Health Inc. 150 Verona Street, Rochester, NY . 14608 --- The DRX Revolution Mobile system is a mobile diagnostic x-ray system for generation of x-rays for examination of various anatomical regions. These products are mobile diagnostic x-ray systems designed to perform radiographic x-ray examinations on all pediatric and adult patients, in all patient treatment areas. |  |  | August 29, 2013 | 2013 | Carestream Health Inc. | It was discovered that Carestream received a report from a customer that 2 different patients were re-exposed due to the fact that the processed images were not able to be found in the system. | Radiation Control for Health and Safety Act | Carestream sent a letter to all affected customers informing them of the issue, how to detect the issue and actions to be taken by Carestream to correct the issue. The actions to be taken by Carestreamn to correct the issue will include the installation of a software patch to mitigate the effects of the observed failure and remove the possibility of a re-exposure in this failure mode. The correction will be implemented within the US to all affected devices at identified customer sites. For further questions call Carestream Customer Care at 1-800-328-2910, 8 am through 8 pm Eastern Standard Time. |  | 310 | 310 systerms domestically | Nationwide | Terminated | November 18, 2015 | 811 | 310 | Undetermined
Z-1313-2014 | 67582 | No | MEF | Teething ring. | Dental | Dental | 510(K) Exempt | Infantino Squeeze &amp; Teethe Monkey, "Go Gaga Squeeze &amp; Teether Coco the Monkey". Non-Fluid Filled Teething Ring. |  |  | January 06, 2014 | 2014 | Infantino LLC | Infantino initiated this voluntary recall of all lots of Infantino Squeeze &amp; Teethe Monkey, "Go Gaga Squeeze &amp; Teethe Coco the Monkey", Model Number 206-647, because the tail of the monkey may pose a choking hazard to young children. | Other | Infantino, LLC initiated this recall by notifying its customer on January 6, 2014 and Consumer Product Safety Commission (CPSC) posted the press release on 02/20/2014. CPSC posted the press release on the CPSC website. The CPSC recall press release dated February 20, 2014 informed consumers should stop using this product unless otherwise instructed and it is illegal to resell or attempt to resell a recalled consumer product. The CPSC press release contains a Recall Summary with the name of product, hazard, remedy, consumer contact information "Infantino toll-free at (888) 808-3111 between 8 a.m. and 4 p.m. PT Monday through Friday or online at www.infantino.com and click on Recall Information on the home page.", Recall Details on the number of units, description, incidents/injuries, remedy, sold exclusively at Target stores nationwide, Importer, Manufactured in China, and CPSC Consumer Information Hotline and Media Contact. The CPSC press release also contained photos of the product.Infantino, LLC also posted on their company website information on this recall. |  | 144000 | 144,000 units | US nationwide. | Terminated | April 23, 2015 | 472 | 144000 | Undetermined
Z-1459-2014 | 67774 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | SmartSite Extension Set, Model No. 20029E, intravascular administration set. |  |  | March 19, 2014 | 2014 | CareFusion 303, Inc. | Disconnections and leakages at the connection of the smallbore tubing and the SmartSite Y-port. | Nonconforming Material/Component | CareFusion sent an Urgent Medical Device Recall Notification letter dated March 2014 to all customers. The letter informs the customers of the problems identified, potential risk and the actions to be taken. Customers are instructed to complete and return to CareFusion the enclosed recall response card. Customers with recall related questions are instructed to contact CareFusion Support Center at (888) 562-6018. Customers with adverse event reports are instructed to contact Customer Advocacy at (888) 812-3266 or email customerfeedback@carefusion.com. Customers with technical questions regarding the Alaris System are instructed to contact Technical Support at (888) 812-3299. |  | 7700 | 7,700 units | USA Nationwide Distribution in the states of: AR, AZ, CA, CO, FL, ID, MD, MI, NC, NV, OK, OR, PA, TX, VA, WI | Terminated | June 06, 2015 | 444 | 7700 | Undetermined
Z-1461-2014 | 67837 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Brilliance CT 16, Brilliance CT 64, Brilliance CT Big Bore Oncology, Brilliance CT Big Bore Radiology, Brilliance iCT, Ingenuity CT, Ingenuity Core, Ingenuity Core 128 and Ingenuity Flex scanners are Computed Tomography X-ray systems. |  |  | March 27, 2014 | 2014 | Philips Medical Systems (Cleveland) Inc | Block assembly-lock stop subframe of service latch was not manufactured to engineering specifications from the supplier, causing premature fracture of the threaded rod on the patient support service latch. As a result, the table top can become free floating due to the disengaged service latch in the table tops subframe. | Nonconforming Material/Component | A Field Safety Notice was sent to all US consignees via certified mail. Market representatives will distribute the letter internationally. Additionally, a Field Service Engineer will contact each site to schedule a time to implement the Field Change Order. |  | 135 | 135 | Nationwide Distribution - USA including the states of CA, FL, IL, IN, KY, MD,SC, and TX. | Terminated | February 02, 2016 | 677 | 135 | Undetermined
Z-2132-2014 | 68766 | No | HQF | Ophthalmic laser. | Ophthalmic | Ophthalmic | 510(k) | MC-500 Multicolor Laser Photocoagulator; Distributed by NIDEK Inc. Manufactured by NIDEK CO., Ltd, Japan. For use in ophthalmic and retinal procedures. |  |  | June 09, 2014 | 2014 | Nidek Inc | Laser aperture label was not affixed to device prior to shipment in the US. | Error in labeling | Nidek sent a Recall/Field Correction letter dated June 9, 2014, to all affected customers. The letter advised customers of the issues and informed customers that Nidek Representatives will visit each location to affix the proper labeling to the device. Customers with questions were instructed to contact Nidek Customer Service at 1-800-223-9044 M-F 8 am to 5 pm PST.For questions regarding this recall call 510-353-7785. |  | 29 | 29 units (26 already re-labeled) | Nationwide Distribution including IL NE NY, NH, CD, MI, TX, CA, PA, FL, KS, UT, MN, WI,. MO. and OR. | Terminated | November 24, 2014 | 168 | 29 | Not_Computer
Z-0019-2015 | 67029 | No | FNL | AC-powered adjustable hospital bed. | General Hospital | General Hospital | 510(K) Exempt | ArjoHuntleigh Contoura C1000 series beds are designed to accommodate patients weighing up to 990 lbs. (450 kg). Model C1000 is the basic bed. Model C1080 includes a weighing system and a power outlet for an ArjoHuntleigh approved mattress pump. |  |  | March 17, 2014 | 2014 | Arjo Hospital Equipment AB | There is a low but existing entrapment risk for a patient being treated on the Contoura 1000 and Contoura 1080 bariatric beds if they have a BMI &lt; 40. | Labeling design | ArjoHuntleigh sent an URGENT FIELD SAFETY NOTICE - IMMEDIATE ATTENTION REQUIRED letters dated March 17, 2014 to all affected cusomers. The letters included instructions for customers to: 1) ensure that the Contoura Bed Field Safety Notice brought to the attention of all caregivers and clinical staff and they are aware that the Contoura 1000 or 1080 medical bed is intended to be used as a bariatric medical bed, designed for the patients with BMI greater than 40; 2) ensure that all users of the Contoura 1000/1080 bed, applicable staff and caregivers are trained to ensure awareness of the revised Instructions for Use Manual and new bed label; 3) discard all previous versions of Contoura 1000/1080 Bed Instructions for Use Manuals and ensure that a copy of this Contoura Bed Field Safety Notice is placed in the enclosed revised Contoura 1000/1080 bed Instructions for Use Manual (746-173-AM-3 - 10/2013); 4) where applicable, include this requirement in subsequent training of applicable staff and caregivers; and, 5) complete and sign the enclosed Customer Response Form and return this form to the ArjoHuntleigh Regulatory Department. An ArjoHuntleigh Service Technician will contact the designated facility representative listed on the Customer Response Form to schedule service to apply the two labels shown above, free of charge. For question contact ArjoHuntleigh at 1-800-323-1245, ext. 57985. |  | 1695 | 1,695 beds (worldwide) | Worldwide Distribution - US: Nationwide in the states of (AZ, CA, CO, CT, GA, ID, IL, MA, MD, MN, NC, NE, OH, PA, RI, SD, TX, VA, WI); and the countries of Australia, Austria, Bahrain, Belgium, Canada, Denmark, France, Germany, Greece, Iceland, India, Ireland, Israel, Italy, Lebanon, Mexico, Netherlands, New Zealand, Norway, Saudi Arabia, Singapore, Spain, Sweden, Switzerland, United Arab Emirates | Open, Classified |  |  | 1695 | Not_Computer
Z-2143-2014 | 68647 | No | IWZ | Radiographic film. | Radiology | Radiology | 510(K) Exempt | Ultra-speed Dental Film, Catalog Number/REF 834 8658, KODAK and Carestream DENTAL labels --- Carestream Health Inc. |  |  | July 02, 2014 | 2014 | Carestream Health Inc. | The Indicator Dot on the film packet's outer printed paper labeling was incorrectly positioned relative to the location of the raised dot on the dental intraoral film. This may cause confusion for the dental practitioner. If the processed film is placed on the view box by referencing the film's raised dot, the image orientation will be rotated 180 degrees. | Package design/selection | Recall notices were mailed to US Customers via overnight carrier on 07/02/2014. Customers are asked to complete an inventory of goods to be returned which will be ultimately destroyed by recalling company. If you have any questions or concerns, please contact the Carestream Customer Care Center in the U.S. at 1-800-933-8031. |  | 1633 | 1633 boxes | Worldwide Distribution. | Open, Classified |  |  | 1633 | Not_Computer
Z-1317-2014 | 67734 | No | MEH | Hip joint metal/ceramic/polymer semi-constrained cemented or nonporous uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Citation TMZF HA 132 degrees Neck Angle, V40 Hip Stem - Right Size #3: catalog # 6265-5113 , Left Size #4: catalog #6265-5104 , and Right Size #4: catalog #6265-5114; Stryker Howmedica Osteonics Hip Implants. |  |  | February 03, 2014 | 2014 | Stryker Howmedica Osteonics Corp. | Stryker received a report from the field indicating that a citation TMZF HA 132 degrees neck angle V40 Hip Stem - Left Size # 4 was identified in a packaging associated with a citation TMZF HA 132 degrees Neck Angle V40 Hip Stem - Right Size #3. | Error in labeling | Stryker sent an Urgent Medical Device Recall Notification letters and Product Accountability Forms to Hospital Risk Management, Chief of Orthopaedics, and Surgeons via Fed Ex with return receipt on February 26, 2014. The letter identified the affected product, problem and actions to be taken. Customer were instructed to return the affected product to the attention of Regulatory Compliance. Please attach the fluorescent orange PRODUCT REMEDIATION sticker to your return. For questions call 201-972-2100. |  | 48 | 48 units | USA Nationwide Distribution | Terminated | December 02, 2014 | 302 | 48 | Not_Computer
Z-2047-2014 | 68565 | No | MBI | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | SUTUREFIX Ultra S with 1 #2 Ultrabraid (blue) |  |  | June 02, 2014 | 2014 | Smith &amp; Nephew, Inc. Endoscopy Division | Sterility of device is compromised due to breach in sterile barrier | Package design/selection | Smith &amp; Nephew issued an Urgent - Product Recall Notification letter dated June 2, 2014 via Federal Express to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to inspect their inventory and locate all affected product and quarantine them immediately.Complete the Return Certification Form .Contact the Smith &amp; Nephew Returns Group at 800-343-5717 (option 3) or send an email to ProductRecoverv@smith-nephew.com. They will provide customers with instructions for returning their product and receiving a replacement.For questions regardin g this recall call 978-749-1073. |  | 907 | 907 units | Worldwide Distribution - USA (nationwide) and Internationally to Australia, Belgium, Denmark, France, Germany, Great Britain, Italy, Netherlands, Norway, Portugal, Spain, and Sweden. | Open, Classified |  |  | 3151 | Not_Computer
Z-1337-2014 | 67665 | [6 empty] | Syphilis Test Strip (Serum) |  |  | March 04, 2014 | 2014 | LumiQuick Diagnostics Inc. | Products shipped to US locations that do not have approved applications for PMA , IDE or 510(k) clearance. | No Marketing Application | Urgent Device Recall notification letters were sent to all US consignees on March 4, 2014 by e-mail and US mail as well as a phone call. The letter identified the products and actions to be taken by the customers. |  |  | ALL | Distributed in the states of CA, FL, NJ, TX, and NY. | Terminated | May 09, 2014 | 66 |  | Not_Computer
Z-1417-2014 | 67757 | No | GCY | Nonpowered, single patient, portable suction apparatus. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | 7mm Drain Kit, Flat Full Perforated, Saber 100ml Bulb and 3C System Sterile (10/case) |  |  | March 20, 2014 | 2014 | Aspen Surgical Products, Inc. | The recall has been initiated due to a potential compromise of the sterile barrier of the Surgidyne Wound Drainage and Surgidyne Bulb Evacuator &amp; Collection System medical devices. | Packaging process control | Aspen Surgical sent an Urgent Medical Device Recall letter dated March 20, 2014, to all affected customers and will be carried out at the distributor and retailer level, which can include doctor's offices, hospitals, and clinics. The distributor was instructed to contact their local FDA District office, conduct a sub-recall and to forward this Recall Notification and Response form to their retailers.A written notification to the consignees was sent by certified mail.Aspen has given the customers the option to either 1.) Return the product 2.) Destroy the product with a Certificate of Destruction to be provided to Aspen. Any returned product to Aspen will be put on MRB and dispositioned by Quality. The returned product will not be reprocessed.In order to determine effectiveness of the recall, any consignee that does not return the response form and/or product, will be contacted within three weeks of receipt of the Recall letter. Non-responders will be contacted a total of three (3) times in order to verify the effectiveness of the recall activity.Customers with questions were instructed to call 888-364-7004, ext 177 or 1-888-364-7004, ext. 156.For questions regarding this recall call 616-698-7100. |  |  |  | Worldwide Distribution - USA including LA, MA, NH, TX, FL, OH, NC, PA, CA, UT, TN, MI, CO, WY, OK, WA, MN, PA, IN, VA, IA, IL, MD, CT, DE, SC, NY, AR, MA, AZ, NJ, MO, GA, KS, WV, NE, NC, Puerto Rico, and Internationally to Canada, UAE, England, Italy, Belgium, Costa Rica, and Ecuador. | Terminated | November 25, 2014 | 250 |  | Not_Computer
Z-0488-2015 | 69642 | No | HHK | Endometrial suction curette and accessories. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | Cooper Surgical CANNULA-CURETTE 12MM; Model Number: MX562 |  |  | October 27, 2014 | 2014 | CooperSurgical, Inc. | The product has been identified to contain an incorrect curette type. | Packaging process control | CooperSurgical sent an Recall letter dated October 27, 2014, via UPS to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. The subject device does not meet the required release specification as it contains an incorrect curette type. No adverse events or injuries have been reported to date in connection with the use of the CANNULA CURETTE. The event was discovered during internal investigation. Products that do not meet the above criteria are safe to use alongside the DFU guideline.The letter requests that customers return the recalled product for a full refund or exchange. If you have any further questions please feel free to call (203) 601.5200. |  | 30 | 30 | US Distribution including the states of NY, MO and NV. | Terminated | May 28, 2015 | 213 | 30 | Not_Computer
Z-2152-2014 | 68749 | No | MTF |  |  | Immunology | PMA | VIDAS TPSA, in vitro diagnostic, for use as an automated quantitative measurement of prostate specific antigen levels in human serum or plasma. |  |  | October 14, 2013 | 2013 | BioMerieux SA | Product is not approved or authorized for distribution in the US. | Packaging process control | A letter was issued on October 14, 2013 to consignees requesting the consignee to check their stock and destroy all kits of the VIDAS TPSA ref. 30428. The letter also requested that the consignee notify their accounts of the recall. An acknowledgment form was also included with the letter which is to be returned to BioMerieux. |  | 1902 | 1902 | FL, LA, MI, MS, NV, OH, OK, TN, TX | Completed |  |  | 1902 | Not_Computer
Z-2041-2014 | 68468 | No | JSS | Microorganism differentiation and identification device. | Microbiology | Microbiology | 510(K) Exempt | BBL CrystaLEAD Enteric/Nonfermentor (E/NF) Identification (ID) System, catalog number 245000. in vitro diagnostic for the identification of aerobic gram-negative bacteria that belong to the family Enterobacteriaceae as well as some of the more frequently isolated glucose fermenting and nonfermenting gram-negative bacilli. |  |  | May 21, 2014 | 2014 | Becton Dickinson &amp; Co. | A portion of the BBL Crystal Enteric/Nonfermentor (E/NF) Identification kits from lot numbers 3224281 and 3290379 are labeled with incorrect expiration dates. A portion of these kits contain Crystal Inoculum fluid bottles that will expire prior to the date that was printed on the kit carton. The expiration date printed on the individual Crystal Inoculum fluid bottles is correct. | Labeling mix-ups | BD Diagnostic sent an Urgent Product Recall letter dated May 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Customers were instructed:Please discard any outer kit boxes that are labeled with the incorrect expiration date. All components of the kit will perform as expected with the expiration dates on each component.Our records indicate you may have been shipped a portion of the above-referenced product between August 23, 2013, and the present day.Please complete the attached form whether or not you have any inventory remaining so that we may acknowledge your receipt of this notification. It is not necessary to contact BD by telephone. Simply complete and fax the form to: Attention: Regulatory Compliance, Fax: 410-316-4258.If further assistance is needed, please contact the BD Customer Service Department at 1-800-675-0908. For all other inquiries, please contact the BD Technical Services Department at 1-800-638-8663.Please accept our apology for any inconvenience this may cause. BD is committed to providing you with the highest quality products. Thank you for your continued support. : -"/ 06 . ;/ i )!.U |  | 18.21 | 18.210 | Worldwide Distribution - US (nationwide), Australia, Belgium, Canada, Colombia, China, Indonesia, India, Japan, Korea, Mexico, Malaysia, Peru, Philippines, Singapore, Thailand, and Taiwan. | Terminated | August 12, 2016 | 814 | 18.21 | Not_Computer
Z-0484-2015 | 69636 | No | HGI | Gynecologic electrocautery and accessories. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | CooperSurgical LEEP RADIUS LOOP ELECTRODE (W 2.0 cm x D 1.0 cm D, 12 cm shaft ). Box of 5 |  |  | October 27, 2014 | 2014 | CooperSurgical, Inc. | Outer box is mislabeled as P/N R2010- Large Radius Loop Electrode instead of P/N R1010- Medium Radius Loop Electrode | Packaging process control | CooperSurgical issued recall letter dated October 27, 2014, to all affected customers via certified mail delivery (UPS). Acccounts were requested to return product for replacement and complete the return response form. Customers with questions were instructed to call 203-601-5200. |  | 595 | 595 units | Worldwide Distribution - USA (Nationwide) and Internationally to Canada, Portugal, and Hong Kong. | Terminated | May 29, 2015 | 214 | 595 | Not_Computer
Z-2151-2014 | 67640 | dam, rubber | EIE | Rubber dam and accessories. | Dental | Dental | 510(K) Exempt | Latex Dental Dams labeled as Non-Latex Dental Dams |  |  | December 11, 2013 | 2013 | Coltene Whaledent Inc | ColtÂżne/Whaledent Inc. received a complaint about the dental dams being the incorrect color. During an internal investigation, the firm found that Latex dental dams had been packaged in Non-Latex packaging. | Packaging change control | The firm issued a press release, called the distributor and sent a certified letter. The firm obtained the distributor's distribution list and sent recall notifications to all of these consignees.The firm included envelopes and return/reply cards, via response form mailed, certified mailing with return receipt. |  | 480 | 32 boxes, 15 individual dams per box | Distributed in Memphis, TN. | Open, Classified |  |  | 480 | Not_Computer
Z-0076-2015 | 69290 | No | DQY | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | Multi-Lumen Central Venous Catheterization |  |  | September 18, 2014 | 2014 | Arrow International Inc | Arrow International, Inc. has initiated a voluntary recall for product code AK-12703, lot number RF3039028 because the label on the outer corrugate is labeled with the incorrect expiration date. All kits packaged within the corrugate are correctly labeled. There is no risk to the patient as the kits within the corrugate are labeled with the correctexpiration date. | Labeling False and Misleading | Arrow International sent an Medical Device Advisory Notifications letter dated September 18, 2014, sent to direct accounts to notify them about the product, problem, and actions to be taken.Arrow International, Inc. ( Arrow ) is notifying customers that the label on the outer corrugate for the above mentioned product code and lot number is labeled with the incorrect expiration date. All kits packaged within thecorrugate are correctly labeled. There is no risk to the patient as the kits within the corrugate are labeled with the correct expiration date.Arrow International, Inc. is committed to providing high quality, safe and effective products. We sincerely apologize for any inconvenience this action may cause your operations. If you have any other questions, feel free to contact your local sales representative or Customer Service at 1-866-246-6990. |  | 65 | 65 | US Distribution including the states of MD, NC , LA , IN , OH, KY, CA , PA , NY and NV. | Terminated | December 02, 2014 | 75 | 65 | Not_Computer
Z-1005-2014 | 66500 | No | KTT | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | VERSA-FX II KEYLESS TUBE, VERSA-FX II KEYLESS SHORT, VERSA-FX II SHORT TUBE PL, VERSA-FX II STD TUBE PLAT and VERSA-FX II SUPRACONDYLR |  |  | October 11, 2013 | 2013 | Zimmer, Inc. | Zimmer recently conducted a review of all historic packaging validations at its Zimmer Manufacturing B.V. (ZMBV) facility in Ponce, Puerto Rico. Based on this review, it was determined that the packaging operations conducted in the Building 1 packaging operation at this facility were not properly validated. As a result, Zimmer conducted seal integrity testing to evaluate if previously packaged product from this facility met the established requirements for sterile barrier integrity. The acceptance criteria established for this testing was 99% reliability with a 90% confidence interval. These packaging tray configurations met the acceptance criteria for sterile barrier integrity of both the inner and outer trays, however, Zimmer has decided to recall all manufacturing lots that have not expired and that were sterile packaged in the ZMBV facility Building 1 packaging operations. | Packaging | Distributors received a letter identifying the issue and their responsibilities dated November 4, 2013. These responsibilities include quarantining the affected fielded inventory in their possession, providing hospital risk managers with the recall notification letter and retrieving affected fielded inventory and direct sales from the hospitals. |  | 7044680 | 7,044,680 total | Worldwide Distribution-USA (nationwide) and the countries of Belgium, Canada, Canary Islands, China, Czech Republic, Denmark, Finland, France, Germany, Greece, India, Italy, Japan, Korea, Luxemburg, Netherlands, New Zealand, Norway, Portugal Reunion, Russian Federation, Singapore, Slovakia, Spain, Sweden, Switzerland, Turkey, Taiwan, Thailand, and United Kingdom. | Terminated | April 14, 2017 | 1281 | 7044680 | Not_Computer
Z-1121-2014 | 67490 | No | MAV | Angiographic injector and syringe. | Cardiovascular | Cardiovascular | 510(k) | Medtronic Everest 20cc Inflation Device; Model Number and Catalog Number: AC2205P |  |  | February 06, 2014 | 2014 | Medtronic Vascular | Incorrect master label. | Labeling mix-ups | Medtronic sent an Urgent Medical Device Recall letter on February 6, 2014, via UPS 3-day delivery. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to remove the affected product that remains in their inventory. Return the affected product to Medtronic: Contact Customer Service at 888-283-7868 and reference RADAR # 170276 to initiate a return and credit of product. Complete the attached Customers Confirmation Certificate and fax it to Medtronic at 651-367-0612 to the attention of Customer Focused Quality. If any affected product has been forwarded to another facility, customers were asked to notify that facility accordingly and facilitate the retrieval of the affected product. Customers with questions were instructed to contact their Medtronic representative.For questions regarding this recall call 888-283-7868. |  | 191 | 191 devices | Nationwide Distribution including TX, ME, IA, PA, OR, and NH. | Terminated | June 02, 2014 | 116 | 191 | Not_Computer
Z-2145-2014 | 68745 | No | HET | Gynecologic laparoscope and accessories. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | Instructions for Use for the blue endo MOREsolution Electromechanical Laparoscopic Morcellator System, Model 55-000011, consisting of the control unit, hand-grip, flexible shaft, trocars/obturators, sleeves, valves, and blades. |  |  | July 16, 2014 | 2014 | Market-Tiers Inc | Some pages of the Instructions for Use were omitted, including pages regarding the Indications for Use, Warning, and Contraindications. | Labeling Change Control | The recalling firm issued a letter dated 7/16/14 via Federal Express requesting the customer destroy the incorrect IFU and replace it with the corrected version. |  | 100 | 100 units | Nationwide Distribution. | Terminated | May 11, 2015 | 299 | 100 | Not_Computer
Z-1576-2014 | 68143 | No | EIA | Dental operative unit and accessories. | Dental | Dental | 510(k) | Enaly 1000 BT-12 Ozone Generator, Model # 1000BT-12 |  |  | April 23, 2014 | 2014 | William C. Domb, D.M.D., A Professional Corporation | William Domb is recalling the Enaly 1000 BT-12 Ozone Generator because it is not approved or cleared by the FDA for medical use. | No Marketing Application | William Domb, DMD sent an Urgent Medical Device Recall letter on April 23, 2014, to all affected customers who purchased the Enaly 1000BT-12 Ozone Generator. The letter informed the customers of the problems identified, risk to health, and the actions to be taken. Customers were instructed to discontinue use of the Enaly 1000 BT-12 Ozone Generator and return their units to the following address:William Domb, DMD190 N Mountain AvenueUpland, CA 91786Customers were also instructed to complete the Medical Device Field Action Return Response form. For questions customers should contact William Domb at (909) 949-2968 or by email at bill@smilestudio.com from Monday-Friday at 8am-5pm Pacific Time.For questions regarding this recall call (909) 949-2968. |  | 16 | 16 units | Worldwide Distribution - USA including IN, HI, WA, TN, CA, TX, FL, NC, and Internationally to Canada, Portugal, and Ireland. | Terminated | January 08, 2015 | 260 | 16 | Not_Computer
Z-2124-2014 | 68744 | No | LRO | Surgical drape and drape accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | Enforcement Discretion | DeRoyal KNEE ARTHROSCOPY PACK, REF 89-6192.04, 1 Per Pack, Distributed by: DeRoyal Industries, Inc.; Manufacturer: DeRoyal Industries, Inc. |  |  | June 04, 2014 | 2014 | DeRoyal Industries Inc | The firm distributed surgical kits which contained Irrigation Sets which were subsequently recalled by Hospira. | Packaging change control | The recall was initiated by letter delivered via UPS on 6/4/2014. The firm requested that the consignee contact Stericycle to arrange for return and replacement of the product. A second notice was sent on 7/22/2014 to those accounts who did not respond to the initial notice. Distributors were requested to notify their end-users or provide a list of end-users to DeRoyal for direct notification. Effectiveness checks will be sent via email for those customers who have provided an email contact and via UPS for those customers who have not responded. |  | 128 | 128 units | Distributed in the states of MI, GA, OH, IL, MO, and IN. | Terminated | May 05, 2015 | 335 | 358 | Not_Computer
Z-0470-2015 | 69839 | No | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | Non-sterile drape |  |  | November 04, 2014 | 2014 | Novadaq Technologies, Inc. | Non sterile-drape LUNA drape was inadvertently mislabeled with an additional contents sterile label. | Error in labeling | Novadaq initiated their recall by telephone and followed by an email and a Medical Device Recall Notice dated November 11, 2014, to all affected customers. Distributors were instructed to notify their customers. The firm is removing the recalled products and replacing them with correctly labeled product. Customers were instructed to confirm receipt of the Recall Notice and their understanding of the actions to be taken by completing the attached confirmation form and faxing it to Novadaq Quality Department at 1-800-886-2419. For questions regarding this recall call 604-432-9861. |  | 324 | 324 units | Nationwide Distribution to IA, LA, GA, MD, ID, FL, TN, OH, TX, NC, MA, KY, KS, WI, IN, SC, OR, CT, CA, WA, PA, OK, and PR. | Terminated | November 06, 2015 | 367 | 1152 | Not_Computer
Z-2021-2014 | 68196 | No | CGA | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | ADVOCATE Redi-Code+ BMB-BA006A Blood Glucose test strips |  |  | May 05, 2014 | 2014 | DIABETIC SUPPLY OF SUNCOAST, INC. | "Advocate Redi-Code + BMB-BA006A Blood Glucose test strips" are inappropriately labeled and marketed as "Advocate Redi-Code + Blood Glucose test strips". | Labeling False and Misleading | Diabetic Supply of Suncoast, Inc. sent an Urgent Voluntary Medical Device Recall letter dated June 6, 2014, to all affected consignees. The firm is asking the distributors to examine their inventory immediately and quarantine any product subject to recall. In addition, if they may have further distributed this product, to please identify their customers and notify them at once of this product recall. Customers who have the BMB-BA006A Advocate Redi-Code+ test strips can contact Diabetic Supply of Suncoast, Inc. directly for further information to determine if the products are being used in the proper manner or if they need to be returned to Suncoast for replacement. Test strips will be replaced with relabeled (with new labels) boxes of test strips to avoid any chance of confusion. Consumers should stop using the above mentioned product and contact Suncoast to find out if this recall pertains to the products they have in their possession and if so how to have them replaced for the relabeled product manufactured by BroadMaster Bio-Tech. Customers with questions should call Diabetic Supply of Suncoast, Inc. at 561-296-0488.For questions regarding this recall call 561-254-1007. |  | 1852800 | 1,852,800 cases (50 strips ea) | Worldwide Distribution - USA (nationwide) and Internationally to Venezuela, Ecuador, Canada &amp; Vietnam | Terminated | June 13, 2017 | 1135 | 1852800 | Not_Computer
Z-0064-2015 | 69173 | No | CHL | Blood gases (PCO2, PO2) and blood pH test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Omni Micro-electrode/reference electrode for cobas b221 analyzer. |  |  | September 12, 2014 | 2014 | Roche Diagnostics Operations, Inc. | The default references for normal values are inconsistent between cobas b 221 and other blood gas instruments. In addition, the reference ranges are different when comparing with corresponding ranges listed in the Instructions for Use for the analyzer and the reference literature source. | Labeling Change Control | The Urgent Medical Device Correction will be sent via UPS Ground (receipt required) to the consignee list. The Urgent Medical Device Correction will be sent via UPS Ground (receipt required) to the consignee list.Consignees are instructed as follows:" Check the current set of reference values on your cobas b 221 blood gas system and, if appropriate, change the reference ranges accordingly." If you have not established your own reference ranges, please follow the instructions in the Workaround section of the UMDC or contact the Roche Support Network Customer Support Center. |  | 492 | 492 devicees | Nationwide Distribution including Puerto Rico and Guam, and the states of AK, AL, AZ, CA, CO, CT, FL, GA, HI, IA, IL, IN, KS, KY, LA, ME, MD, MI, MN, MO, MS, MT, NC, NH, NJ, NM, NY, OH, OK, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA and WI. | Open, Classified |  |  | 492 | Not_Computer
Z-0821-2014 | 67163 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Puritan Bennett 840 Series Ventilator PB840 Operators Manual Addendum PN 066009A 09/02 |  |  | January 25, 2013 | 2013 | Nellcor Puritan Bennett Inc. (dba Covidien LP) | Covidien is correcting labeling to clarify the operational life of the oxygen sensor (O2 sensor), PN 4-072214-00 used in the Puritan Bennett" 840, and PN, G-062010-00, used in Puritan Bennett" 740, and760 ventilator systems. The Operators Manual for PB840 states a nominal life of one year and the Addendum incorrectly states 2 years. | Labeling Change Control | The firm, Covidien, sent an "URGENT MEDICAL DEVICE VOLUNTARY FIELD CORRECTION" letter dated January 24, 2014 to its Consignees/customers via certified mail. The letter described the product, problem and actions to be taken. The customers were instructed to read and understand the enclosed information; communicate this information to those who need to be aware of this labeling update; for the Puritan Bennett 840 ventilator only: If you have addendum P/N 066009A 09/02, remove and discard the addendum and replace it with the new information in Attachment I; for the Puritan Bennett 740 and 760 ventilators only: Add the new information in Attachment I; Ensure all O2 sensors in use or in inventory conform to the instructions in this notification. to your current operators manual and service manual, and read and understand the warnings, cautions and notes related to safety, ventilator use and operation that are contained in the Preface of the Operators Manuals. are asked to remove incorrect labeling and replace with the revised labeling.Covidien enclosed a labeling update (Attachment I) which provides detailed information related to the life of oxygen sensors used with the Puritan Bennett 840, 740 and 760 ventilators. Covidien has validated another oxygen sensor, P/N 10097559, for use in the Puritan Bennett 840, 740 and 760 ventilators. Information about that sensor is also included in the update in Attachment I.Should you have any questions regarding this letter or to report any issues with the Puritan Bennett" 840, 740 or 760 ventilators contact our Technical Support Department @ 1-800-255-6774, option 4, then option 1, to ensure proper device reporting procedures are followed. |  | 70160 | 70,160 | Worldwide Distribution: US (nationwide) and Internationally to: Albania, Argentina, Australia, Bahrain, Belgium, Belize, Bermuda Bolivia, Bosnia and Herzegovina, Brazil, Cambodia, Canada, Chile, China, Colombia, Costa Rica, Croatia, Ecuador, Egypt, El Salvador, Finland, France, French Polynesia, Georgia, Germany, Greece, Guam, Guatemala, Honduras, Hong Kong, Hungary, India, Indonesia, Iran, Ireland, Israel, Italy, Jamaica, Japan, Jordan, Kazakhstan, Kenya, Kuwait, Lebanon, Libya, Lithuania, Macedonia, The former Yugoslav, Malaysia, Mexico, Moldova, Republic of, Montenegro, Morocco, Netherlands, Nicaragua, Northern Mariana Islands, Oman, Panama, Paraguay, Peru, Philippines, Poland, Puerto Rico, Qatar, Romania, Russia, Saudi Arabia, Singapore, Slovenia, South Africa, South Korea, Spain, Sri Lanka, Sudan, Switzerland, Syria, Taiwan, Thailand, Tunisia, Turkey, Uganda, Ukrainian SSR, United Arab Emirates (UAE) United Kingdom, Uruguay, Venezuela and Yugoslavia. | Terminated | November 24, 2014 | 668 | 78119 | Not_Computer
Z-0049-2015 | 69404 | No | GEX | Laser surgical instrument for use in general and plastic surgery and in dermatology. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | StarMedTec LightTrailÂż Reusable Fibers, 365 um; Material/Part Number: 6453 |  |  | September 30, 2014 | 2014 | Boston Scientific Corporation | Product is not cleared for use with lasers other than the Auriga XL system | No Marketing Application | Boston Scientific sent an " Urgent Medical Device Recall Removal Immediate Action Required" letter dated September 30, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. The notification instructs the account to check their inventory, remove any affected product from their inventory, and contact Boston Scientific for further instructions on returns. The notification also requests the completion of a Reply Verification Tracking Form as a method of documenting the presence or absence of affected product in their inventory.We ask that you complete and return the Reply Verification Tracking form, included with this letter, according to the instructions on page 3, even if you no longer have inventory of the recalled lots. If you identify any of the affected lots within your inventory, please segregate it immediately and you will be contacted by Boston Scientific with additional instructions after you return your form. You will receive credit, as appropriate, for any affected product in your inventory. Fax to : Field Action Center at 1-866-213-1806.For further questions call (508) 382-9555 |  | 5 | 5 units | US Distribution including the states of NY and TX. | Terminated | November 26, 2014 | 57 | 45 | Not_Computer
Z-2093-2014 | 68701 | No | BYG | Oxygen mask. | Anesthesiology | Anesthesiology | 510(K) Exempt | Oxygen Mask, Model No. 8120-7 |  |  | June 25, 2014 | 2014 | Salter Labs | The case label correctly states partial rebreather but the insert label incorrectly states non-rebreather. | Error in labeling | Salter Labs sent a Medical Device Recall letter dated June 25, 2014 to all affected customers. The letter identified the affected product, problems and actions to be taken. Customers are instructed to complete the attached Return Response Form and return it to Salter as soon as possible. Customers are instructed to send completed form to Salter by either Fax (800) 628-4690, attn: Sarah Velasquez, email to SVelasquez@us.salterlabs.com, or mail to Salter Labs, attn: Sarah Velasquez, 100 W. Sycamore Rd. Arvin, CA 93203. |  | 584 | 584 units | Worldwide Distribution - US (nationwide) and the countries of Brazil, United Kingdom, Indonesia, Morocco, Mexico, Norway, Venezuela, and South Africa. | Terminated | September 10, 2014 | 77 | 584 | Not_Computer
Z-1073-2014 | 67405 | Yes | KPO | Hemodialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | NxStage PureFlow-B Solution; Model/Catalog Number: RFP-401. Product is packaged in a case containing two individual 5L bags. |  |  | January 24, 2014 | 2014 | NxStage Medical, Inc. | Mislabeled product. | Packaging process control | NxStage sent a Voluntary Medical Device Recall letter dated January 24, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. They were asked to check all boxes in inventory for the recalled lot and segregate and quarantine them. In addition, they were asked to complete the recall reply form and return it by fax to NxStage Medical as well as to contact NxStage Customer service to arrange for return of all affected product and for replacement to be sent. Distributor was asked to notify all downstream customers and to provide NxStage Medical with a list of customers. Please complete and fax this form to 978-687-4810 even if you do not have any of the affected product on hand.If you have any questions or comments, please feel free to contact NxStage Customer Service at 1-866- NXSTAGE (1-866-697-8243). |  | 969 | 969 cases (2 bags per case) | US Distribution including the states of NY, TX, ID, AR, CA and AL | Terminated | August 29, 2014 | 217 | 969 | Not_Computer
Z-2074-2014 | 68590 | Yes | KDI | High permeability hemodialysis system. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Baxter MiniCap with Povidone-Iodine Solution, 5C4466P, Baxter Healthcare Corporation. For dialysis. |  |  | June 17, 2014 | 2014 | Baxter Healthcare Corp. | Inadequate iodine and packaging related defects. | Packaging process control | The firm notified their direct consignees on 06/17/2014 by letter requesting customers to locate and remove all affected product from their facility, and contact Baxter Healthcare Center for Service to arrange for return and credit. |  | 585360 | 585,360 units | Distributed USA (nationwide) and Canada. | Terminated | April 03, 2015 | 290 | 585360 | Not_Computer
Z-0066-2015 | 68987 | No | DQY | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | Arrow Select Kits (ASK) and Maximal Barrier Precautions Kits (CDC) with BD Medical Posiflush 10mL Saline Flush Syringe. |  |  | August 06, 2014 | 2014 | Arrow International Inc | The component supplier recalled their 0.9% Sodium Chloride Injection USP BD Posiflush SF Saline Flush Syringe 10mL because some unit packages may exhibit open seals which impact package integrity and potentially product sterility. | Packaging | Arrow International sent an Urgent Medical Device Recall Notification letter, dated August 6, 2014, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to immediately discontinue use and quarantine any affected product. Customers were instructed to remove the BD pre-filled saline syringe from the exterior of the packaging. To return the BD syringes, customers were instructed to complete the enclosed Recall Acknowledgement Form and fax it to 1-855-419-8507, Attn: Customer Service. A customer service representative will contact the customer with a Return Good Authorization (RGA) Number and will provide instructions for the return of producet to Arrow International. All customers were asked to complete the enclosed Recall Acknowledgement Form and fax it to 1-855-419-8507. Attn: Customer Service. Customers with questions were instructed to contact their local sales representative or Customer Service at 1-866-246-6990. |  | 71255 | 71255 | Nationwide Distribution | Open, Classified |  |  | 71255 | Not_Computer
Z-2115-2014 | 68703 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | Pad-Pak and Pedi-Pak accessories for the HeartSine Samaritan PAD 300P and 350P device (SAM 300P and 350P). |  |  | June 12, 2014 | 2014 | Heartsine Technologies, Limited | A small number of sealed foil pouches containing the electrodes were found to be difficult to open, resulting in potential delay in therapy. | Packaging process control | HeartSine notified distributors in the United States by e-mail on June 12, 2014 and by follow-up letter on June 13, 2014. Distributors were given a letter to use in communications with customers on June 17, 2014. Customers were instructed to return affected product to there distributors. Distributors were asked to return the affected product to HeartSine. In order to effect an immediate return, replacement Pad-Paks and Pedi-Paks have already been delivered to all respective distributors. A seconde letter, dated June 30th, was sent to end users to clarify where on the label to find lot specific information.For question contact Quality Manager at +44 (0)28 90 93 94 19. |  | 3245 | 3245 | Worldwide Distribution - US (nationwide) and in Canada, Mexico and Latin America. | Terminated | April 16, 2015 | 308 | 3245 | Not_Computer
Z-2113-2014 | 68709 | No | MVC |  |  | Pathology | PMA | InSite HER-2 Detection Kit, an immunohistochemistry assay for in vitro diagnostic use. |  |  | July 01, 2014 | 2014 | Biogenex Laboratories, Inc. | Product was distributed without 510(k) approval. | No Marketing Application | Each customer contacted by phone - neither is in possession of the product. |  | 4 | 4 kits | Brazil and Ecuador. No US distribution | Terminated | July 24, 2014 | 23 | 4 | Not_Computer
Z-2072-2014 | 68742 | No | GWN | Nystagmograph. | Neurology | Neurology | 510(k) | Integra Camino 110-4 Intracranial Pressure (ICP) Monitoring Catheters, |  |  | July 09, 2014 | 2014 | Integra LifeSciences Corp. | Integra LifeSciences is recalling the Camino Intracranial Pressure Monitoring Kits because of revised labelling to indicate they are unsafe in an MR environment. | No Marketing Application | A customer notification letter dated 7/9/14 was sent to customers to inform them that Integra LifeSciences has recently revised the labeling of certain Camino Intracranial Pressure Monitoring Kits to indicate they are UNSAFE in an MR environment. The letter informs the customers of the problems identified and the actions to be taken. Customers are instructed to complete the form attached the letter and return it by email or fax. Customers with questions are instructed to contact Customer Service at (855) 532-1723. |  | 28752 | 28,752 units | Algeria, Austria, Belgium, Bulgaria, Cyprus, Egypt, Finland, France, Germany, Great Britain, Greece, Hungary, Greece, Hungary, Ireland, Israel, Italy, Jordan, Lebanon, Lithuania, Luxemborg, Netherlands, Poland, Portugal, Romania, Saudi Arabia, Slovenia, South Africa, Spain, Sweden, Switzerland, Tunisia, Turkey, United Arab Emirates, Argentina, Brazil, Brunei, Chile, China, India, Japan, Korea, Mexico. | Terminated | October 14, 2014 | 97 | 28752 | Not_Computer
Z-0476-2015 | 69386 | No | LBH | Cavity varnish. | Dental | Dental | 510(k) | Dharma Cavity Varnish 15mL |  |  | September 29, 2014 | 2014 | Dharma Research, Inc. | During an FDA inspection it was found that the products are marketed without a cleared 510k. | No Marketing Application | On September 29, 2014 the firm sent a notification to its consignees informing them of the product being discountued and removed. The firm asked for the product to be returned directly to Dharma's office at 5220 NW 72 Avenue, Suite 15, Miami FL 33166. For any questions please do not hesitate to call (305) 482-9669. |  | 916 | 916 | Worldwide Distribution: US Distribution in states of: CA, PA, FL, IL, including Puerto Rico; and countries of: Honduras, Dominican Republic, Turkey, Nicaragua, Colombia, Pakistan, Trinidad &amp; Tobago, Guatemala, Phillipines, Saudi Arabia, and United Arab Emirates. | Open, Classified |  |  | 3758 | Not_Computer
Z-1552-2014 | 68074 | No | LFL |  |  | General &amp; Plastic Surgery | 510(k) | TX1 Tissue Removal System. TX1 Procedure Pack P/N 554-1003-001 (outer box), TX1 MicroTip P/N 554-1002-001 (inner tray). For use with the TX System Console as an ultrasonic surgical aspirator of soft tissue. |  |  | April 14, 2014 | 2014 | Tenex Health Inc | The sterile barrier in the packaging may be compromised due to cracks in the thermoform tray that may occur during transport of the product in extreme cold weather conditions, in conjunction with mechanical impact. | Packaging | Tenex Health sent an Urgent Medical Device Recall notification letter dated April 15, 2014, to all consigness. The letter informed consignees of the recall, manufacturer, product name, product code, reason for recall, recall carried out to the user level, instructions to customers, contact information, label examples, and response form was provided. Consignees were instructed to immediately examine their inventory per the enclosed instructions, if they identify any of the affected product within their inventory, they should segregate the product and immediately return only MicroTips with damaged trays to the address below, using the Tenex Health, Inc. FEDEX account number and referencing the Returned Goods Authorization (RGA) number.RGA#Tenex Health, Inc.26902 Vista TerraceLake Forest, CA 92630Customers were asked to complete and return the enclosed response form as soon as possible.Customers with questions should contact Tenex at (949) 454-7500, ext. 209or email tanaka@tenexhealth.com.For questions regarding this recall call (949) 454-7500. |  | 9927 | 9,927 units | Worldwide Distribution - USA (Nationwide) and Internationally to Panama and Singapore. | Terminated | July 16, 2014 | 93 | 9927 | Not_Computer
Z-0715-2014 | 67177 | No | HBE | Powered simple cranial drills, burrs, trephines, and their accessories. | Neurology | Neurology | 510(k) | Anspach MCA Cutting Burr, DFU Rx Only |  |  | March 29, 2012 | 2012 | The Anspach Effort, Inc. | Current directions for use for the MCA cutting burs (item 18-0001) do not provide enough details regarding the intended use specifically for the MCA burrs. | Error in labeling | Anspach sent an Urgent Medical Device Labeling Correction letter dated Marach 29,, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to include the item 18-0079 with their inventory immediately, complete the attached reply form indicating their receipt of the letter. Return the completed form by fax to the number or email provided on the form. If product was further distributed to other facilities, the correction letter should be forwarded to them. Customers with quesitons should call 1-800-327-6887 or email customer.support@synthes.com or ProductSupportTeam@Synthes.com.For questions regarding this recall call 561-627-1020. |  | 492 | 492 (321 US) | Worldwide Distribution - USA (nationwide) and Internationally to Great Britain, Australia, Switzerland, Sweden, France, Japan, Spain, Italy, Canada, and Netherlands. | Terminated | April 15, 2016 | 1478 | 492 | Not_Computer
Z-0819-2014 | 67149 | No | FTL | Surgical mesh. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Composix L/P with Echo PS 10" X 13" Product Code 0144113 |  |  | January 03, 2014 | 2014 | Davol, Inc., Subs. C. R. Bard, Inc. | Product labeling does not match product configuration. | Labeling Change Control | Davol, Inc., Subs. C. R. Bard, Inc.sent a customer notification letter on January 3, 2013, to all affected customers via Fedex to inform consignees that the Instructions for Use (IFU) and graphic contained onthe inside of the product carton bottom references two features added to the inflation tube that are not present on the product they received. Please contact the Davol Customer Service Department at 1-800-556-6275 or C. R. Bards Medical Services &amp; Support Department at 1-800-562-0027 if you have any questions. Consignees are asked to complete an effectiveness check form and return it via fax to 401-825-8753.For questions regarding this recall call 800-556-6275, |  | 19 | 19 | Nationwide Distribution including CA, PA, IL, VA, FL, GA, NY, TX, MO, NE, IN, NM, OH, IA. | Terminated | December 23, 2015 | 719 | 65 | Not_Computer
Z-0070-2015 | 69368 | No | OMP | Powered suction pump. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | V.A.C. GranuFoam Silver Large Dressing is barrier to bacterial penetration and may help reduce infection. Featuring SensaT.R.A.C. Technology REF M8275099/10 STERILE Single Use only Latex Free Rx Only. |  |  | September 18, 2014 | 2014 | KCI USA, Inc. | A portion of the V.A.C. GranuFoam Silver Dressing M8275099/10 lot 50775977 was incorrectly manufactured with an incorrect expiry date on the sterile pouch. The correct expiry date of 2015-11 is on the case label, however the pouch label identifies the expiry date of 2016-11. | Error in labeling | KCI will send a customer letter to each consignee notifying them of the issue, and to return the affected units. |  | 1090 | 1090 pouches (109 boxes) | The affected product was distributed to Mexico, Turkey, and United Arab Emirates only. The affected product was not distributed to the US. | Terminated | July 24, 2015 | 309 | 1090 | Not_Computer
Z-2069-2014 | 68480 | No | EJK | Calcium hydroxide cavity liner. | Dental | Dental | 510(k) | Calcium Hydroxide Powder, USP |  |  | June 05, 2014 | 2014 | Axcentria Pharmaceuticals, LLC | One lot of Calcium Hydroxide Powder is being recalled because it was manufactured and distributed without an approved 510k, and the current label does not include instructions for use. | No Marketing Application | Axcentria Pharmaceuticals sent an urgent medical device recall notice, dated June 5, 2014, was sent to customers and described the product, problem, and actions to be taken. Customers were asked to fill out and return the form included.For questions customers were instructed to call 215-453-5055, ext 317. |  | 528 | 528 | US distribution to NY and CA | Open, Classified |  |  | 528 | Not_Computer
Z-0183-2014 | 66610 | No | GCJ | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | Applied Medical's InziiÂż 12/15mm Retrieval System. |  |  | October 21, 2013 | 2013 | Applied Medical Resources Corp | Applied Medical is conducting a voluntary recall of the InziiÂż 12/15mm retrieval system. During shipment, the retrieval system packaging has the potential to become punctured with small holes, which could compromise the sterile barrier. The likelihood of this situation to occur and result in patient harm is highly unlikely; however, out of an abundance of caution for patient safety and a commitmen | Package design/selection | Customer notifications of the recall began the week of 10/21/2013 via UPS overnight with letters requesting that all affected products be returned to Applied. The notification letter is dated October 21, 2013 and titled "URGENT: MEDICAL DEVICE RECALL". The notification was accompanied by a Recall Notification Confirmation Form. The letter informed customers of the recall by providing the product description with codes (Model and Lot Numbers). reason for the recall, actions to take and contact information. |  | 29506 | 29,506 units | Worldwide Distribution: US (nationwide) and countries of: Australia, Austria, Belgium, Brazil, Canada, Chile, Colombia, Czech Republic, Denmark, Ecuador, Finland, France, Germany, Hong Kong, India, Ireland, Israel, Italy, Japan, Lebanon, Malaysia, Mexico, New Zealand, Netherlands, Norway, Peru, Poland, Portugal, Saudi Arabia, Singapore, Spain, South Africa, Sweden, Switzerland, United Arab Emirates, and United Kingdom. | Terminated | September 15, 2014 | 329 | 29506 | Not_Computer
Z-1574-2014 | 68109 | No | HTT | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | ANSPACH Cutting bur, 3 mm Diamond Ball, Standard Length. REF L-3SD Use with LONG Attachments. Used for cutting and shaping bone, including the spine and cranium. |  |  | January 08, 2014 | 2014 | The Anspach Effort, Inc. | Anspach product code L-3SD contained product code S-3SD which is shorter in length than the labeled product. | Labeling mix-ups | Anspach sent an "URGENT: MEDICAL DEVICE REMOVAL" letter dated January 13, 2014, to all affected customers. The letter described the product, the problem and actions to be taken. If they have distributed any of the products to other services or facilities, customers were instructed to please forward this information as appropriate. Customers with questions you contact Anspach Customer Support at (800) 327-6887, option 3 or email customer.support@synthes.com.. Customers were asked to complete the attached reply form confirm receipt of the letter, return the completed form by fax to 800-327-6661.For questions regarding this recall call 561-494-3706. |  | 30 | 30 cutting burs affected | Nationwide Distribution including FL, TX, NC, IL, CO, and PA. | Terminated | January 20, 2015 | 377 | 30 | Not_Computer
Z-2123-2014 | 68706 | No | GEX | Laser surgical instrument for use in general and plastic surgery and in dermatology. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | PHAROS Excimer Laser, Model No. EX-308 |  |  | June 30, 2014 | 2014 | Ra Medical Systems Inc | RA Medical Systems is recalling the PHAROS EX-308 Excimer Laser because there is a possibility that the laser may calibrate with the iris closed. | Error in labeling | Ra Medical Systems sent a customer letter dated June 30, 2014, to who purchased the PHAROS EX-308 Excimer Laser Phototherapy System Hand Piece. The letter informed the customers of the problems identified and the actions to be taken. Customers were also provided with the Instructions for Use (IFU) for the PHAROS EX-308 Excimer Laser Phototherapy System Hand Piece and a replacement iris indication label for the handpiece. |  | 435 | 435 units | Worldwide Distribution - USA (Nationwide) and Internationally to: Korea, Kuwait, and Saudi Arabia. | Terminated | September 09, 2014 | 71 | 435 | Not_Computer
Z-2090-2014 | 68718 | No | DRM | External cardiac compressor. | Cardiovascular | Cardiovascular | 510(k) | Zoll Circulation LifeBand Single Pack (accessory), and Nickel-metal hydride (NIMH) Battery. |  |  | July 16, 2013 | 2013 | Zoll Circulation, Inc. | Typographical error in the Japanese legal label of the AutoPulse Resuscitation System Model 100. The label is specific to the Japanese version of the product. | Labeling Change Control | Zoll Japan handled the recall. |  | 201 | 161 Nickel Metal batteries, 40 LifeBand | Distributed in Japan. | Terminated | July 22, 2014 | 371 | 201 | Not_Computer
Z-2156-2014 | 68807 | No | LFL |  |  | General &amp; Plastic Surgery | 510(k) | Sonopet Ultrasonic Surgical System Console, Ultrasonic Aspirator, Stryker Medtech. |  |  | June 24, 2014 | 2014 | Stryker Instruments Div. of Stryker Corporation | The firm received complaints regarding insufficient irrigation to the tip of the Sonopet system. Lack of irrigation can potentially lead to excessive heat at the tip. | Labeling design | The firm sent Medical Device Upgrade Notification letters dated July 11, 2014 to their customers. |  | 2460 | 2,460 units | Nationwide Distribution-including DC and the states of AK, AL, AR, AZ, CA, CO, CT, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, TN, TX, UT, VA, VT, WA, WI &amp; WV. | Terminated | October 27, 2015 | 490 | 2460 | Not_Computer
Z-0709-2014 | 66869 | No | CAE | Oropharyngeal airway. | Anesthesiology | Anesthesiology | 510(K) Exempt | King Systems |  |  | December 13, 0013 | 13 | King Systems Corp. | KLTD214 airways found inside KLTSD414 packages. The KLTSD414 airway has a lumen for use of a gastric suction tube; the KLTD214 airway does not include a lumen for a gastric suction tube. | Labeling Change Control | December 13, 2013 Urgent Recall Notices were sent to all Direct Accounts identifying the recalled item, risks associated with the incorrectly labeled product and asking customers to stop use, quarantine and destroy any recalled product on hand. Customers were asked to froward the notifcation if the product was further distributed and FAX the completed recall acknowledgement form to 317-776-5175 or email wnethercutt@kingsystems.com. Call the customer service representative (800.262.8673) to arrange for credit of your affected product(s). |  | 108 | 108 cases( 540 products) | Distributed in the states of CA, NY | Terminated | May 15, 2014 | 730638 | 108 | Not_Computer
Z-0487-2015 | 69648 | No | BZD | Noncontinuous ventilator (IPPB). | Anesthesiology | Anesthesiology | 510(k) | Cannulaide, Model 101. The Cannulaide product is composed of die-cut hydrocolloid coated polyurethane film with an integral hook strip (i.e., Velcro) to allow for the use of a loop material to help secure a nasal interface. |  |  | January 18, 2007 | 2007 | Beevers Manufacturing &amp; Supply, Inc. | Cannulaide model CA 101 was recalled because the firm discovered this model was packed and labeled as CA 102. | Labeling Change Control | Beevers Manufacturing and Supply, Inc. (BMS) of McMinnville, Oregon, began notifying consignees about this recall on 1/18/2007.Customers were advised to inspect blue box that say size 2 (CA 102) and check if size 1 (CA 101) was packed inside. Salter Labs provided BMS' recall report to FDA in October 2014. |  | 393 | 393 units | Worldwide Distribution: US (nationwide) and country of: ME-Musaffah. | Terminated | December 01, 2014 | 2874 | 393 | Not_Computer
Z-2592-2014 | 68993 | Yes | MSD | Blood access device and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Split Cath III Trays, 16F x 28CM |  |  | August 04, 2014 | 2014 | Medical Components, Inc dba MedComp | The Split Cath III Trays were incorrectly labeled as 16F x 32CM Split Cath III instead of 28CM. | Labeling False and Misleading | Medcomp sent a Recall Notification letter, dated August 4, 2014, to all affected customers. The letter identified the affected product, problem, and action to be taken. Customers were requested to respond and return all affected product. For questions call contact Susan Smith at ssmith@medcompnet.com. |  | 110 | 110 | USA Nationwide Distribution | Terminated | December 10, 2014 | 128 | 110 | Not_Computer
Z-2281-2014 | 68346 | No | HQL | Intraocular lens. | Ophthalmic | Ophthalmic | PMA | AMO TECNIS CL Foldable Silicone IOL with OptiEdge Design, Model Z9002. Indicated for primary implantation for the visual correction of aphakia in adult patients. Tecnis CL Mono 13 mm with the following diopter sizes: 19.5-20 |  |  | May 13, 2014 | 2014 | AMO Puerto Rico Manufacturing, Inc. | Units may be labeled with the incorrect diopter power. | Error in labeling | AMO sent an Urgent Medical Device Recall dated May 13, 2014 to the affected customers. The letter identified the affected product, problem and actions to be taken. This recall only affects the specific Silicone Tecnis IOLs Model Z9002. Implanted IOLs with appropriate post-op refractive outcomes are not affected by this recall. Please underatake the following actions:1. Stop using and remove from your inventory the affected product.2.Complete and return the included Facsimile Form as AMO. If you have inventory to return, please complete the Facsimile Form, and returned to AMO. For questions contact AMO Customer Service Representative 877-266-4543. |  | 77055 | 77,055 IOLs in total | Worldwide Distribution - Nationwide (USA and Puerto Rico), and foreign: Latin America, Asia Pacific, Canada, Europe, Africa and Middle East, and Japan. | Open, Classified |  |  | 77055 | Not_Computer
Z-1501-2014 | 67976 | [6 empty] | Hot Ice, 1 fl. oz., Meta Labs LLC, UPC 80705-2 |  |  | February 03, 2014 | 2014 | Altasource LLC dba Meta Labs LLC | Firm is not registered as a drug manufacturer. | No Marketing Application | Meta Labs notified customers of the Voluntary Recall Notification letter on February 4, 2014. The letter requested that the products be returned. For questions call 770-645-0602. |  | 155 | 155 | USA Nationwide in the states of AZ, AR, CA, CO, FL, GA, ID, IL,KY, LA, MI, NE, NV, NY, NC, OH, PA, TN, TX, UT, and VA. | Open, Classified |  |  | 274 | Not_Computer
Z-0934-2014 | 67302 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | 3.5mm Locking Cortical Bone Screw, T15 Hexalobe, Self Tapping, 20mm |  |  | January 08, 2014 | 2014 | OrthoPediatrics Corp | OrthoPediatrics has initiated a Voluntary RECALL of Part # 00-0903-2620, 3.5MM Locking CorticalScrew, Lot # 009SV because the device within the package may be incorrect. These packages incorrectly contain a 3.5mm Non Locking Cortical Screw # 00-1050-3524. | Labeling mix-ups | Orthopediatrics Corporation has requested sales representatives and distributors who have been identified via shipping records as having the possibility of having received one or more discrepant devices affected by this Recall were requested to quarantine the devices identified in this Recall by telephone on January 8, 2014. The decision to Recall these devices was made and a formal written Recall Notice sent to the consignees via email on January 13, 2014.Notification will not be sent to hospitals at this time provided that all devices areretrieved or accounted for within 60 days of the initiation of recall. Hospitals will be notified on an as needed basis if OrthoPediatrics determines that all recalled devices have not been recovered or accounted for, and that contacting the hospitals involved will expedite the recall efforts. All discrepant devices subject to this voluntary Recall are planned to be returned from distribution to OrthoPediatrics Corp. in Warsaw, Indiana and maintained in quarantine. All remaining unshipped devices from these lots have been placed in quarantine at OrthoPediatrics Corp.All quarantined devices at OrthoPediatrics will be inspected and repackaged to correct the mislabeling per normal documented procedures. Objective evidence will be maintained of the rework of the devices.For questions regarding this recall call 574-268-6379. |  | 12 | 12 units | Worldwide Distribution - USA including OK, KY, FL, LA, OH, CO, MO, MN and Internationally to Ireland, Saudi Arabia, Italy, Turkey, South Africa, and Germany. | Terminated | April 23, 2014 | 105 | 12 | Not_Computer
Z-1456-2014 | 67499 | No | GXL | Skullplate screwdriver. | Neurology | Neurology | 510(K) Exempt | Synthes CMF Battery Powered Driver, Part no. 305.840. Used to place screws in the cranialfacial bones, and to drill and/or tap holes to prepare cranialfacial bones for screw insertion. |  |  | January 30, 2014 | 2014 | Synthes, Inc. | The CMF Battery Powered Driver graphic case contains outlines of the Battery Powered Driver, taps, screwdriver blades, and a screw holding sleeve. However, the device is not intended to be used in conjunction with a tap for tapping holes in a patient's skull, upper and/or lower jaw. Therefore some outlines in the graphic case suggest a potential use of this device that is considered off-label. | Labeling design | Synthes sent an Urgent Notice: Field Safety Notification letter dated January 30, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Please take the following actions;Please see the Battery Powered Driver Users Manual for on-label procedures. You can access an electronic version of the Battery Powered Driver Users Manual by accessing the following : linlc http://www.synthes.com/pages/default.aspx. Click on North America Tab at the top of the screen, and then type Battery Powered Driver Users Manual in the search bar to download the Users Manual Forward this Field Safety Notification to anyone in your facility that needs to be informed. If the graphic case has been forwarded to another facility, contact that facility. " Maintain a copy of this notice.If you have any questions, please call 610-719-5450 or contact your Synthes Sales Consultant. Thank you for your attention to this issue. |  | 20081 | 20081 | Nationwide Distribution | Open, Classified |  |  | 20081 | Not_Computer
Z-2558-2014 | 68951 | No | KQT | Gastroenterology-urology evacuator. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | CollectEVAC, Sterile device with tubing and adaptor for use with ACMI and STORZ/Wolf resectoscopes. Part number 610-954. Rx only. Sterile EO. |  |  | August 01, 2014 | 2014 | Civco Medical Instruments Inc | A pinhole leak identified in the packaging of the recalled product potentially could compromise the sterility of the product. | Packaging | Civco Medical Instrucments contacted customers by phone and sent an Urgent Medical Device Recall letter dated August 1, 2014, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were requested to return all the affected product. Customers were instructed to contact Customer Service at 319-248-6520 with any concerns or questions. |  | 52 | 52 kits | Worldwide Distribution - USA including AZ, IL, MA, MD, MI, MN, MO, MS, LA, NC, NJ, NY, SD, TN, VA, WI, and Internationally to Brazil. | Terminated | January 28, 2015 | 180 | 52 | Not_Computer
Z-1494-2014 | 67831 | No | LJS | Percutaneous, implanted, long-term intravascular catheter. | General Hospital | General Hospital | 510(k) | Arrow Pressure Injectable PICC Powered by Arrow VPS Stylet Catalogue number CDC-35052-VPS, peripherally inserted central catheter. |  |  | March 24, 2014 | 2014 | Arrow International Inc | Kits are packaged with the incorrect catheter. The kits contain a 40cm catheter and should contain a 50cm catheter. | Labeling mix-ups | A recall notification letter, dated March 24, 2014 was sent to end users." Urgent Medical Device Recall Notification " |  | 789 | 789 | Nationwide Distribution- Including the states of Az, Tn, Fl, Ga, Ca | Terminated | September 21, 2015 | 546 | 789 | Not_Computer
Z-2302-2014 | 68536 | No | OEQ | Angiographic injector and syringe. | Cardiovascular | Cardiovascular | Enforcement Discretion | Customed Angiodrape Pack |  |  | May 20, 2014 | 2014 | Customed, Inc | Customed has determined that there is the possibility that packaging integrity may be compromised on different catalog numbers (multiple lots). This could result in an injury to the patient due to product contamination or loss of sterility condition. | Package design/selection | Customed Inc. sent a Recall Notification letter dated June 3, 2014 via email and followed by letter to all affected customers. The letter identified the affected products, problem and actions to be taken. The letter informed customers if the affected products are in their inventory to: 1. Remove products from sale (quarantine them) and cease distribution. 2. Complete and return the enclosed acknowledgement form to Customed Inc .Recall Coordinator. Affected products should be returned to Customed Inc for credit.Customers with questions were instructed to contact Customed Inc. Recall Coordinator at 1-787-801-0100 Ext. 7540 |  | 39936 | 26 lots, 1536 units (multiple units per lot) | Worldwide Distribution - US (nationwide) in the state of Florida including Puerto Rico (USA); &amp; USVI (St John) and the countries of Tortola, Trinidad, Dominican Republic, Costa Rica and Mexico. | Terminated | August 14, 2017 | 1182 | 15299586 | Not_Computer
Z-2678-2014 | 69102 | No | JAX | Pneumotachometer. | Anesthesiology | Anesthesiology | 510(K) Exempt | No packaging. The device is a computed tomography scanner. |  |  | January 09, 2007 | 2007 | Philips Medical Systems (Cleveland) Inc | When the Extended Display Firled ov View reconstruction is used, the CT numbers and anatomy in teh extended region are not accurate. | Labeling design | Philips sent a Product Safety Notification letter dated December 21, 2006 to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to distribute the Product Safety Notification to all affected users and keep a copy of the notification with the Instructions for Use until further notice. Customers with questions were instructed to contact Customer Care Center at 1-800-722-9377, option 5 - Diagnostic Imaging, option 1 - CT.For questions regarding this recall call 440-483-7600. |  | 61 | 61 | Worldwide Distribution - US including MO, TX, MD, AR, VA, KY, RI, NJ, MA, NY, LA, PA, VT, SC, CA, OH, FL, MN, MI and Internationally to Netherlands, Canada, Finland, Austria, France, Malaysia, Denmark, Norway, UK, Australia, Belgium, New Zealand. | Terminated | November 06, 2015 | 3223 | 61 | Not_Computer
Z-1457-2014 | 67560 | No | GAD | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | The Synthes Hohmann Retractor, 43 mm Long Shank, Angled 235 mm |  |  | February 11, 2014 | 2014 | Synthes, Inc. | The Synthes Hohmann Retractor was mis-etched on the product and package as part number 399.24 instead of part number 399.22. | Error in labeling | Synthes sent an Urgent Notice: Medical Device Recall letter dated February 11, 2014, to affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to Complete the attached Verification Section form. For questions call 610-719-5450. |  | 20 | 20 | Worldwide Distribution - USA Nationwide Distribution in the states of: TN, MD, NY, IN, NJ, WA, VA,and the country of Canada. | Terminated | September 10, 2015 | 576 | 20 | Not_Computer
Z-2670-2014 | 68989 | No | NGT | Intravascular catheter. | General Hospital | General Hospital | 510(k) | 0.9% Sodium Chloride Injection, USP BD PosiFlush(TM) SF Saline Flush Syringe 10 mL REF#306553 used in Centurion Convenience Kits. Recalled product is attached to the outside of the Convenience Kits. |  |  | July 25, 2014 | 2014 | Centurion Medical Products Corporation | Unit packages may exhibit open seals which impacts package integrity and potentially product sterility. | Packaging | Centurion sent an Urgent Product Recall Notice dated July 25, 2014, to all affected customers via Certified Mail. The notification discussed the affected product issue and provided instructions for initiating product return. Customers were instructed to identify and remove all inventory of these kits that they may have within their possession. Customers were asked to forward a copy of the notice to any customers to whom this product may have been further distributed. Customers were also asked to complete the enclosed accountability record, whether or not they have the product, and fax or email at their earliest convenience to 517-546-3356 or lcarpenter@centurionmp.com.Customers with questions were instructed to call (517) 546-5400, Ext 1135.For questions regarding this recall call 517-546-5400, ext 1156. |  | 21388 | 21388 | The product was distributed domestically to the following states: UT, WI, TX, NE, AZ, CA, IA, MD, NY, OK, PA, TN, TX, VA, NC, IL, MA, OR, WA, IN, SD. | Terminated | November 25, 2014 | 123 | 21388 | Not_Computer
Z-1491-2014 | 64791 | No | MVU | Analyte specific reagents. | Pathology | Hematology | 510(K) Exempt | The CEP 8 SpectrumGreen (SG) ASR Probe Kit, 20ul (list 06J37-018) is an Analyte Specific Reagent (ASR) containing a single vial of CEP 8 spectrum green probe and a Certificate of Analysis (COA). Chromosome enumeration probes (CEP) are chromosome-specific FISH probes that hybridize to highly repetitive human satellite DNA sequences, usually located near centromeres. CEP signals enable the identification and enumeration of human chromosomes in interphase and metaphase cells from fresh and archived samples. The label for CEP 8 SG ASR Probe Kit indicates it is an Analyte Specific Reagent, and analytical and performance characteristics are not established. |  |  | March 28, 2013 | 2013 | Abbott Molecular | 20 kits of CEP 8 SG ASR Probe (list 06J37-018, lot 440023), contained Vysis CEP 4 (Alpha Satellite) SG Probe, 20u, (part 32-112004 lot 438814), instead of the expected CEP 8 probe. | Packaging change control | Customers were sent Field Correction Notices dated 3/28/13 instructing them to inspect their CEP 8 SpectrumGreen Probe kits for the presence of CEP4 probe vials. If any CEP4 vials are discovered, customers should contact their local Abbott Molecular Technical Support Represtative for coordination of reagent replacement with the correct CEP8 vial. If customers have any specific questions they can contact Abbott Molecular at 1-800-553-7042, Option 2. |  | 20 | 20 kits | Nationwide (CA, CO, CT, FL, HI, IL, IN, MA, MD, MI, MN, MO, MT, NC, NE, NJ, NM, NV, NY, OH, OR, PA, TN, TX, VA, WA, WI); Canada | Terminated | March 08, 2016 | 1076 | 20 | Not_Computer
Z-2671-2014 | 68974 | No | MZO |  |  | Microbiology | PMA | VITROS Immunodiagnostic Products Anti-HCV Reagent Pack for in vitro qualitative detection of immunoglobulin G antibody to hepatitis C virus (anti-HCV) in human serum and plasma. REF 680 2454, IVD -- Ortho-Clinical Diagnostics High Wycombe UK. Note that one sales unit has 5 Reagent Packs contained within it. Each Reagent Pack can test for 100 samples (patient, Quality Control, etc.). |  |  | August 06, 2014 | 2014 | ORTHO-CLINICAL DIAGNOSTICS | VITROS Immunodiagnostic Products Anti-HCV kit lot 9090 has been observed producing sporadic lower than expected VITROS Anti-HCV test results due to an approximate signal loss of up to 66.2% which can result in unexpected negative test results being obtained for reactive sample fluids (quality control and patient samples). | Packaging process control | All US consignees (users) who were shipped the affected lot VITROS HCV Reagent Pack were sent communications via FedEx and/or US Priority Mail (PO Boxes only) on 8/06/14, and advised of the issue and to immediately discontinue using all remaining inventory of VITROS Anti-HCV Reagent Packs Lot 9090. For additional information, please contact OCD's Customer Technical Service representatives at 1-800-421-3311. |  | 463 | 463 sales units | NY, NJ, CT, MA, PA, VA, MD, FL and GA | Terminated | April 14, 2017 | 982 | 463 | Not_Computer
Z-2568-2014 | 69072 | No | DSP | Intra-aortic balloon and control system. | Cardiovascular | Cardiovascular | 510(k) | LINEAR 7.5 FR. 40cc IAB Catheter |  |  | August 11, 2014 | 2014 | Datascope Corporation | During label reconciliation a duplicate serial number label was detected to be printed for the LINEAR 7.5 FR 40 cc Intra-Aortic Balloon Catheter. | Labeling False and Misleading | Maquet Inc.is sued an Urgent Medical Device Recall Letter/ Fax Back Form dated August 11, 2014. Customers are asked to quarantine the affected product and upon return they will be provided with an exchanged unaffected product. The attached form should be completed an returned to Maquet Inc. Questions can be directed to Maquet Technical Support at (888) 627-8383 - Press option "3" followed by option "1" Monday through Friday between the hours of 8 am and 5 pm (EST). |  | 1 | one unit | US Distribution to TX. | Terminated | February 05, 2015 | 178 | 1 | Not_Computer
Z-1489-2014 | 68034 | No | DQY | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | ev3 EverCross 0.035" OTW PTA Dilatation Catheter, 6mm x 120mm x 135cm, REF: AB35W06120135. and ev3 EverCross 0.035" OTW PTA Dilatation Catheter, 6mm x 40mm x, 80cm REF: AB35W06040080. Peripheral Dilatation Catheter |  |  | April 09, 2014 | 2014 | Ev3, Inc. | Covidien is recalling 2 production lots of EverCross 0.035 PTA Balloon Catheter because of a product labeling error. There is a discrepancy in the PTA Balloon length and the delivery catheter length printed on the product labeling. | Labeling Change Control | Consignees were sent a Covidien "Urgent Product Recall" letter dated April 09, 2014. The letter described the problem and the product involved in the recall. The letter described the required actions which included to immediately quarantine and discontinue use of the affected devices and to return affected product. Requested consignees to complete the Verification Form and return it to their Covidien Sales Representative. For questions they can contact Covidien Customer Service at 1-800-716-6000. |  | 39 | 39 units (31 OUS, 8 US) | Worldwide Distribution-USA including the states of WV, TX, OK, IN, NY, TX, and NC, and the countries of AUSTRALIA, ISRAEL, TURKEY, UNITED KINGDOM, ITALY, FRANCE, and CHINA. | Terminated | October 06, 2014 | 180 | 39 | Not_Computer
Z-2591-2014 | 69022 | No | MRD |  |  | General &amp; Plastic Surgery | 510(k) | EN MENTOR Round High Profile Single Use Saline Breast Implant Sizer Single Use Saline Breast Implant Sizer |  |  | July 02, 2014 | 2014 | Mentor Texas, LP. | The device labeling is stamped with 330cc and the device is actually 380cc device. | Error in labeling | Mentor sent an Important Customer Notification letter dated July 17, 2014 to all affected customers and follow-up with a visit by the local sales representative. The letter identified the affected product, problem and actions to be taken. Customers were instructed to immediately identify all unsused product from lot list provided and maintain a copy of this notice with the affected product. For question contact your local Sales Representative or Mentor Customer Service at 1-800-235-5731. |  | 34 | 34 units | U.S Nationwide Distribution in the states of Michigan, Maryland, Texas, New Jersey, Indiana, Nevada, Florida, California, New York, Tennessee, Pennsylvania, and Colorado. | Terminated | April 06, 2015 | 278 | 34 | Not_Computer
Z-1441-2014 | 67752 | No | DSA | Patient transducer and electrode cable (including connector). | Cardiovascular | Cardiovascular | 510(k) | FAS-LOC 6 Foot Disposable Extension Cable with Safe Connect/Screw Down. |  |  | March 05, 2014 | 2014 | Remington Medical Inc. | Remington Medical discovered an error on the FL-601-97 IFU; part number IFU-RM-0154. Specifically, the polarity for the black chuck in the picture is wrong. It shows (+), but should show (-). | Labeling Change Control | Remington Medical Inc sent an Field Safety Notice dated March 5, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Please be aware that the typographical error described above is included in the FL-601-97 Instructions for Use (RM-0154 Rev 05) provided with the lots that you have received as identified below. RMI has corrected the typographical error in the revised FL-601-97 Instructions for Use (RM-0154 Rev 5.1)which is attached to this communication. Please incorporate the information contained in the attachment with the existing FL-601-97 labeling (Instructions for Use) provided with the product. If you have distributed these products further, please forward copies of this notice and the revised Instructions for Use to anyrecipients as appropriate.If you have any questions regarding this communication, please contact RMIs sales representatives or customer service team at 800-989-0057, Monday - Friday, between the hours of 8 am - 4pm EST.Patient safety is the highest priority for Remington Medical, Inc. We appreciate your assistance with this matter and regret any inconvenience this may cause. |  | 41499 | 41,499 cables | USA (nationwide) Distribution including Canada | Open, Classified |  |  | 41499 | Not_Computer
Z-0920-2014 | 66953 | Yes | KDI | High permeability hemodialysis system. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Fresenius Bicarbonate Jug Adapter to the 2008 Series Hemodialysis machine |  |  | December 17, 2013 | 2013 | Fresenius Medical Care Holdings, Inc. | Bicarbonate Jug Adapter distributed prior to510(k) approval | No Marketing Application | Fresenius notified customers by telephone and Formal notification on 12/17/13 via the Urgent Bicarbonate Jug Adapter Recall letter sent certified mail with signature confirmation and fax-back form.Customers were instructed to examine their stock immediately to determine whether they have any of the affected parts on hand. If customers have the affected parts, they are instructed to discontinue use and place all Bicarbonate Jug Adapters in a secure area for return to Fresenius Medical Care Renal Therapies Group, LLC (FMC-RTG). Customers instructed to contact their Fresenius Medical Care Technical Services Team for instructions on how to return the recalled product. |  | 703 | 703 units | Nationwide Distribution. | Open, Classified |  |  | 703 | Not_Computer
Z-2661-2014 | 69182 | No | KOD | Urological catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Rusch, Soft Rubber (Nelaton and Nelaton Robinson), Bladder Catheter, Sterile. |  |  | September 04, 2014 | 2014 | Teleflex Medical | Misbranding: Although the product labeling identifies the catheters as &quot;soft rubber,&quot; the Natural Rubber Latex caution statement was not printed on the product label. | Labeling Change Control | Consignees were notified via letter on/about 09/04/2014. Subsequently, consignees were notified of the expanded recall on 01/13/2015. |  | 193363 | 193,363 ea. + 1,140 units (total 194,503 units) | Worldwide Distribution-USA (nationwide) and the countries of Austria, Belgium, Bahrain, Belgrad, Canada, Croatia, Cypress, Czeck Republic, Denmark, Finland, Germany, Great Britain, Hungary, Ireland, Italy, Libya, Netherlands, Norway, Romania, Russia, Saudi Arabia, Slavonia, Slovakia, Spain, Switzerland, Turkey, Portugal and Zambia. | Open, Classified |  |  | 193363 | Not_Computer
Z-2570-2014 | 68754 | No | LKF | Obstetric-gynecologic specialized manual instrument. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | CooperSurgical H/S CATHETER (PART # 61-5007) |  |  | July 03, 2014 | 2014 | CooperSurgical, Inc. | Unsealed pouches. | Packaging process control | Customer notification letter was sent on July 1, 2014 via UPS with confirmed delivery receipt. The letter identified the product, the problem, and the action needed to be taken by the customer. At CooperSurgical's expense, arrangements were made to replace any of the affected products that customers may have on hand. For questions contact James Keller, Vice President, Regulatory Affairs/Quality Assurance at 203-601-5200. |  | 1400 | 1,400 units | Worldwide Distribution -- USA, in the states of CA, TX, IL, MI, MS, ID, NY, FL, MD, PA, OR, GA, CO, WA, NJ, HI, and NC, including DC; and the countries of the countries of Spain, England, Holland, and United Kingdom. | Terminated | May 28, 2015 | 329 | 1400 | Not_Computer
Z-2595-2014 | 69066 | No | LKB |  |  | General Hospital | Enforcement Discretion | DualCap IV Pole Strips Disinfectant Caps for Luer Access Valves, REF 500-LB, Sterile. |  |  | August 05, 2014 | 2014 | Catheter Connections, Inc. | Catheter Connections is recalling DualCap IV Pole Strips disinfectant caps for luer access valves due to potential improper packaging cap seal which may render the sealed cap non-sterile. | Packaging | The firm, Catheter Connections, sent an "Urgent: Medical Device Field Action" letter dated August 5, 2014 to its consignees/customers. The letter described the product, problem and actions to be taken. The consignees/customers were instructed to identify any customers who received this product having lot no. 1406509; make all potential users in your facility are made aware of this notification and the recommended actions; forward the attached URGENT MEDICAL DEVICE CORRECTION LETTER to customers immediately via fax or email; search your inventory and instruct all users to check each cap coming from the specific affected lot; quarantine and discard any defective product; and complete and return the attached URGENT MEDICAL DEVICE FIELD ACTION REPLY FORM via fax to 1-888-862-2693 or email to: info@cathconn.com.Catheter Connections does not require the return of devices having the affected lot number.If you have questions, please contact Catheter Connections at 1-801-906-0820 (M-F, 8:00 am - 5:00 pm MST) and ask for Quality Department. |  | 24910 | 24,910 units | US Distribution to states of: CT, GA, MD, and PA.. | Terminated | October 08, 2014 | 64 | 24910 | Not_Computer
Z-2668-2014 | 68874 | No | BRZ | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Pedi-Pak Pedi-Syringe Filter 60 mL Becton Dickinson/Monoject Syringe |  |  | July 17, 2014 | 2014 | Genesis BPS, LLC. | Genesis BPS is recalling numerous devices since they did not notify the agency of its intent to introduce material changes. | No Marketing Application | Genesis BPS sent an Urgent Medical Device Recall letter with a Recall Letter/Confirmation Form on July 17, 2014 to all affected customers. Customers were asked to immediately examine their inventory and quarantine product subject to recall. In addition, if there is further distribution of the product, customers should be identified and be notified of the recall. Product use should be discontinued and disposed of according to applicable federal regulations and internal protocols.The attached confirmation form should be completed and returned to Genesis by either fax or (201) 708-1104 or e-mail (info@genesisbps.com). Questions can be directed to (201) 708-1400 ext 2004 (9:00 and to 5:00 pm EST).For questions regarding this recall call 201-488-1174. |  | 7272 | 7272 units | Nationwide Distribution &amp; Switzerland | Terminated | March 04, 2016 | 596 | 19440 | Not_Computer
Z-2645-2014 | 69117 | No | OLI | Low level laser system for aesthetic use | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Ultraslim. These include all products labeled as UltraSlim, Luxurian, BioSlim Lipo, Cellulize, Slim 'n Trim, UltraRed, UltraBlue, Slim-PRO, Dazzle Light, UltraSlim Cold Light, SlenderRay, and U-Slim It, as well as those products with private labelling branding. |  |  | May 06, 2014 | 2014 | Advanced Photonic Systems, LLC | Marketing without marketing clearance or approval. | No Marketing Application | Advanced Photonic Systems sent an Urgent Medical Device Recall letter dated July 12, 2014 via UPS to all those who are known by APS to have possession, custody, or control of the affected device. Customers were instructed to return the recalled product at the sole expense of APS. Each 90 days for the following year, APS plans to provide a progress report to the FDA that identifies which recipients have complied. For more information , customers may write to Advanced Photonic Systems, LLC at the address provided or call 1-800-210-5060 (9am to 5pm, Monday through Friday).For questions regarding this recall call 800-210-5060. |  | 235 | 235 Total | Worldwide Distribution - US- AL, AZ, CA, CO, FL, GA, ID, IL, IN, LA, MA, MD, MI, MN, MO, NV, NJ, NY, NC, OH, PA, TX, UT, and VA.Internationally to Canada, France, and Switzerland | Terminated | June 24, 2015 | 414 | 235 | Not_Computer
Z-1472-2014 | 67803 | No | HSB | Intramedullary fixation rod. | Orthopedic | Orthopedic | 510(k) | Fassier-Duval IM Telescopic System. Orthopedic intramedullary rod implant. Catalogue number FD-040(L)-SS. |  |  | March 05, 2014 | 2014 | Pega Medical Inc. | The outer package label contained the incorrect length of the device. There is a label on the outer package and an additional label on the product contained in the package. | Error in labeling | There is an error on the outside label of the implant. The length was mistaken on the label. It was indicated 245mm instead of 338mm. It has to be clarified that the Fassier-Duval implant is provided with two labels, one that is outside the tube used for packaging and one inside the tube. |  | 48 | 48 | Worldwide distribution, USA nationwide, Greece, United Arab Emirates, Peru, Canada, Israel, United Kingdom, France, Germany, Colombia, Mexico, Singapore, Lithuania, Denmark, and Spain. | Terminated | May 19, 2014 | 75 | 48 | Not_Computer
Z-2625-2014 | 69043 | No | LGM | Patient care reverse isolation chamber. | General Hospital | General Hospital | 510(k) | Life Force Chamber |  |  | June 27, 2014 | 2014 | Life Force of Tampa, LLC | The Life Force chamber is marketed and promoted as a medical device, intended for use in the diagnosis of disease or other conditions or in the cure, mitigation, treatment, or prevention of disease, or is intended to affect the structure or function of the body. Life Force of Tampa does not have an approved application for premarket approval in effect, or an approved application for an investigat | Labeling False and Misleading | Life Force of Tampa LLC sent an URGENT: MEDICAL DEVICE RECALL letter dated June 27, 2014 to all affected customers. The letter instructed the customers to: 1) Immediately remove and discard all copies of the labeling material. 2) Remove all electronic copies of the materials from any publicly accessible locations (e.g. website). 3) Complete and return the Acknowledgment form to Life Force of Tampa. For questions or concerns contact Douglas G. Cornell at (813) 961-4655. |  | 10 | 10 | Nationwide Distribution in the states of AZ, FL, MD, NV, OK, PA, and TX. | Terminated | July 30, 2015 | 398 | 10 | Not_Computer
Z-2655-2014 | 69054 | No | MEB | Infusion pump. | General Hospital | General Hospital | 510(k) | B.Braun Accuflo 100ML/H 100ML Elastomeric Infusion Device. |  |  | August 06, 2014 | 2014 | B. Braun Medical, Inc. | The primary packaging label (inner pouch label) indicates that the package contents are Catalog Item: CT-1000-250, while the secondary packaging label (outer carton label) indicates the contents are CT-1000-100. The actual contents are a CT-1000-100. | Labeling Change Control | B. Braun Medical Inc. provided the recall notification received from B. Braun Melsungen AG to Progressive Medical Inc. who is the initial importer of the product in the U.S. market. The notification was sent via FedEx and e-mail on/about August 6, 2014 and identified the product, problem, and action to be taken. The notification indicated that if the affected units were further distributed that any distributors are required to forward the letter on to their consignees. |  | 8640 | 8,640 units | Distributed in the state of Missouri. | Terminated | January 29, 2015 | 176 | 8640 | Not_Computer
Z-1463-2014 | 67835 | No | ONG | Laser surgical instrument for use in general and plastic surgery and in dermatology. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | FRAXEL DUAL 1550/1927 Laser System, (FRAXEL SR 1500 DUAL Laser System), for dermatological procedures. |  |  | March 20, 2014 | 2014 | Solta Medical, Inc. | Obsolete versions of the Laser Treatment Settings card for the Fraxel DUAL 1550/1927 Laser system may have contained incorrect indication for use for treatment of Melasma. | Labeling Change Control | Solta Medical sent an Correction Important Product letter dated March 20, 2014 to affected customer. The letter identified the affected product, problem and actions to be taken. Customers were instructed to complete and return the enclosed response form. |  | 1077 | 1077 units | Worldwide distribution | Terminated | April 27, 2015 | 403 | 1077 | Not_Computer
Z-0901-2014 | 67206 | No | BWN | Manual operating table and accessories and manual operating chair and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Radial Assist RAD BOARD, used to support the |  |  | December 18, 2013 | 2013 | Merit Medical Systems, Inc. | Merit Medical Systems, Inc. is voluntarily recalling one lot (B507171) of RAD BOARD RB 100 devices. The affected devices are missing the main label which graphically depicts patient orientation on the board. The issue may allow the board to be oriented in the wrong direction where the semi-radiopaque material in the board (embedded Xenolite) could obscure patient anatomy during fluoroscopy. There | Labeling Change Control | Merit Medical Systems, Inc.sent a letter voluntarily recalling one lot (B507171) of RAD BOARD" RB100 devices. The affected devices are missing the main label which graphically depicts patient orientation on the board. The issue may allow the board to be oriented in the wrong direction where the semi-radiopaque material in the board (embedded Xenolite) could obscure patient anatomy during fluoroscopy. There have been no reports of patient harm or injury from Merit customers as a result of this issue. Merit has chosento remove this lot from the field by replacing the unlabeled RAD BOARDs. Merit requests that customers stop using these devices immediately until the product has been replaced by their Merit sales representative.Customers with questions were instructed to call 1-801-208-4344 or 1-801-208-4365.For questions regarding this recall call 801-208-4301. |  | 11 | 11 | Worldwide Distribution - USA including NJ, TX, WV, MN, GA, and Internationally to Mexico and Japan. | Terminated | February 13, 2014 | 57 | 11 | Not_Computer
Z-1465-2014 | 67925 | No | OEZ | Diagnostic intravascular catheter. | Cardiovascular | Cardiovascular | Enforcement Discretion | Terumo Cardiovascular Procedure Kit- X-Coated Pump and Table Pack with FX25RE |  |  | March 28, 2014 | 2014 | Terumo Cardiovascular Systems Corp | Incorrect customer notification label applied to the Tyvek lid of this specific lot of TerumoÂż Cardiovascular Procedure Kits | Labeling mix-ups | Terumo CVS contacted the affected user on 3/28/14 by phone via phone script to discontinue use and return all affected units in inventory. Terumo CVS will replace or issue credit for returned product. |  | 37 | 37 units | CA | Terminated | December 12, 2014 | 259 | 37 | Not_Computer
Z-2640-2014 | 69083 | No | MGQ |  |  | General &amp; Plastic Surgery | 510(k) | Amerigel Post-Op Surgical Kits . Kit includes: Amerigel Wound Dressing, Amerigel Wound Wash, 22x2 Gauze Pads, and 1x3 Fabric Bandages |  |  | July 11, 2014 | 2014 | Amerx Health Care Corp. | Not approved labeling claims (antimicrobial and autolytic debridement) | Labeling design | On July 21, 2014, Amerx Health Care Corporation telephoned all affected customers advising them of the recalled products. During the call the customers were instructed to: 1) Discontinue dispensing the lot(s) and promptly return via parcel post to the Clearwater facility. 2) If you dispose of the product on your own, provide proof of disposal. 3) You will be reimbursed by check or credit memo for returned or disposed goods and postage. |  | 116633 | 116,633 units | Worldwide Distribution-USA (nationwide) including the District of Columbia, Virgin Islands, Puerto Rico, and Guam, and the states of AK, AL, AZ, CA, CO, CT, FL, GA, HI, IL, IN, KS, KY, LA, MA, ME, MD, MI, MT, MO, NC, NH, NE, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, WA, WI, WY, and WV, and the countries of Canada, England, Africa, and Australia. | Terminated | July 22, 2015 | 376 | 522101 | Not_Computer
Z-2636-2014 | 69049 | No | DQO | Diagnostic intravascular catheter. | Cardiovascular | Cardiovascular | 510(k) | Volcano TrakBack II Disposable Pullback Device; |  |  | August 13, 2014 | 2014 | Volcano Corporation | Internal testing had identified a potential breach to the sterile barrier that may impact the sterility of the product. | Packaging | Volcano sent an Urgent Product Recall Notification letter dated August 13, 2014, to all customers who received the device in the last 25 months. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to immediately cease use of the affected product and return unused devices to the Volcano distribution center. Customers were asked to prepare the attached inventory log and then contact Volcano Customer Service at 800-228-4728, option 4 to arrange for return of any affected product in their inventory. A credit would be issued for any returned product. Customers with questions were asked to call 800-228-4728.For questions regarding this recall call 978-439-3586. |  | 8251 | 8251 total devices (3777 US) | Worldwide Distribution - USA (nationwide) and Internationally to Korea, Thailand, Canada, Australia, Chili, Mexico, Argentina. | Terminated | June 30, 2015 | 321 | 8251 | Not_Computer
Z-2673-2014 | 69047 | No | NGT | Intravascular catheter. | General Hospital | General Hospital | 510(k) | BD PosiFlush SF Saline Flush Syringe 0.9 Sodium Chloride Injection, USP 10 mL REF 306553. |  |  | July 22, 2014 | 2014 | Becton Dickinson &amp; Company | BD received some reports of open seals found on the BD Posiflush SF Flush Syringe 10 mL (Ref #306553). | Packaging | BD sent recall letters/recall response cards dated 7/22/2014 via e-mail or UPS 2nd day mail. Customers are asked to examine their inventory and quarantine any affected product subject to recall. Customers should complete the Recall Response Card and fax it back to BD at 1-201-847-6990. Any questions can be directed to 1-888-237-2762 option 3. |  | 3088320 | 3,088,320 units | Nationwide Distribution. | Terminated | April 21, 2015 | 273 | 3088320 | Not_Computer
Z-2550-2014 | 68818 | No | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Trauma Tubes and Plates: |  |  | July 22, 2014 | 2014 | Zimmer, Inc. | Zimmer is recalling sterile-packaged implants and instruments packaged in six different package configurations due to packaging design verification test failures. Specifically, multiple test samples fromeach of the six configurations failed simulated distribution and shipping testing when the devices either breached inner or outer package trays or seal integrity was lost. | Package design/selection | Zimmer sent an URGENT MEDICAL DEVICE RECALL letter dated July 2014 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Please call the Customer Call Center at 1-800-348-2759 for questions or concerns involving this notice. |  | 972125 | 972,125 units | Worldwide Distribution. | Terminated | September 15, 2015 | 420 | 972125 | Not_Computer
Z-2290-2014 | 68753 | [6 empty] | PENTAX Upper/Lower G.I. Endoscope OF-B130 and OF-B194 Gas/Water Feeding Valve |  |  | July 17, 2014 | 2014 | Pentax Medical Company | The Instructions for Use for the Pentax OF-B130 and OF-B194 Gas/Water Valve IFU if not followed carefully may cause a potential risk to health. | Labeling design | Pentax Medical sent an Urgent Medical Device Recall Field Safety Corrective Action letter dated July 16, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.If the Instruction for Use are not followed carefully by the user there is a potential risk to health for the patient, such as a perforation.Customers are asked to complete the attached response form and fax to (201) 799-4063. Customers should see the attached revised Instructions for Use Z632 R09. If you have any questions regarding this action, please feel free to contact your Sales Representative at : 800-431-5880 ext 2064 |  | 1206 | 98 units US; 1108 units OUS | Worldwide Distribution - US Distribution including the states of CA, CO, IN, KS and ID., and the countries of Asia, Canada, Europe and Japan. | Terminated | August 20, 2015 | 399 | 1206 | Not_Computer
Z-2689-2014 | 68955 | No | BZE | Breathing system heater. | Anesthesiology | Anesthesiology | 510(k) | CareFusion AirLife Heated Infant Breathing Circuit |  |  | July 31, 2014 | 2014 | Carefusion 2200 Inc | The CareFusion AirLife Heated Infant Breathing Circuit is being recalled due to a regulatory compliance risk involving a material change. The changes to the gas pathway may potentially impact the safety of the device and quality of the gas condensates to the patient. | No Marketing Application | CareFusion sent an URGENT: PRODUCT RECALL letters dated July 31, 2014 to all affected customers. The letters instructed customers to: 1) perform a 100% physical inventory to verify if any of the affected product codes are in stock; 2) complete and return the attached Recall Response Form within 15 days; 3) notify any customers/parties that the recalled products were distributed to of the recall; 4) distributors should have all their customers return the products and a completed Recall Response Form to the distributor for credit; and, 5) customers should destroy the recalled products according to their disposal procedure. If a customers procedure(s) does not allow for the destruction of the recalled products, the customers was advised to contact CareFusion AirLife Customer Support at 800-323-9088 (Monday - Friday, 8:00 AM - 5:00 PM CST) for product return options. Customers will receive credit for all destroyed products. Direct accounts with questions about this recall should contact CareFusion AirLife Customer Support at 800-323-9088 (Monday - Friday, 8:00 AM - 5:00 PM CST) for additional information. |  | 318280 | Approximately 318,280 devices | Worldwide Distribution - US Nationwide in the states of AZ, CA, FL, IL, KY, LA, MI, MO, NE, NY, OH, OK, PA, RI, TN, TX, UT including Puerto Rico and the countries of Canada, China, Kuwait, Mexico, Saudi Arabia, and Switzerland. | Terminated | February 08, 2016 | 557 | 318280 | Not_Computer
Z-2171-2014 | 68814 | No | NAY | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | EndoWrist Stapler System User Manual Addendum and Wall Chart; |  |  | June 02, 2014 | 2014 | Intuitive Surgical, Inc. | Changes were made to the Stapler EndoWrist Stapler 45 System Instruments and Accessories User Manual Addendum (550986-03) and the EndoWrist Stapler 45 Wall Chart (PN 551524-01) in which the pH range specification for the enzymatic cleaner used throughout the cleaning process was modified from pH-neutral to pH-neutral to mild alkaline (mild alkaline is defined as pHd11). Mild Alkaline enzymatic | Labeling Change Control | Intuitive Surgical representatives hand delivered corrected manuals to the two affected sites.. |  | 2 | 2 | Distribution to Florida and South Carolina | Terminated | August 08, 2014 | 67 | 2 | Not_Computer
Z-1525-2014 | 66841 | No | MHY |  |  | Neurology | PMA | Medtronic Activa RC (37612), Activa PC (37601), Activa SC (37602), Activa SC (37603), Activa PC+S (37604), Multi-program Rechargeable Neurostimulator for Deep Brain Stimulation for movement disorders. |  |  | February 27, 2014 | 2014 | Medtronic Neuromodulation | Medtronic neurostimulators (used with pain and deep brain stimulation therapy) have a cycling feature that, when enabled, allows stimulation output to be toggled on and off based on the programmed settings. The use of this feature is determined by the physician and set using the clinician programmer. Current labeling indicates the use of cycling improves device longevity and recharge interval | Labeling False and Misleading | Medtronic sent a Medical Device Correction letter dated February 2014 to all affected consignees. The letter identified the product, the problem, and the action to be taken by the consignee. Consignees were instructed to do the following:If you have patients with cycling enabled or wish to use cycling, use the updated calculations and graphs available online at professional.medtronic.com/cycling to assess the impact of cycling on device longevity. Medtronic will update the product labeling in the second half of 2014." Remind your patients to continue to check their battery status. Instructions on how to check battery status can be found in the patient programmer or recharger system manual." Pain Stimulation Trialing Only: The cycling feature should not be enabled on the External Neurostimulator (ENS) when estimating device longevity because the resulting longevity estimate for non-rechargeable devices may not provide accurate information.Consignees were also instructed to acknowledge receipt of the notice by completing the attached reply form and returning it to Medtronic using the contact details on the reply form. For questions consignees can contact Medtronic Neuromodulation Technical services at 1-800-707-0933 weekdays 7am - 6pm CST.For questions regarding this recall call 800-707-0933. |  | 244207 | 170,286 US, 73,921 OUS for all devices | Nationwide Distribution - all states including DC and Puerto Rico. | Terminated | November 07, 2014 | 253 | 244207 | Not_Computer
Z-2734-2014 | 69053 | Yes | CBF | Hyperbaric chamber. | Anesthesiology | Anesthesiology | 510(k) | The Rejuvenator Chamber manufactured by Baro-Therapies, Inc. |  |  | August 08, 2014 | 2014 | Baro-Therapies, Inc | The firm is distributing the Rejuvenator device without an approved 510(k). | No Marketing Application | Baro-Therapies issued a Correction Notice letter on August 11, 2014, to all affected customers. The letter stated they have suspended marketing of the device and plans to shut down their website until they receive clearance from the FDA. For questions or concerns contact the firm at (850) 654-0051. |  | 10 | 10 | US Nationwide Distribution in the states of FL, GA, MS, TN and CO. | Terminated | August 06, 2015 | 363 | 10 | Not_Computer
Z-2176-2014 | 68798 | No | IZL | Mobile x-ray system. | Radiology | Radiology | 510(k) | Agfa Digital Radiography X-Ray System DX-D 100 |  |  | May 13, 2014 | 2014 | AGFA Healthcare Corp. | Unit was mislabeled with a factory label showing 32 KW output power. | Labeling mix-ups | Only one customer was affected no customer letter was issued field correction was conducted.. |  | 1 | 1 | US in the state of WV | Terminated | August 13, 2014 | 92 | 1 | Not_Computer
Z-2181-2014 | 68803 | No | PDH | Neonatal phototherapy unit. | General Hospital | General Hospital | 510(k) | Luma Wrap Phototherapy Swaddler is designed to provide centered and comfortable boundaries as a benefit for restive babies in the Neonatal Intensive Care Unit. |  |  | July 11, 2014 | 2014 | Beevers Manufacturing &amp; Supply, Inc. | Beevers Manufacturing &amp; Supply Inc. is recalling Luma Wrap neonatal phototherapy blankets because the product was inadvertently misclassified and marketed without a cleared 510(k) submission from the US FDA. | No Marketing Application | BMS is contacting their customers and requesting that distributors report the amount of affected product they have in stock, prior to returning affected product. BMS is also asking its distributors to forward the recall notification to its customers to whom affected product was sold. BMS expects those customers to also report to BMS the amount of affected product they have. |  | 22080 | 240 total (112/ lot 66812; 92/ lot 66813; 17/ lot 30660; and 19/ lot 67879). | Product was distributed to US, Germany, Austria, France, Italy, Spain, Australia. | Terminated | July 17, 2015 | 371 | 22080 | Not_Computer
Z-2729-2014 | 68725 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Reusable laparoscopic instruments are packaged as non-sterile. Cleaning and sterilization of devices must occur prior to use. |  |  | May 30, 2014 | 2014 | Instrumed International, Inc. | Additional caution statements have been added to the Instructions For Use for monopolar electrodes. 1) To avoid tissue carbonation, the operation voltage of the HF generator must not exceed 650 peak voltage (Vp) for all monopolar electrodes. and 2) For all monopolar electrodes, the useful life of this product is less than or equal to 50 cycles and less than or equal to 2 years. | Labeling design | Instrumed sent an URGENT: MEDICAL DEVICE CORRECTION letters dated May 30, 2014 to all consignees. The letters advised the consignees of the new warning statements on the product Instructions For Use and an instruction to complete and return the attached Response Request Form to Instrumed International. Consignees were further instructed to extend the recall to the retail level. For questions contact Instrumed Quality Manager at 847-908-6119. |  | 15 | 15 instruments | Distribution US Nationwide in the states of: AL, AR, AZ, CA, CO, CT, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, WA, WI, WV, and WY. | Open, Classified |  |  | 4504 | Not_Computer
Z-0622-2014 | 67025 | No | JXG | Central nervous system fluid shunt and components. | Neurology | Neurology | 510(k) | Integra Large Rickham -Style Reservoir |  |  | November 13, 2013 | 2013 | Integra LifeSciences Corp. | Integra became aware through an adverse complaint trend that Rickham Style Reservoirs with bent or broken connectors that were shipped in sleeve-type packages were found prior to use but after the sterile package was open. | Packaging | Integra sent an Urgent Voluntary Medical Device Recall letter dated November 13, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to inspect their inventory for the affected product and if located stop using immediately. Customers were asked to complete the attached form and return it to Integra as indicated and keep a copy of the form for their records. Once the form was received by Integra they will send an RMA number, directions to return the product, and input an order to replace the quantity they indicated on the form.Customers with questions were instructed to call 1-855-532-1723.For questions regarding this recall call 855-532-1723. |  | 78 | 78 packages | Worldwide Distribution - USA including FL, NE, WI, WA, CT, ND, AL, NY, NH, OR, OK, MN, and TX. Internationally to Belgium, Czech Republic, Germany, France, Great Britian, Italy, Poland, and Portugal. | Terminated | February 13, 2015 | 457 | 272 | Not_Computer
Z-0627-2014 | 67093 | No | ODP | Intervertebral body fusion device. | Orthopedic | Orthopedic | 510(k) | C-Plus" 10mm x 12mm Webless implants |  |  | October 07, 2013 | 2013 | RTI Surgical, Inc. | As the result of a recent internal review of regulatory documents, we have determined that the 10 mm x 12 mm Webless C-Plus implants were incorrectly documented as meeting the requirement for labeling as a partial Vertebral Body Replacement Device (VBR). While the use of this device as a partial VBR is unlikely due to the size of the device, use of this device as a partial VBR could result in a s | Labeling False and Misleading | RTI Surgical, Inc. sent an Important Notice letter dated October 8, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed remove fall affected product from their inventory and send back to Customers Service at 375 River Park Circle, Marquette, MI 49855 by October 16, 2013. Customers with questions were instructed to call 906-226-4489.For questions regarding this recall call 906-226-4812. |  | 1906 | 1906 | Nationwide Distribution - USA including FL, MO, NY, CA, MI,and IL. | Terminated | February 26, 2014 | 142 | 1906 | Not_Computer
Z-0071-2015 | 69333 | No | DKZ | Amphetamine test system. | Toxicology | Toxicology | 510(k) | ArcPoint Labs 5 Panel Dip Drug Screen, Part No. APD-5M. |  |  | September 23, 2014 | 2014 | Ameditech Inc | Ameditech is recalling the ArcPoint Labs Dip Drug Screen Dip Card because the positive interpretation illustration was incorrect on the cap labels as well as on the card cap graphic in the package insert. | Labeling Change Control | A customer notification letter dated 9/23/14 was sent to all customers who purchased the ArcPoint Labs Dip Drug Screen Dip Card. The letter informs the customers of the problems identified and the actions to be taken. Customers are instructed to complete and return the Verification Form. Customers with questions are instructed to contact Ameditech at (858) 805-8714 or email at amd.complaint@alere.com. |  | 400 | 400 kits | Distributed in AL. | Terminated | February 27, 2015 | 157 | 800 | Not_Computer
Z-0986-2014 | 67390 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Synthes 3.2 mm Guide Wire 400mm. |  |  | January 27, 2014 | 2014 | Synthes, Inc. | Synthes 3.2 mm Guide Wire 400mm included an incorrect raw material listed on the label. | Labeling Change Control | A recall notification letter, dated January 27, 2013, was sent to End Users and sales consultants. |  | 1132096 | 1,132,096 | Worldwide Distribution-USA (nationwide) and the countries of Canada and Switzerland. | Terminated | December 18, 2015 | 690 | 1132096 | Not_Computer
Z-2215-2014 | 68654 | No | JDG | Hip joint femoral (hemi-hip) metallic cemented or uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Moore Hip Prosthesis Long Fenestrated Stem 38 mm (1-1/2 Inch) Head Diameter 203 mm Length. |  |  | July 01, 2014 | 2014 | Zimmer, Inc. | Zimmer is initiating a voluntary recall of 20 production lots of Moore Hip Prosthesis Long Fenestrated Stems due to packaging design verification test failures. Specifically, 9 of 16 test samples failed simulated distribution and shipping testing when the Moore hip stem breached either the inner or outer package tray. Zimmer has not received any complaints of damaged package trays for these 20 lot | Packaging process control | On July 1, 2014 Zimmer, Inc. issued recall notification to their consignees via letter. This notice was issued to affected distributors and hospitals with instructions for returning the affected product. Distributors were sent a letter identifying the issue and their responsibilities. These responsibilities include locating and removing the affected product in their territory, as well as identifying hospitals and surgeons who have previously implanted the affected product and providing related contact information (name, address, and phone number). Distributors are to return on-hand affected products to Zimmer and ensure all of their products are accounted for using the form provided in the letter. Hospital risk managers and surgeons were provided with a letter identifying the issue and their responsibilities. These responsibilities include locating the affected product and returning it to their Zimmer sales representative. |  | 68 | 68 | Worldwide Distribution - US (nationwide) in the states of IN, WI, MN, OR, and TX, and one foreign consignee in Canada. | Terminated | January 09, 2015 | 192 | 68 | Not_Computer
Z-2233-2014 | 68892 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | BD Multi Check CD4 Low Control Catalog No. 340916 is intended as a complete process control for immunophenotyping by flow cytometry. |  |  | July 18, 2014 | 2014 | BD Biosciences, Systems &amp; Reagents | The outer box label contains the incorrect distribution date: Box states 2017-09-02 ; should be 2014-09-02 | Labeling mix-ups | Global hold initiated on July 18, 2014. Customer letters sent by certified mail on July 18, 2014 to advised users of the problem. |  | 85 | 85 | IN, MA, TX, VA, Ca, KY, MD, PA, GA, and BD locations in Singapore and Belgium | Terminated | May 19, 2015 | 305 | 85 | Not_Computer
Z-0975-2014 | 67384 | No | KKY | Polytetrafluoroethylene with carbon fibers composite implant material. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Stryker CMF Customized Implant Kit X-Large; 5444-0-401 X-Large Implant (2) |  |  | December 06, 2013 | 2013 | Stryker Howmedica Osteonics Corp. | Stryker received a report from the field that an incorrect patient-specific Simplex P CMF Customized Implant Kit was delivered to a hospital. | Labeling Change Control | Stryker Orthopaedics sent Urgent Medical Device Recall Notification letters/Product Recall Acknowledgement Forms dated 12/6/2013 to US branch/agency. surgeons, hospital risk managers, and sales representatives. |  | 1 | one unit | Distributed in PA. | Terminated | May 06, 2015 | 516 | 1 | Not_Computer
Z-2719-2014 | 69025 | No | GZB | Implanted spinal cord stimulator for pain relief. | Neurology | Neurology | 510(k) | Percutaneous Trial Lead Kit Model 3086 (8 ch percutaneous lead). |  |  | July 29, 2014 | 2014 | St. Jude Medical, Inc. | The Directions for Use (DFU) manual contains information for devices that are not approved in the US. The correct US approved versions of the manuals were sent to the customers to replace the incorrect international manuals. | Packaging process control | A field action was initiated with customer letters, dated July 29, 2014, and attached replacement DFUs. The letter identified the affected product and the reason for the recall. The recalling firm requested the consignees refer to the replacement DFU for guidance while using the device, and to discard the international manuals. Questions regarding the letter should be directed to customers' St. Jude Neuromodulation Representative or call the Technical Services Department at 1-800-727-7846. |  | 4 | 4 units | Distributed in the states of Texas and Illinois. | Terminated | August 21, 2015 | 388 | 6 | Not_Computer
Z-0977-2014 | 67267 | No | MBP | Resorbable calcium salt bone void filler device. | Orthopedic | Physical Medicine | 510(k) | Medtronic GRAFTON (R) DBM, Flex 2.5cm x 10cm, REF T42100, Rx only, STERILE A, QTY: 1 EA, Medtronic Sofamor Danek, USA, Inc., 1800 Pyramid Place, Memphis, TN 38132 |  |  | January 06, 2014 | 2014 | Medtronic Sofamor Danek USA Inc | Product was labeled as being measured as 2.5 cm x 10 cm, but contained product that measured as 2.5 cm x 5 cm. | Packaging process control | The firm, Medtronic, sent an "URGENT-V voluntary Market Removal (Recall)" communication letter and questionnaire packet (via FedEx) dated January 3, 2014 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to remove the product from normal storage locations; Do not use the product; and as applicable, contact your Medtronic Sales Representative to coordinate the return and replenishment of affected product or call our Global Quality Department at 800-876-3133 ext 3197. In addition, the customers were instructed to disseminate this information to additional personnel within your facility as appropriate. A Medtronic Sales Representative will contact you to facilitate the return of any impacted Set Screw Retaining Drivers you may have in your possession.If you need assistance or if you have any related questions or concerns, please do not hesitate to contact your Medtronic Sales Representative or our Global Quality Department at 800-876-3133 extension 3197. |  | 30 | 30 units | US Distribution in the states of: MO, FL, MN, OH, MI, TX, and NJ. | Terminated | June 24, 2014 | 169 | 30 | Not_Computer
Z-2244-2014 | 68819 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | Magna-Fx Cannulated Screw Fixation System |  |  | July 22, 2014 | 2014 | Zimmer, Inc. | Zimmer is recalling various trauma screws that failed a leak test due to damage of the inner cavity only. | Package design/selection | On 7/22/14, Zimmer sent recall notification to the consignees explaining the issue, identifying affected product, providing the health risks, and instructing the consignees what to do for the recall. Questions or concerns, please contact Zimmer at 1-877-946-2761. |  | 47682 | 47,682 units total | National Distribution: AK, AL, AZ, CA, CT, DC, FL, GA, US VIRGIN ISLANDS, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, NE, NJ, NM, NV, NY, OH, OK, OR, PA, PR, SC, SD, TN, TX, VA, WA, WI, WV, WYForeign Distribution: Canada, Mexico, ARGENTINA, AUSTRALIA, BELGIUM,BRAZIL, CHINA, EL SALVADOR, ENGLAND, GERMANY, INDIA, ITALY, JAPAN, JORDAN, KOREA, NEW ZEALAND, SAUDI ARABIA, SINGAPORE, SPAIN, SWITZERLAND, TAIWAN, UAE | Terminated | June 18, 2015 | 331 | 47682 | Not_Computer
Z-0215-2015 | 69619 | No | IYO | Ultrasonic pulsed echo imaging system. | Radiology | Radiology | 510(k) | AxoTrack I Sterile Procedure Kit 18G, K1001-18XTW, (24/case labeled for US, 5/case labeled for US, 24/case labeled for EU, 5/case labeled for EU). |  |  | March 24, 2014 | 2014 | Soma Access Systems LLC | There is a possibility the sterile packaging may fail before the expiration date specified on the package. | Packaging | Consignees were notified via USPS Certified Letter, email and FedEx on 3/24/2014. |  | 3530 | 3530 units | Worldwide Distribution-US (nationwide) including the states of AL, AZ, CA, CT, FL, GA, IN, KY, LA, MA, MI, MO, NC, NJ, NY OH, OR, PA, SC, TN, TX, VA, and WA and the countries of Belgium, Hong Kong, India, Japan, Panama and Singapore. | Terminated | August 06, 2015 | 500 | 3530 | Not_Computer
Z-1508-2014 | 67818 | No | DIO | Cocaine and cocaine metabolite test system. | Toxicology | Toxicology | 510(k) | ToxCup Drug Screen Cup, Catalogue Numbers: HT15, HT15A, PT11, PT11A, PT15A, PT23A, a screening device for the detection of drugs and drug metabolites in human urine. For in vitro diagnostic use only. |  |  | March 26, 2014 | 2014 | Branan Medical Corporation | May be labeled with incorrect result interpretation graphics on the ToxCup lid label. | Labeling Change Control | The firm, Branan Medical Corp (BMC), sent an "URGENT MEDICAL DEVICE RECALL" letter dated March 26, 2014 to all the customers who received the ToxCup Drug Screen Cup. The letter describes the product, problems identified and the actions to be taken. Customers are instructed to discontinue use and/or sale of product; quarantine all units of these product lots; destroy unused product(s) in compliance with your laboratory and state requirements; share this information with your staff and retain notification as part of your QS documentation; if distributed to your customers, provide a copy of this letter to them; complete and return the enclosed Customer Verification Form via Fax to: Branan Medical Corporation at 949-598-7167 or by e-mail at: branan.complaints@brananmedical.com within 10 days. Should you have any questions about the information contained in this notification, contact Branan Medical Corporation at (949) 598-7166 or email branan.complaints@brananmedical.com. |  | 9800 | 9,800 units | Worldwide Distribution: US (nationwide) and country of Canada. | Terminated | July 23, 2014 | 119 | 9800 | Not_Computer
Z-2253-2014 | 68898 | No | KWY | Hip joint femoral (hemi-hip) metal/polymer cemented or uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Pivot Bipolar Cup 22x42/43-51/53mm, REF 133-YYYY |  |  | July 11, 2014 | 2014 | Ortho Development Corporation | Ortho Development is recalling various Pivot Bipolar Cup, Patella, PS Tibial Insert, UC Tibila Insert, and CK Tibil Insert because the foil pouches did not properly seal, allowing air to seep back into the pouch after sealing which increases the risk that the material properties of the implant would be substantially degraded due to oxidation. | Packaging process control | Ortho Development notified all affected consignees/distributors of the recall via email on July 11, 2014 which included a list of the affected product (part numbers, lot numbers and quantities) under their control. An RMA number and return shipping label were also included with the email. Emails were followed up with phone calls to verify receipt of the notification. |  | 188 | 188 | Worldwide Distribution - US Nationwide and the country of Japan. No military/govt/VA consignees. | Terminated | October 06, 2014 | 87 | 227 | Not_Computer
Z-2264-2014 | 67814 | No | PDZ | Laser surgical instrument for use in general and plastic surgery and in dermatology. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | All brochures printed before November 20, 2013 for the HyperBlue 1530 medical laser, including the product brochure and the online information regarding the product at www.hyperionMed.com before December 3, 2013. |  |  | November 20, 2013 | 2013 | Hyperion Medical | Brochures and web site information for the HyperBlue 1530 diode laser promoted features which had not been cleared in the product's premarket notification to the FDA. The web site and brochures had claims that it could use 940 nm, 808 nm, and 532 nm wavelengths as single wavelengths or in combination with other wavelengths, and a claim that the product could be used with a 15 mm hand piece. | Labeling False and Misleading | Hyperion Medical notified all sales representatives by private courier letter to stop distributing the the sales brochure and return or destroy existing stocks of the brochure. Sales representatives were instructed not to demonstrate the 15 mm hand piece with the HyberBlue 1530 and to remove from any of their individual internet web sites any information about multiple wavelength use or the 15 mm handpiece accessory. The letter to the sales representatives included delivery confirmation. The information on the firm's web site for this product regarding multiple wavelengths and the 15 mm handpiece was removed.For further information please call (214) 914-3133. |  | 3346 | 3,346 brochures. Quantity not applicable to web site. | US Distribution including the states of PA, MD, UT, TX, OH, GA, MO, FL, CO, and WA. | Terminated | May 15, 2015 | 541 | 3346 | Not_Computer
Z-2694-2014 | 68487 | No | HMT | Surgical drape and drape accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | eye-pak 7407 Tray Support Cover REF 8065740745 |  |  | June 03, 2014 | 2014 | Alcon Research, Ltd. | The peel pouches of the Tray Support Covers may be insufficiently sealed so that sterility cannot be assured. | Packaging process control | Alcon Research sent an Urgent Medical Device Recall letter dated June 3, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Please follow these important steps:1) Review your inventory of Tray Support Covers to determine if you have any units from the affected Lots in stock.2) Place the product aside to ensure it is not used.3) Return the attached Response Form via fax or email to Alcon.Please return the Response Form even if you do not have any inventory from the affected lots.4) Alcon will contact you to arrange the return of your inventory.If you have any further questions, concerns, or need assistance with obtaining replacement product, please contact Alcon Customer Service at 1-800-862-5266. |  | 2462 | 2462 | Worldwide Distribution - Nationwide Distribution and the countries of Australia, Belgium, and Canada. | Terminated | July 27, 2015 | 419 | 2462 | Not_Computer
Z-2266-2014 | 68952 | No | DSY | Vascular graft prosthesis. | Cardiovascular | Cardiovascular | 510(k) | Gelweave (TM) Vascular Prostheses |  |  | July 15, 2014 | 2014 | Vascutek, Ltd. | The cautery that is provided with the Gelweave graft in the box has a shorter shelf life indicated on its label, versus that of the graft (4 years and 5 years, respectively). | Incorrect or no expiration date | Vascutek sent an Urgent Field Safety Notice dated July 15, 2014, to all affect customers. The letter identified the product the problem and the action needed to be taken by the customer. User Instructions:A. Vascutek advises that you ensure suitable equipment is available prior to implantation to trim the graft if required.B. For product received with a cautery in the interim, Vascutek request that you check the shelf life of each product before every operation. If you find that the cautery is out of date, remove and discard the cautery. If you require an appropriate cautery supplied as a separate item, please contact your local sales representative for a referral.C.Vascutek are taking this opportunity to reinforce the cautions from the Gelweave TM Vascular Prostheses IFU regarding soaking of the graft in saline to prevent focal burning if a cautery is to be used. This statement is already included in the IFU and now appears in bold, red font to ensure appropriate awareness.This notice needs to be passed on to all persons who need to be aware within your organization or to any organization where the devices have been transferred or distributed.Please maintain awareness of this Field Safety Notice while all actions are taken in your organization and until these products are provided without the cautery.Please return the User Return Slip by e-mail or fax to the Distributor's address given on page 3.This action by Vascutek Ltd. is being taken with the knowledge of the National Competent AuthorityÂż Medicines and Healthcare Products Regulatory Agency (MHRA).If you have any further questions or comments, please do not hesitate to contact us at tcvs.recall@terumomedical.com. Return Completed Form immediately fax to (734) 741-6149. If required, your Vascutek Ltd. representative can discuss and provide more information on the use of cautery with Vascutek products. For further questions call (800) 262-3304 ext. 6056 |  | 63162 | 63,162 | Nationwide Distribution including the states of : AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, WA, WI and WV. | Terminated | March 19, 2015 | 247 | 63162 | Not_Computer
Z-0966-2014 | 66995 | No | GWN | Nystagmograph. | Neurology | Neurology | 510(k) | Normative Data Template CD for use with the I-Portal devices. |  |  | November 21, 2013 | 2013 | Neuro Kinetics, Inc. | The data provided on the Normative Data Template CD for use with the I-Portal devices has not received clearance by FDA through the premarket 510(k) notification process. | No Marketing Application | A recall notification letter, dated November 2013, was sent to Users. A follow up letter regarding scheduling removal of the software is planned for December 2013. |  | 148 | 148 | Distributed USA (nationwide) and the country of Canada. | Terminated | July 20, 2015 | 606 | 148 | Not_Computer
Z-1452-2014 | 67808 | No | BZO | Ventilator tubing. | Anesthesiology | Anesthesiology | 510(K) Exempt | Reusable latex breathing bags with multiple part numbers and sold as part of test lung part number and reusable anaesthesia sets. |  |  | March 25, 2014 | 2014 | Draeger Medical, Inc. | Reusable latex breathing bags with various part numbers and sold as part of test lung and reusable anaesthesia sets were recalled when Draeger became aware of an incident where a user suffered an acute allergic reaction after using a DrÂżger latex test lung. The user did not suffer permanent injury. The breathing bag that belongs to the test lung and other latex breathing bags are labelled NR .A | Labeling design | A recall notification letter, dated March 2014, was sent to wholesale distributors and end users. |  | 433 | 433 | USA Nationwide Distribution in the states of: MI, SC, CO, OH, NE, WI, NC, SD, AZ, TX, MA, WA, PA, KY, MN, GA, OK, NY, IN, VA, UT, MT, ME,FL, CA, MD,MO, IA, TN, MD, WV, IL,NJ, CT, and AR. | Terminated | October 08, 2014 | 197 | 433 | Not_Computer
Z-0237-2015 | 69448 | No | MCZ | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | Enforcement Discretion | Surgical packs (includes trays, bags and kits) |  |  | October 08, 2014 | 2014 | Customed, Inc | Customed products are being recalled because of compromised sterility, not limited to potentially damaged packaging. Serious deficiencies in the manufacturing processes (QSR) and uncontrolled and inadequate storage conditions increase the risk of contamination and may result in patient infection. | Packaging | Customed initiated a recall on October 8th, 2014 of sterile convenience surgical Packs-trays, bags (known as Convenience Packs) and new lots of products added to the initial June 3, 2014 recall of 233 sterile surgical packs. The products are being recalled because of significant compromised sterility, not limited to potentially damaged packaging. Customers were instructed to quarantine and cease use of the affected product identified. Customers are to return the response form included with the letter, to the contact information provided. Upon receipt, the firm will contact customers with further instructions to obtain credit, replacement product, and for return of the recalled product or any other disposition action. If customers have any questions, they should contact the Customed Recall Coordinator at 407-850-5558 extensions 8001 or 8005, or via e-mail: pablo.diaz@customedhealing.com. |  | 185000 | 185000 | US Distribution to Florida, New York and Puerto Rico.. | Terminated | August 14, 2017 | 1041 | 1079228 | Not_Computer
Z-2269-2014 | 68768 | No | KNH | Contraceptive tubal occlusion device (TOD) and introducer. | Obstetrics/Gynecology | Obstetrics/Gynecology | PMA | Gyrus ACMI Falope-Ring Dilator, item 000878-4. The Falope-Ring Dilator is a conical plastic device that is a component of the Falope-Ring Band Contraceptive Tubal Occlusion Device procedure pack (FRB-30). The Falope-Ring Band Contraceptive Tubal Occlusion Device (TOD) consists of a cylindrical radiopaque silicone band that occludes the lumen of the Fallopian tube when placed over a looped segment of the tube. The FRB-30 (used with the reusable applicators) consists of one reusable Falope-Ring Guide, an instruction manual and 30 procedure kits. Each kit is provided sterile and includes one dilator and two Falope Bands. |  |  | July 08, 2014 | 2014 | Gyrus Medical, Inc | Falope-Ring Dilator for Reusable Falope-Ring Band Applicators have been sold as a stand-alone non-sterile component. The dilator is intended to assist placement of the Falope-Ring band on the reusable Falope applicator. | Packaging process control | Olympus sent a "Medical Device Recall" letter dated July 7, 2014 to affected customers. The letter was addressed to Surgery Department &amp; Risk Management Department. The letter described the product being recalled, problem and actions to be take. Consignees were advised to immediately cease further use of any affected product and to remove and discard it. Customers were to fax the completed questionnaire to 484-896-7128. For questions contact Laura Storms at 484-896-5688 or at laura.storms@olympus.com |  | 507 | 507 units | Worldwide - US Nationwide Distribution in the states of AL, AR, CA, CO, CT, DE, FL, GA, IA, IL, IN, LA, MD, MI, ,MN, MS, NH, HY, OH,OR, PA, TN TX, UT, VA, WA, WI,and WV. and the country of CANADA. | Terminated | May 07, 2015 | 303 | 507 | Not_Computer
Z-2688-2014 | 69137 | No | OMP | Powered suction pump. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | RENASYS EZ Negative Pressure Wound Therapy (NPWT) System (RENASYS EZ) for patients who would benefit from a suction device particulary as the device may promote wound healing. |  |  | August 13, 2014 | 2014 | Smith &amp; Nephew, Inc. | Modification were made to the RENASYS EZ NPWT Systems without a 510(k) or PreMarket Approval notification. | No Marketing Application | All RENASYS EZ Products will be removed from the field and customers will be transitioned to other product. On August 13, 2014 the firm sent an "Urgent- Medical Device Removal RENASYS EZ" notification via FedEx overnight asking that the use of the RENASYS EZ should be immediately discontinued and the device should be returned to Smith &amp; Nephew. If you have questions, please reach out to your Customer Care representative or call 1-800-876-1261 Monday through Friday, 8:00 AM to 6:00 PM, Eastern Time. |  | 2142 | 2,142 units | Worldwide Distribution- US (nationwide) including MA, MN, NY and Puerto Rico; Canada, Mexico, Dubai, Australia, Chile, Colombia, Germany, France, England, Spain, Italy, Norway, Hong Kong, Shanghai, New Zealand, South Africa, Singapore, Japan, Thailand and the Caribbean Islands. | Terminated | August 06, 2015 | 358 | 2142 | Not_Computer
Z-1507-2014 | 67834 | No | FOZ | Intravascular catheter. | General Hospital | General Hospital | 510(k) | Arrowgard Blue Plus Antimicrobial Multi-lumen Central Venous Catheter, catalogue number AK-45854-CDC. |  |  | July 17, 2013 | 2013 | Arrow International Inc | Arrow is recalling the Arrowgard Blue Plus Antimicrobial Multi-lumen Central Venous Catheter due to a labeling inconsistency. One portion ofthe lidstock states that the kit contains 8.5FR x 16CM catheter, while the rest of the lidstockcorrectly identifies the catheter as 8.5FR x 20CM | Labeling False and Misleading | Arrow International Inc, ("Arrow") sent a Urgent Medical Device Recall Notification dated July 17, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. 1. If you have affected stock, immediately discontinue use and quarantine any products with the catalog and lot number listed above. 2. To return product, complete the enclosed Recall Acknowledgement Form and fax it to : 1-866-804-9881. Attn: Customer Service. This will allow us to document the amount of product you have on hand for return Goods Authorization (RGA) number and will provide instructions for the return of product to Arrow International.3.If you have no affected stock please complete the enclosed Recall Acknowledgement Form and fax it to : 1-866-804-9881, Attn: Customer Service. This will allow us to document your receipt of this letter.For further questions please contact Arrow Customer Service at 1-800-343-2935. |  | 170 | 170 | US Distribution including the states of IA, OR, WV, MI, AZ, IL, PA, GA, DE, NC, LA, and VA. | Terminated | June 02, 2015 | 685 | 170 | Not_Computer
Z-0847-2014 | 67207 | No | DZE | Endosseous dental implant. | Dental | Dental | 510(k) | Tapered Screw-Vent Implant, MTX, 4.7mm x 11.5mm, 4.5mm Platform. Catalog Number TSVWB11, Lot 62470008. |  |  | January 06, 2014 | 2014 | Zimmer Dental Inc | Zimmer Dental is conducting a voluntary recall of a single lot of the Tapered Screw-Vent Implant, Catalog TSVWB11, Lot 62470008, because some of the packages of this lot may have the Cap Label state Âż3.7 x 10mm instead of Âż4.7 x 11.5mm. | Labeling Change Control | Zimmer Dental initiated the recall on 01/06/2013 by sending out recall notifications via FedEx to domestic customers and via electric mail to international customers.Customers are sent a letter identifying the issue, product description with codes, risk, instructions, and contact information. The letter's instructions include locating the affected product and return unused product to Zimmer Dental and notifying the amount of product that has been consumed. |  | 307 | 307 | Worldwide Distribution-USA (nationwide) and the countries of Thailand, Japan, Hong Kong, Morocco, and UAE. | Terminated | July 07, 2014 | 182 | 307 | Not_Computer
Z-1603-2014 | 68017 | No | DWL | Medical support stocking. | General Hospital | General Hospital | 510(K) Exempt | DROP LOK" Knee Brace***LATEX FREE" |  |  | March 19, 2014 | 2014 | Zimmer, Inc. | During a transfer of products from a recently shutdown facility, the firm discovered raw material labeled as latex free actually contained latex. | Labeling mix-ups | Zimmer sent an Urgent: Device Removal Letters dated April 8, 2014 to their customers. The letter identified the affected product, probllem and actions to be taken. Customers were instructed to review this notification, iIdentify and quarantine the affected product to prevent further distribution or use, complete the attached Response Form and return it via email. orporateQuality_PostMarket@zimmer.com. Return affected product to: Zimmer SurgicalAttn: QA/RA Dept.- Recall200 West Ohio AvenueDover, Ohio 44622 USAPlease include a copy of the Response Form with the shipment.For returns outside the US, please email Rhonda.duncan@zimmer.com obtain an IRA (international return authorization) number. The IRA request should include the part number(s) being returned and the quantity. Please write the associated IRA number on the outside of the box.4. Zimmer will credit your account for returned Drop-Lok" Knee Braces, Cartilage Knee Braces, Hinged Knee Supports, or Neoprene Tennis Elbow Supports. Please return a copy of the completed response form along with your returned product to ensure proper credit.Important: Please distribute this notification to all personnel within your organization who need to be aware. If you have further transferred affected product(s), please provide the customer's information on the Business Response Form to Zimmer.For questions call 330-354-0989 |  | 131 | 131 units | Worldwide Distribution - US Nationwide in the states of: AK, AZ, CA, FL, GA, IA, IL, IN, LA, MD, MI, MO NC, NE, NV, NY, OH, OK, OR, PA, RI, SD, TX, UT, VA, VT, WA, WI &amp; WV and countries of: AUSTRALIA, CANADA, GERMANY, IRAQ, ITALY, SAUDI ARABIA, TUNISIA &amp; UNITES ARAB EMIRATES. | Terminated | April 14, 2015 | 391 | 3548 | Not_Computer
Z-2073-2014 | 68674 | No | GEX | Laser surgical instrument for use in general and plastic surgery and in dermatology. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | LP,Surgical Fibers, Bare Laser Fiber 600 micron, DBLF-60-2, Sterile EO. |  |  | June 20, 2014 | 2014 | Laser Peripherals LLC | Laser Peripherals LLC is recalling DBLF-60-2 Bare Laser due to a mislabeling that indicated a sterility expiration date as '2016-16' as opposed to the intended date of '2016-06'. | Labeling False and Misleading | Laser Peripherals contacted all consignees via telephone on June 20, 2014, in regard to the product being recalled by following the firm's Standard Operating Procedures. The SOP's included Instructions to immediately stop the use of the recalled product and provided other communications related to the recall.For questions regarding this recall call 612-280-1318. |  | 25 | 25 | Nationwide Distribution to IL, IN, and MD | Terminated | October 22, 2014 | 124 | 25 | Not_Computer
Z-1166-2014 | 67308 | No | JOJ | Sterilization process indicator. | General Hospital | General Hospital | 510(k) | VERIFYÂż SixCess 270FP Challenge Pack, STERIS Corporation, 5960 Heisley Road, Mentor, OH 44060, 440-354-2600. |  |  | December 19, 2013 | 2013 | Steris Corporation | STERIS has identified that the outer shipping package for one lot (Lot #024552A) of VERIFYÂż SixCess Challenge Packs incorrectly identifies the lot as Lot #024522A. | Labeling False and Misleading | The firm, Steris, sent an "URGENT VOLUNTARY RECALL NOTICE" dated December 20, 2013 to their customers. The letter describes the product, problem and actions to be taken. The customers were instructed to verify if the outer packaging of your on-hand inventory contains lot number 024522A; destroy any remaining inventory of the affected product in your possession and contact STERIS Customer Service at 1-800-548-4873 to obtain no cost replacement product and if you do not have the affected product in your inventory, no further action is required.If you have any questions regarding this matter, please contact your STERIS Account Manager or STERIS Customer Service at 1-800-548-4873. |  | 1880 | 94 boxes/20 units | Worldwide Distribution: US (nationwide) including states of: IL, LA, ME, MI, NC, OH, OK, VA, WI &amp; WV; and to the country of: Canada.Product was also shipped to Canada. | Terminated | March 26, 2014 | 97 | 1880 | Not_Computer
Z-1793-2014 | 68229 | No | OIE | Specimen transport and storage container. | Pathology | Clinical Chemistry | Enforcement Discretion | APTIMA Urine Collection Kit; Catalog Number |  |  | April 17, 2014 | 2014 | Gen-Probe Inc | Incorrect labeling was used on each pouch of the affected lots of APTIMA Urine Specimen Collection Kits. The outer box of the kits are identified correctly as "APTIMA Urine Collection Kits"; however, the inner pouches are incorrectly labeled as "APTIMA Unisex Swab Collection Kit for Endocervical and Male Urethral Swab Specimens." | Labeling mix-ups | Hologic sent a Recall Notice letter dated April 14, 2014 via overnight mail (FedEx) to affected customers. The letter identified the affected product, problem and actions to be taken. The notification letter requests customers to: 1. segregate the entire contents of the affected kit lots2. reconcile and destroy all partial and/or full kits remaining in inventory and to document; the amount destroyed on the Customer Response Form provided with the Recall Notice; 3. if the urine collection kits were further distributed to other collection sites, then to notify them of this issue; and 4. complete and return the Customer Response Form to Technical Support (contact information is provided in the notification letter).For additional information, questions, or concerns, contact Hologic Molecular (Technical) Support.Toll Free: 888-484-4747Direct: 858-410-8511Fax: 858-410-8250Email: molecularsupport@hologic.com |  | 4740 | 4,740 units | Worldwide Distribution - US Nationwide and the countries of Canada (CA), Denmark (DK), France (FR), Germany (DE), and the United States (US), Belgium (BE), United Kingdom (GB). | Terminated | August 26, 2014 | 131 | 4740 | Not_Computer
Z-0145-2015 | 69224 | No | IYO | Ultrasonic pulsed echo imaging system. | Radiology | Radiology | 510(k) | Parker Eclipse Probe Cover, Latex Free, REF 38-01, 10 Eclipse Probe Covers, Size: 2.37/1.75"W x 9.5"L (60/44 mm x 241 mm), Lot #D424, Expiry date 04/2016, Pre-gelled inside with Aquasonic 100 Ultrasound Transmission Gel |  |  | August 06, 2014 | 2014 | Parker Laboratories, Inc. | Parker Laboratories, Inc. received a report of an incorrect expiration date on the inner boxes and on the plastic bags of the Eclipse Probe Covers. | Incorrect or no expiration date | Parker Laboratories Inc. sent an "Urgent Medical Device Product Recall" Letter/Response Form via Fedex to their customers. The letter indentified the affected product, problem and actions to be taken. For questions contact Parker Laboratories Quality Assurance at 973-276-9500, ext. 7142. |  | 367 | 367 boxes | Worldwide Distribution: US (Nationwide) and country of: Canada. | Terminated | April 29, 2015 | 266 | 367 | Not_Computer
Z-0212-2015 | 69588 | No | JJY | Quality control material (assayed and unassayed). | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | IMMULITE Calcitonin Control Module (IMMULITE, IMMULITE 1000, IMMULITE 2000, IMMULITE 2000 XPi) |  |  | October 20, 2014 | 2014 | Siemens Healthcare Diagnostics, Inc. | A typographical error regarding the expiration date was identified in the control module instructions for use. | Error in labeling | Siemens Healthcare Diagnostics issued an Urgent Medical Device Correction letter October 17, 2014, to US customers, and an Urgent Field Safety Notice issued October 16, 2014, outside the US. The letter identified the product the problem and the action needed to be taken by the customer.Customers are advised to review the notice with their medical directors. Customers should discard the control module IFU listed in table 1 and refer to table 2 for the list of revised IFU versions as noted on the notice. Finally, the customers should check the system for the expiration date and if correct, no further action is required. Field Correction Effectiveness Check should be returned via fax (312) 275-7795 within 30 days. For further questions please call 914-631-8000. |  | 2212 | 2212 units (92 units domestically &amp; 2120 units internationally) | Worldwide Distribution - US Distribution including the states of FL, KY, MN, MA, PA, UT, NY, VA and CA., and the countries of Algeria, Australia, Argentina, Austria, Belgium, Brazil, Bulgaria, Cameron, Chile, China, Colombia, Costa Rica, Cyprus, Czech Republic, Denmark, Estonia, Finland, France, Germany, Great Britain, Greece, Hong Kong, Hungary, India, Iraq, Ireland, Israel, Italy, Jordan, Kazakhstan, Latvia, Macedonia, Mexico, Netherlands,Pakistan, Peru, Poland, Portugal, Rep. Korea, Romania, Russian Fed, Saudi Arabia, Serbia, Singapore, Slovakia, Slovenia, South Africa, Spain, Sweden, Switzerland, Tadjikistan, Taiwan,Turkey, Unit. Arab Emir and Vietnam. | Terminated | March 31, 2017 | 893 | 2212 | Not_Computer
Z-1174-2014 | 67391 | No | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Synthes Small Notch Titanium Reconstructive Plate |  |  | January 27, 2014 | 2014 | Synthes, Inc. | Certain modules containing the Synthes Small Notch Titanium Reconstructive Plate set was distributed by Spine sales consultants although it was contraindicated for certain spine applications. | Labeling design | Synthes sent an Urgent Medical Medical Device Recall letter, dated January 27, 2013, to Users and Sales Consultants. The letter identified the affected product, reason for recall, potential hazard, and actions to be taken. Customers were instructed to complete the verification section. For questions call 610-719-5450 or contact your Synthes Sales Consultant. . |  | 18747 | 18747 | USA Nationwide Distribution | Open, Classified |  |  | 18747 | Not_Computer
Z-0772-2014 | 67098 | No | FGY | Introduction/drainage catheter and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | BD Interlink threaded lock cannula |  |  | November 26, 2013 | 2013 | Becton Dickinson &amp; Company | BD Interlink Threaded Lock Cannula REF# 303369 and lot #3021392 is being recalled due to the potential that a small percentage of the lot may exhibit open seals which adversely impacts unit package integrity and product sterility. | Packaging | BD Diagnostics sent an Urgent: Voluntary Product Recall letters/Notice of Return dated November 26, 2013, to all affected customers. Customers were asked to identify and segregate stock of the affected number, inform the appropriate personnel in their organization to discontinue use of the affected product, return the afected product, following the instructions on the enclosed packing slip, BD would issue a credit for product returned. If product was further distributed customers were instructed to identify their customers and notify them at once of the product recall using the provided customer communication. Customers were asked to complete the attached notice and fax to 855-731-0157. Customers with questions were instructed to call 1-855-731-0157.For questions regarding this recall call 755-731-0157. |  | 268500 | 268,500 products | Worldwide Distribution - USA (nationwide) and Internationally to Canada and Japan. | Open, Classified |  |  | 268500 | Not_Computer
Z-0708-2014 | 67124 | No | HBE | Powered simple cranial drills, burrs, trephines, and their accessories. | Neurology | Neurology | 510(k) | Multiples Anspach Micro Curved Attachments (MCA) and Curved Burr Support S leves (Cutters and Attachments) |  |  | October 26, 2012 | 2012 | The Anspach Effort, Inc. | During a recent review of documentation Anspach determined that the current directions for use for the MCA cutting burrs and MCA Micro Curved Attachment do not provide clear details regarding the intended use specifically for the MCA burrs. The MCA burrs are intended for use in delicate bone in otologic procedures such as cochleostomy. The use of the device in dense bone removal may result in cu | Labeling False and Misleading | Anspach sent an Urgent Medical Removal letter dated October 26, 2012, to all affected customers. The letter stated that they were initiating a Medical Device removal regarding the MCA Cutting Burrs. During a recent review of documentation it was determined that the current directions for use for the MCA cutting burrs (item 18-0079) and MCA Micro Curved Attachment (item 18-0060) did not provide clear details regarding the intended use specifically for the MCA burrs. The MCA burrs are intended for use in delicate bone in otologic procedures such as cochleostomy. The use of the device in dense bone removal may result in cutter/ burr fracture. If a fracture occurs, contact Customer Support at (800) 327-6887 or email ProductSupportTeam@synthes.com. If the procedure was done successfully, no additional patient follow up is necessary.Customer Immediate Actions:1. Please remove and return any MCA items from your inventory immediately.2. Complete the attached reply form indicating your receipt of this letter. Return the completed form by fax or email to the number or email address provided on the form. If you distribute any of the products to other services or facilities, please forward this information as appropriate. Customers with questions were instructed to contact Anspach Customer Support at (800) 327-6887 or email ProductSupportTeam@synthes.com.For questions regarding this recall call 1-800-327-6887. |  | 14330400 | 6,824 imits (4,7324 in USA / 2,100 Foreign) | Worldwide Distribution - USA (nationwide) and Internationally to Canada, Netherlands, Great Britain, France, Sweeden, Italy, Spain, Japan, South Africa, Australia, Switzerland, | Terminated | June 23, 2014 | 605 | 14330400 | Not_Computer
Z-0147-2015 | 69446 | No | DRC | Trocar. | Cardiovascular | Cardiovascular | 510(K) Exempt | Blunt Hurd Cannula Blunt Tip 10 mm with 5 mm Reducer simplifies open laparoscopy. |  |  | September 13, 2013 | 2013 | CooperSurgical, Inc. D.B.A. Lone Star Medical Products | CooperSurgical is recalling two lots of 900-863 because the wrong package insert (IFU) was placed into the shipper. | Error in labeling | CooperSurgical sent an Urgent Field Safety Notice ( Correction ) dated September 27, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.The recalling firm provided correct package inserts (IFU) to affected consignees. The recalling firm intends to conduct effectiveness checks by reviewing return receipts from the initial recall communication.We apologize for any inconvenience this communication may cause. If you have any questions or concerns, please contact CooperSurgical at (203) 601-9818. |  | 90 | 90 units | Worldwide Distribution: US (nationwide) to North Carolina; and Internationally to Australia and Italy. | Terminated | September 23, 2015 | 740 | 90 | Not_Computer
Z-0131-2015 | 67987 | No | CHI | Lipase test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Lipase Reagent, 7D80. |  |  | April 07, 2014 | 2014 | Abbott Laboratories, Inc | Package inserts have incorrect SmartWash parameters for Triglyceride. | Labeling Change Control | Firm issued a Product Correction Letter to all Lipase reagent customers who received lots 43827UQ12, 43819UQ03, 44449UQ05, 44127UQ06, 44334UQ07, 43994UQ07, 44596UQ08, 44775UQ09, and 44930UQ11. The letter identified the affected product and provided an explanation of the problem, as well as, the patient impact and actions to be taken. The Product Correction Letter provided guidance for the laboratory to correctly configure the Architect cSystem's Lipase assay parameters. Customers are to retain a copy of the letter for their lab records, and forward a copy to any laboratories that the product was further distributed to. Questions should be directed to Customer Service at 1-877-4ABBOTT or your local customer service for customers outside of the US. |  | 11916 | 11916 units | Worldwide Distribution: US (nationwide) and countries of: Algeria, Angola, Argentina, Armenia, Australia, Austria, Bangladesh, Belgium, Bosnia and Herzegovina, Brazil, Bulgaria, Canada, Chile, China, Columbia, Croatia, Czech Republic, Dominican Republic, Egypt, Estonia, France, Germany, Greece, Guatemala, Hungary, India, Indonesia, Iraq, Ireland, Israel, Italy, Jamaica, Jordan, Kenya, Kuwait, Kyrgyzstan, Latvia, Lebanon, Libya, Lithuania, Luxembourg, Moldova, Montenegro, Morocco, Netherlands, New Zealand, Norway, Pakistan, Panama, Paraguay, Philippines, Poland, Portugal, Qatar, Republic of Yemen, Romania, Saudi Arabia, Senegal, Serbia, Slovenia, South Africa, South Korea, Spain, Sweden, Switzerland, Taiwan, Thailand, Trinidad and Tobago, Tunisia, Turks and Caicos, Turkey, Uganda, United Arab Emirates, United Kingdom, and Uruguay. | Terminated | April 01, 2015 | 359 | 11916 | Not_Computer
Z-1161-2014 | 67477 | No | OMP | Powered suction pump. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Smith &amp; Nephew RENASYS-G Sterile, Gauze Dressing Kit with Soft Port REF #'s: 66800961 (The RENASYS Soft Port Dressing Kits are used in conjunction with negative pressure wound therapy pumps) |  |  | January 09, 2014 | 2014 | Smith &amp; Nephew Inc. | Reports of wound fluid and /or blood that were not evacuated from beneath the wound dressing and the pump blockage alarm did not activate has been associated with inadequate Soft Port sealing at the wound site or improper positioning of the Soft Port orifice over the wound contact dressing drape. Inadequate sealing can result in air leaks and contribute to inefficient fluid evacuation from the wo | Packaging change control | A product bulletin titled "Soft Port Enhancements" dated March 14, 2013, was released for external distribution describing the changes to the Soft Port products, e.g. increasing the Soft Port orifices size and adding a second slit to the release paper covering the port adhesive, and stating the purpose of the modifications to the device. These product bulletin was provided to sales representatives for distribution to customers requesting information regarding the design enhancement made to the Soft Port product range. The bulletin was not distributed to all the customers.To ensure that the product modification and its purpose was appropriately communicated to consignees of affected product a Medical Device Correction Notice, dated 01/22/2014, retrospectively advising customers of the modification to the Soft Port products has been prepared and was sent to consignees of the affected products on the same date. |  | 7353 | 7,353 individuals units | Worldwide Distribution-USA (nationwide) including Puerto Rico and the states of AK, AL, AR, AZ, CA, CO, CT, FL, GA, ID, IL, KY, LA, MA, MS, MT, NC, NJ, NV, NY, OH, OR, PA, TN, TX, VA, WA, WI, WV, HI and the countries Dubai, Curacao, Argentina, Australia, Chile, Germany, France, England, Hong Kong, Kuwait, Lebanon, Libya, Sri Lanka, Malaysia, New Zealand, Manila, Saudi Arabia, South Africa, Singapore, Turkey, Taipei and Vietnam. | Terminated | September 26, 2014 | 260 | 2423739 | Not_Computer
Z-1833-2014 | 68072 | No | KTT | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Synthes External Fixation Systems (Small, Medium, Large and DO), bone fixation appliance components. |  |  | April 10, 2014 | 2014 | Synthes, Inc. | Labeling changes have been made related to MR (Magnetic Resonance imaging) conditions as a result of changes in required protocols to designate a product MR Safe, MR Conditional, or MR Unsafe. | Labeling Change Control | Synthes sent an Urgent Notice dated April 10, 2014, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to:All Synthes External Fixation devices should be treated as MR Conditional."Synthes asks that you review the information contained in this Labeling Notification and complete the Verification Section located on Page 3."Discard outdated revisions of the Technique Guides noted in the table below."Update your records with updated Labeling Infonnation."Forward this Medical Device Labeling Update to anyone in your facility that needs to be infonned, especially those personnel that conduct MR testing."If the Verification Fonn is answered on behalf of more than one facility and/or individual, please clearly indicate the name and address of the facility and/or individual in pg. 3."Updated product literature can be located on the Synthes website athttp://www.syntheskyo.com/trauma kyo/home/home.htm."Please see the attached inseii for current conditions for use in the MR environment."Maintain a copy of this notice.Customers with questions were instructed to call 610-719-5450.For questions regarding this recall call 610-719-5000. |  | 5114052 | 5,114,052 | Worldwide Distribution - USA (nationwide) and Canada. | Open, Classified |  |  | 5114052 | Not_Computer
Z-1164-2014 | 67658 | No | DQY | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | Boston Scientific, RunWay Guide Catheter, UPN H749389696640, Catalog Number 38969-664. |  |  | March 03, 2014 | 2014 | Boston Scientific Corporation | The sterile pouch of the devices from this lot may not be completely sealed and sterility may be compromised. The most common adverse health consequence expected from an unsealed pouch is a transient bacteremia, which may require antibiotic therapy. | Packaging | Consignees were sent a Boston Scientific "Urgent Medical Device Recall Removal " letter dated March 3, 2014. The letter described the problem and the product involved in the recall. The letter advised consignees to return the affected product and to complete and return the Account Reply Verification Tracking Form. The local Sales Representative can answer any questions that you may have regarding this Recall Removal. |  | 7 | 7 | TX and Russia | Terminated | June 09, 2014 | 98 | 7 | Not_Computer
Z-1843-2014 | 68352 | No | PAM | Multiplex nucleic acid assay for identification of microorganisms and resistance markers from positive blood cultures. | Microbiology | Microbiology | 510(k) | FirmArray Blood Culture Identification (BCID) Panel, Model 2.0 |  |  | May 27, 2014 | 2014 | BioFire Diagnostics, Inc. | BioFire has identified an increased risk of false positive results when the FilmArray Blood Culture Identification (BCID) Panel is used with bioMÂżrieux BacT/ALERT Standard Anaerobic (SN) blood culture bottles. False positive results have been observed for Pseudomonas aeruginosa and Enterococcus. | Labeling design | BioFire Diagnostics sent an Urgent Recall/Correction letter on May 28, 2014. Customers were instructed If the BCID Panel is used to test BacT/ALERT SN bottles, positive results for Pseudomonas aeruginosa and Enterococcus should be confirmed by another method prior to reporting the test results.Customer with questions were instructed contact customer support at Support@biofiredx.com or via telephone by dialing 1-800-735-6544 and selecting option 5 for customer support and then option 1 for FilmArray.For questions regarding this recall call 801-736-6354, ext 360. |  | 1084 | 1084 | Worldwide Distribution - USA (nationwide) European Union, Hong Kong, and VA. | Terminated | November 07, 2014 | 164 | 1084 | Not_Computer
Z-1152-2014 | 67418 | No | NOT | Immunohistochemistry reagents and kits. | Hematology | Pathology | 510(k) | HER2 Image Analysis |  |  | February 07, 2014 | 2014 | Leica Biosystems Imaging, Inc. | In an abundance of caution, Aperio (now known as Lecia Biosystems) is providing additional information to clarify that the use of Digital IHC Breast/ER/Dako Image Analysis, Digital IHC Breast/HER2 Digital Read, Digital IHC Breast/HER2/Dako Image Analysis, Digital IHC Breast/PR Digital Read and Digital IHC Breast/PR/Dako Image Analysis with the RUO products identified in that earlier notification i | Labeling False and Misleading | Aperio (now known as Leica Biosystems) sent an recall notification letter on February 7, 2014, via USPS mail. The letter dated January 31, 2014, informed customers of the reason for the notification, product description, provided additional information to clarify that the use of Digital IHC Breast/ER/Dako Image Analysis, Digital IHC Breast/HER2 Digital Read, Digital IHC Breast/HER2/Dako Image Analysis, Digital IHC Breast/PR Digital Read and Digital IHC Breast/PR/Dako Image Analysis with the RUO products identified in that earlier notification is also for Research Use Only and not for use in diagnostic procedures. The notice also included updated manuals for the image analysis algorthithms so that the intended use in the hardware and software manuals is consistent. An acknowledgement form and contact information was also provided.We sincerely apologize for any inconvenience this may have caused. If you have any questions, please do not hesitate to contact Aperio Technical Support at (866) 478-3999. |  | 78 | 78 | US Distribution including the states of FL, WA, PA, MS, TX, MN CA, NY, NJ, NC, CO and OH. | Terminated | April 15, 2014 | 67 | 78 | Not_Computer
Z-0790-2014 | 66933 | No | LLH | Microorganism differentiation and identification device. | Microbiology | Microbiology | 510(K) Exempt | BD GeneOhm Cdiff Assay, Part #441401, containing Lysis Kit sample buffer tubes, 200 tests/kit; labeled as ***GeneOhm Sciences Canada, Inc. 2555 boul. du Parc-Technologique Quebec, Qc, Canada, G1P 4S5*** |  |  | October 11, 2013 | 2013 | Becton Dickinson &amp; Co. | An incorrect reagent was packed with an in vitro diagnostic kit, which may cause an increase in unresolved results or false negative results for Clostridium difficile in patient samples. | Error in labeling | Becton Dickinson sent an Urgent Product Recall Letter dated October 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customers.They were advised that a portion of the BD GeneOhm Cdiff assay test kits from one lot, 08T13108, may contain incorrect bags of sample buffer tube. Customers were requested to return response forms indicating acknowledgement of the recall and quantity of remaining product that was discarded. Please fax to 410-316-4258. |  | 33 | 33 kits | US Distribution including the states of NY, IN, UT, MN, OH, MA, SD, CA, NJ, and PA. | Terminated | February 19, 2014 | 131 | 33 | Not_Computer
Z-1851-2014 | 68387 | Yes | BTR | Tracheal tube. | Anesthesiology | Anesthesiology | 510(k) | RUSCH, LaserTube (Rubber), Laser Resistant Tracheal Tube, cuffed, Teleflex Medical, Willy-Rusch, GmbH. |  |  | May 23, 2014 | 2014 | Teleflex Medical | Although the Natural Rubber Latex caution symbol is clearlydisplayed on the product label, the accompanying caution statement "Caution: This product contains Natural Rubber Latex which may cause Allergic Reactions" was not printed on the product label. | Error in labeling | Teleflex Medical sent a Urgent Medical Device Recall Notification dated June 6, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Our records indicate that you have received products that are subject to this recall. We are now notifying our customers to take the following actions:1. If you have affected stock, immediately discontinue use and quarantine any products with the catalog numbers listed above from the lot numbers listed on the attachment.2. To return product, complete the enclosed Recall Acknowledgement Form and fax it to 1-866- 804-9881, Attn: Customer Service. This will allow us to document the amount of product you have on hand for return. A customer service representative will contact you with a Return Goods Authorization (RGA) Number and will provide instructions for the return of product to Teleflex Medical.3. If you have no affected stock, please complete the enclosed Recall Acknowledgment Form and fax it to 1-866-804-9881, Attn: Customer Service. This will allow us to document your receipt of this letter.Teleflex Medical is committed to providing high quality, safe and effective products. We sincerely apologize for any inconvenience this action may cause your operations. If you have any other questions, feel free to contact your local sales representative or Customer Service at 1-866-246-6990. |  | 1658 | 1658 ea | Nationwide Distribution and the country of Colombia. | Terminated | January 13, 2016 | 600 | 1658 | Not_Computer
Z-0138-2015 | 69279 | No | NAD | Occlusive wound dressing. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Biatain Super Non-Adhesive, Catalog number 46300, 10 pcs, REF 0463001400. Country of Origin:HU. |  |  | September 08, 2014 | 2014 | Coloplast Manufacturing US, LLC | Coloplast is conducting a recall for one lot (4201931) of the Biatain Super Non-Adhesive wound dressing which was mislabeled. | Labeling mix-ups | Coloplast sent an "Urgent Voluntary Recall Notification" dated September 9, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. If they have further distributed the product they need to contact their customers immediately. For questions they can contact Coloplast Customer Care Representative at 800-533-0464.A second follow up "Urgent Voluntary Recall Notification" letter dated September 23, 2014 was also sent to consignees. |  | 6 | 6 | US Distribution to MO and RI. | Terminated | January 21, 2015 | 135 | 6 | Not_Computer
Z-1273-2014 | 67638 | No | DQK | Programmable diagnostic computer. | Cardiovascular | Cardiovascular | 510(k) | Spacelabs Healthcare G2 Clinical Access Mislabeled CDs for PN: 063-1829-11 Rev A and Update Kit, PN: 040-1565-10. |  |  | February 14, 2014 | 2014 | Spacelabs Healthcare, Llc | Spacelabs has voluntarily recalled G2 Clinical Access Software CD because the programs installed on the Clinical Access CD are incorrect and it will prevent the customer frominstalling Clinical Access. | Labeling Change Control | Spacelabs Healthcare began sending out the Medical Device Correction letter, dated 28 February 2014, to domestic customers on 28 February 2014 and to international customers by 7 March 2014. The letter identified the product, the product, the problem, and the action to be taken by the customer. Customers are also directed to call 1-800-522-7025 and select 2 for Technical Support.Customers are advised to dispose the CD, PN: 063-1829-11 Rev A previously supplied and replace it with the replacement CD (Rev B) provided. |  | 36 | 36 units total: 17 units were distributed in the US and 19 units distributed outside the US. | Worldwide Distribution- USA (nationwide) including the states of Alabama, Florida, Georgia, Louisiana, Missouri, New York, South Carolina, Tennessee, Texas and Wyoming, and the countries of Australia, Canada, Czechoslovakia, China, Colombia, France, Great Britain, Italy, Netherlands, Poland, Saudi Arabia, and Turkey. | Terminated | May 20, 2014 | 95 | 36 | Not_Computer
Z-0132-2015 | 69101 | No | MWS | Cardiovascular surgical instruments. | Cardiovascular | Cardiovascular | 510(K) Exempt | Acrobat V Vacuum Off-Pump System |  |  | August 13, 2014 | 2014 | Maquet Cardiovascular, LLC | An internal finding identified the presence of pinholes/cuts in some of the Tyvek pouches of the vacuum tubing sets included as an accessory to Acrobat V Vacuum Off-Pump System, Acrobat SUV Vacuum Off-Pump System, Axius XPOSE 3 Access Device and XPOSE 4 Access Device. | Package design/selection | Maquet Inc. issued Recall letters/Acknowledgement forms dated via Fed-Ex on 8/13/2014. |  | 16532 | 16,532 units (19,270 units international) | Worldwide Distribution: US (nationwide) and countries of: ARGENTINA, AUSTRIA, BANGLADESH, BELARUS, BRAZIL, BRUNEI, CANADA, CHILE, CHINA, COLOMBIA, CROATIA, CZECH REPUBLIC, DENMARK, ECUADOR, GERMANY, FINLAND, FRANCE, GREECE, HONG KONG, INDICA, INDONESIA, ISRAEL, ITALY, JAPAN, MEXICO, MALAYSIA, NETHERLANDS, PANAMA, PERU, PHILIPPINES, POLAND, PORTUGAL, RUSSIA, SAUDI ARABIA, SINGAPORE, SLOVAK REPUBLIC, SOUTH KOREA, SPAIN, SWITZERLAND, TAIWAN, TURKEY, UNITED ARAB EMIRATES, UNITED KINGDOM AND VIETNAM. | Terminated | March 30, 2016 | 595 | 16532 | Not_Computer
Z-0124-2015 | 69417 | No | KNW | Gastroenterology-urology biopsy instrument. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Spirotome" Soft-Tissue Biopsy Needle Set. |  |  | September 26, 2014 | 2014 | Cook Medical Incorporated | MEDINVENTS, the manufacturer of the Spirotome&quot; Soft-Tissue Biopsy Needle Set distributed by Cook Medical, has initiated a voluntary recall of these products. The recall hasbeen initiated due to customer complaints that the protective caps were dislodged in the pouch. Furtherinvestigation disclosed potential dislodging of protective caps on the needle tips on other distributedproducts. | Package design/selection | On 9/26/2014, URGENT: MEDICAL DEVICE RECALL notifications were sent to the affected distributors and institutions Risk Managers/Recall Administration with instructions for returning the affected product. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. For questions, call Customer Relations at 1-800-457-4500 or 1-812-339-2235. |  | 1239 | 1239 units | Worldwide Distribution-USA (nationwide) including the states of Arizona, California, Colorado, Hawaii, Iowa, Illinois, Massachusetts, Maryland, Michigan, Missouri, New Mexico, Pennsylvania, Texas, Virginia, Vermont, and Wisconsin, and the countries of AT; BE; CH; CY; CZ; DE; DK; ES; FI; FR; GB; HU; IC; IE; IS; IT; LB; LT; LU; NL; NO; PT; QA; SE; TR; and ZA. | Terminated | February 02, 2016 | 494 | 1239 | Not_Computer
Z-1868-2014 | 68330 | No | MBP | Resorbable calcium salt bone void filler device. | Orthopedic | Physical Medicine | 510(k) | Accell Evo3C Demineralized Bone Matrix, 5 cc, Model Number: 02-6000-050 , Lot Number: 132160. |  |  | May 19, 2014 | 2014 | Integra LifeSciences Corp. | This lot of Accell Evo3C Demineralized Bone Matrix, 5cc was incorrectly packaged in boxes that describe the product as Accell Evo3 Demineralized Bone Matrix. Evo3C is contraindicated for treatment of vertebral compression and not indicated for use in the posterolateral spine. | Packaging | Integra sent an Urgent Voluntary Medical Device Recall dated May 19, 2014 to customers via mail (traceable courier service). The recall notification letter identified affected product, reason for recall and actions to be taken Customers were instructed to complete the "RECALL ACKNOWLEDGEMENT AND RETURN FORM". and return to:FCA1@integralife.com or FAX to 1-609-275-9445. For questions contact Customer Service at 1-800-550-7155. |  | 61 | 61 | US Nationwide Distribution | Terminated | September 05, 2014 | 109 | 61 | Not_Computer
Z-0117-2015 | 69379 | No | DZE | Endosseous dental implant. | Dental | Dental | 510(k) | Straumann Bone Level Tapered Implant 04.1 mm RC, SLActiveÂż 8mm, RoxolidÂż, Loxim |  |  | September 19, 2014 | 2014 | Straumann USA, LLC | Assembled with Narrow Connect (NC) transfer piece instead of Regular Connect (RC) | Packaging process control | Strauman notified consignees of this Field Corrective Action by letter dated September 19, 2014 and delivered via hand by Straumann Representative. Accounts are requested to return inventory. Contact (978) 747-2514 for assistance. |  | 6 | 6 units | US Distribution in states of: FL and MA. | Terminated | March 13, 2015 | 175 | 6 | Not_Computer
Z-1928-2014 | 68469 | No | NBC | B-type natriuretic peptide test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Alere Triage BNP Calibrators for the Beckman Coulter Access Family of Immunoassay Systems, Model No. 98202 |  |  | May 30, 2014 | 2014 | Alere San Diego, Inc. | Alere San Diego is recalling Alere Triage BNP Calibrators for the Beckman Coulter Access Family of Immunoassay Systems because the bottle labeled as Calibrator Level 5 actually contains Calibrator Level 4. This results in a calibration failure when attempting to calibrate using a calibrator kit which contains the affected Level 5 Calibrator. | Error in labeling | Alere sent an Urgent Medical Device Recall letter dated May 30, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. The primary consignees were instructed to block all stock of the affected lot in their warehouses; discontinue use and return all stock on hand, complete the provided verification form to document receipt of the notice and quantity to be returned, contact their primary consignees and advise them of the requireddiscontinuation and return of the affected lot. For customers (or their sub-distributors) who have inventory on hand, Alere San Diego, Inc. will provide replacement product to the customer for the quantity of units that they must return.Customers are instructed to complete the attached reply form with the amount of unused product that will be returned for replacement. Customers with questions are instructed to contact Alere San Diego at 9975 Summers Ridge Rd, San Diego, CA 92121, (877) 308-8287, triage.support@alere.com |  | 3121 | 3,121 kits | Worldwide Distribution - Nationwide and the countries of Australia, Canada, China, EU, Mexico, Russia, Singapore, South Africa, South Korea, Taiwan and United Arab Emirates. | Terminated | February 11, 2015 | 257 | 3121 | Not_Computer
Z-1930-2014 | 67011 | No | JJY | Quality control material (assayed and unassayed). | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | MAS Omni Immune PRO Controls |  |  | December 02, 2013 | 2013 | Microgenics Corporation | The constituent, Inhibin A, was removed from MAS Omni IMMUNE and MAS Omni IMMUNE PRO controls; however the package insert was not updated to reflect the change. | Labeling Change Control | Urgent Medical Device Correction letters were sent on December 2, 2013 to all affected customers and distributors via trackable mail. |  | 1393 | 1393 | Worldwide Distribution-USA (nationwide) and the countries of Australia, China, Germany, Hong Kong and Malaysia. | Terminated | July 08, 2014 | 218 | 4578 | Not_Computer
Z-1149-2014 | 67562 | No | NZW | Intravascular catheter. | General Hospital | General Hospital | 510(k) | Heparin I.V. Flush Syringe is available is pre-filled polypropylene syringes, 100 units/mL, 3 mL fill in a 12 mL syringe. |  |  | February 19, 2014 | 2014 | Medefil Incorporated | The bar code that is printed on the Master Carton label does not correspond to the product code fully displayed on the Master Carton label. The non-matching information can only be detected by using a bar code reader. The defective bar codes will be read as Product No. MIH-3335 ( Heparin I.V. Flush Syringe 100 units/mL; 5 mL fill in 12 mL Syringe) whereas this product is actually Product No. MIH-3 | Labeling Change Control | URGENT MEDICAL DEVICE CORRECTION letters dated February 19, 2014 were sent to all direct accounts. The letters included instructions to: quarantine the recalled products; and,complete and return the attached customer response form (via Fax to 630-681-9100 or via mail to Medefil Inc., 250 Windy Point Drive, Glendale Heights, IL 60139) so that the correct number of replacement labels can be sent from Medefil. |  | 1608 | 1,608 Master Cartons | USA (nationwide) including the states of AZ, CA, CO, FL, GA, IL, LA, MA, MD, MI, MN, MO, MS, NC, NJ, NY, OH, OK, PA, SC, TX, UT, and WI, and the country of Canada. | Open, Classified |  |  | 1608 | Not_Computer
Z-1942-2014 | 68482 | No | HSX | Knee joint femorotibial metal/polymer non-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Stride Femoral Component Size 2, PFSI-00053, Right Lateral /Left Medial. Unicondylar Knee prosthesis. |  |  | June 04, 2014 | 2014 | Blue Belt Technologies MN | Mislabeled. Packages labeled PFSI-00053, Right Lateral/Left Medial contains PFSI-00054, Right Medial/Left lateral implants. Packages labeled PFSI-00054, Stride Right Medial/Left Lateral contain PFSI-00053, Left Medial/Right Lateral. | Labeling mix-ups | Firm initiated recall on 6/4/2014 by Blue Belt Technologies representatives hand delivering an Urgent Medeical Device Recall letter. The letter was addressed to the attention of Safety Officer/Regulatory Compliance Officer. The letter described the issue of the labeling mix-up, asked for devices to be quarantined, and stated that a Blue Belt Technologies Clinical Specialist will be contacting the consignee or return and replacement. A Receipt Acknowledgement form was asked to be returned by fax or mail. |  | 2 | 2 | US Nationwide Distribution in the states of CA, PA, and TX | Terminated | August 04, 2014 | 61 | 2 | Not_Computer
Z-1633-2014 | 67833 | No | FOZ | Intravascular catheter. | General Hospital | General Hospital | 510(k) | Arrowgard Blue and Arrowgard Blue Plus Pressure Injectable Central Venous Catheter, a multiple-lumen catheter that permits venous access to central circulation. Catalog number ASK-45703-PHF. |  |  | July 17, 2013 | 2013 | Arrow International Inc | The labeling states that the kit contains Introducer Needle: Echogenic 18 Ga. x 2.5 (6.35cm) XTW and 5mL ArroW Raulerson Spring-Wire Introduction Syringe; however, the kit contains, and should state, Introducer Needle: 18 Ga. x 2.5 (6.35cm) XTW and 5mL ArroW Raulerson Spring-Wire Introduction Syringe. As a result of this labeling inconsistency, if the clinician were to place the needle under x-r | Labeling False and Misleading | Arrow sent an Urgent Medical Device Recall Notification letter dated July 17, 2013, was sent to the customer requesting they discontinue use and quarantine any remaining affected product for return to Arrow.Our records indicate that you have received products that are subject to this voluntary recall. We are now asking that you take the following actions:1. If you have affected stock, immediately discontinue use and quarantine any products with the product and lot numbers listed above.2. To return product, complete the enclosed Recall Acknowledgement Form and fax it to 1-800-343-2935,Attn: Customer Service. This will allow us to document the amount of product you have on hand for return. Your sales representative will contact you with a Return Goods Authorization (RGA) Number and will provide instructions and assistance for the return of product to Arrow International.3. If you have no affected stock, please complete the enclosed Recall Acknowledgment Form and fax it to 1-800-343-2935, Attn: Customer Service. This will allow us to document your receipt of this letter. |  | 4285 | 4285 | US Distribution only to MI. | Terminated | March 30, 2015 | 621 | 4285 | Not_Computer
Z-1792-2014 | 68294 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | cobas 8000 analyzer series Operators Manual Roche cobas |  |  | May 21, 2014 | 2014 | Roche Diagnostics Operations, Inc. | Under certain conditions loose ProCell/CleanCell aspiration tube filters may cause: consecutive discrepant results; sipper alarm; foam on reservoirs and system reagent alarms. | Labeling design | Roche sent an Urgent Medical Device Correction (UMDC) letter dated May 20, 2014 via UPS Grount (receipt required) to affected customers. The letter identified the affected product, problem and actions to be taken. The letter instructed customers to: 1. Follow the instructions in the Precautions Regarding Handling the Aspiration Tube Filter section in this UMDC; 2. Discontinue performing ProCell/Clean Cell aspiration tube filter check as this maintenance item willbe performed by your Roche FSR as a part of service maintenance; 3. If you have distributed the products listed within this UMDC to other sites or facilities, please provide them with a copy of the UMDC; and 4. Complete the faxback form enclosed with this UMDC and fax it to 1-888-410-6620. For questions contact the Roche Support Network Customer Support Center, 24 hours a day, seven days a week at 1-800-428-2336 if you have questions about the information contained in this UMDC. |  | 149 | 149 | US Nationwide Distribution | Terminated | June 04, 2015 | 379 | 1758 | Not_Computer
Z-1787-2014 | 65483 | No | MUE | Full-field digital mammography system. | Radiology | Radiology | 510(k) | Beekley Spots Light Image O-Spots for Mammography |  |  | June 19, 2013 | 2013 | Beekley Corporation | Light Image O-Spots Dispenser mislabeled as Soft n' Stretchy Light Image S-Spots | Error in labeling | Beekley sent an Urgent Medical Device Recall letters dated June 19, 2013 via FedEx with signature and tracking confirmation. The letter identified the affected product, problem and actions to be taken. Firm requests users to examine inventory, remove and return recalled product. For questions contact Chelsea Fuller, Regulatory Manager at 1-800-233-5539 x 491 |  | 81 | 81 boxes | Worldwide Distribution - USA in the states of CA, FL, MA, NC, NY, and PA. | Terminated | June 16, 2014 | 362 | 81 | Not_Computer
Z-1945-2014 | 68640 | Yes | KDI | High permeability hemodialysis system. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Baxter, MiniCap with Povidone-Iodine solution, 5C4466P. For dialysis. |  |  | July 08, 2013 | 2013 | Baxter Healthcare Corp. | the white printed paper layer on the pouch packagingmay not be fully adhered to the pouch foil layer. The upper paper layer may serve as a protective barrier and help support overall packaging. | Packaging change control | An Urgent Product Recall notification was mailed via U.S.P.S. first-class mail to facilities, (dialysis centers and distributors) on July 8, 2013 and mailed by first-class mail to home patients on July 11, 2013. |  | 20940 | 20,940 units | Distributed in the states of AR, TN MO, KS, TX, MS, and AL. | Terminated | June 30, 2014 | 357 | 20940 | Not_Computer
Z-0749-2014 | 63624 | No | HBE | Powered simple cranial drills, burrs, trephines, and their accessories. | Neurology | Neurology | 510(k) | Micro Curved Attachment (MCA) Cutting Burr, DFU 18-0001. All lots are included in recall. Part Number MCA |  |  | March 29, 2012 | 2012 | The Anspach Effort, Inc. | The Anspach Effort, Inc. is recalling all lots of the Anspach Micro Curved Attachment (MCA) Cutting Burrs and Dissection Tools. The recall was initiated after the indications for use were found to be broader than the validation of the MCA Cutting Burrs and its directions for use did not clearly state the correct intended use of the device. After a labeling correction, Anspach Effort, Inc. decide | Labeling Change Control | Firm is performing a Medical Device Recall by sending customers a recall letter and requesting the return of Anspach Micro Curved Attachments and Dissection Tools currently on hand. |  | 163 | 163 | Worldwide Distribution: US (nationwide) and Internationally to: AU, BE, CA, CH, ES, FR, GB, IT, JP, NL, SE, and ZA. | Terminated | March 11, 2015 | 1077 | 1412 | Not_Computer
Z-1698-2014 | 68065 | No | NJT | Immunohistochemistry reagents and kits. | Hematology | Pathology | 510(K) Exempt | Novocastra Liquid Mouse Monoclonal Antibody Glutathione S-Transferase pi (NCL-L-GSTpi-438), a liquid tissue culture supernatant. For in vitro diagnostic use. |  |  | February 26, 2014 | 2014 | Leica Microsystems, Inc. | Novocastra Liquid Mouse Monoclonal AntibodyGlutathione S-Transferase pi (Product Code: NCL-L-GSTpi-438) does not function as intended up to the expiry date on the product labeling. This was detected through an ongoing stability program. There is a linkage between the age of the product and staining intensity. | Incorrect or no expiration date | Leica Biosystems sent a Medical Device Recall Notification letter dated February 26, 2014 to all direct accounts. The letter advised that clinicians may want to review results obtained from the affected products. The letter also included instructions for customers to: 1) appropriately destroy any unused or partially used NCL-L-GSTpi-438 (lots 6018777, 6023260,6026297, 6026980) reagent or confirm that they are not in stock; 2) complete and return the attached Medical Device Recall Notification Acknowledgement Form to 847-236-3747; and, 3) contact the local Leica representative immediately if replacement is required. Customers with questions or concerns about this recall can contact Leica via e-mail at LMGRA@leica-microsystems.com.For questions regarding this recall call 847-405-5413. |  | 33 | 33 units | Nationwide Distribution including AZ, CA, FL, and NY. | Terminated | August 14, 2017 | 1265 | 33 | Not_Computer
Z-1697-2014 | 68284 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | Integra Newdeal SURFIX Fixed Angle Locking Screw 3.5mm x 22mm (catalog number 285322S) and 3.5mm x 30mm (catalog number 285330S). Screws are provided in different diameters and lengths and are used to anchor various sizes and shapes of bone fixation plates when treating bone fractures or reconstruction of the ankle, foot or hand. |  |  | March 28, 2014 | 2014 | Integra LifeSciences Corporation | New Deal has identified through a complaint report that some 3.5 mm Diameter x 22 mm Length Surfix screws in Lot no. F626 (Catalogue No. 285322S) have been packaged and labeled as 3.5mm Diameter x 30 mm Length Surfix screws (Catalogue No. 285330S). | Package design/selection | Integra issued an Urgent Voluntary Medical Device Recall letter dated March 28, 2014, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to review their inventory for affected product, immediately stopo using or distributing them. Customers were asked to complete the Acknowledgement Form and return by email or fax as indicated on the form. When the form is received customer service will contact customers to provide an RMA number and directions to return the affected products. Customers with questions were instructed to call 1-855-532-1723.For questions regarding this recall call 609-275-2700. |  | 417 | 417 units | Worldwide Distribution - USA including AR, CA, NC, TX and Internationally to Austria, Chile, Denmark, Finland, France, Germany, Greece, Ireland, Italy ,Lebanon, Portugal, Russia, South Africa, Spain, Switzerland, United Arab Emirates, and United Kingdom. | Terminated | April 18, 2016 | 752 | 417 | Not_Computer
Z-0718-2014 | 63071 | No | HTW | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Anspach Single Use Sterile Bone Cutting Burs and Irrigation Clips. Rx Only |  |  | September 09, 2011 | 2011 | The Anspach Effort, Inc. | Recall was initiated due to the observation of separation between clear layers of the pouch during real time aging studies. As a result, it is possible to have a breach in the sterile barrier. | Packaging change control | Recalling firm notified direct accounts by letter on 09/09/11. Direct accounts were instructed to screen their inventory immediately for product packaged with the Tyvek Pouch/Tube Configuration, complete the reply form, and return the products to the firm. |  | 1860048 | 1,805,022 cutting burs and 55,026 irrigation clips | Worldwide Distribution - USA (nationwide) and the countries of Japan, Canada, Nigeria, Guam, China, Czech Republic, Thialand, Ecuador, Lithuania, Argentina, Turkey, Cayman Island, Jamaica, India, Jordan, Germany, Brazil, Mexico, Israel, Cameroon, Switzerland, South Africa, Greece, Finland, Taiwan, Hong Kong, France, Spain, Cyprus, Australia, United Kingdom, Sweden, Bahama, Korea, Malaysia, Iceland, Portugal, Polan, Austria, Belgium, Egypt, Singapore, Netherland, Arab Emirates, Kenya, Italy, Venezula, Columbia, Dominican Republic, Chile, Panama, Honduras, Costa Rica, Philippines. | Terminated | April 15, 2016 | 1680 | 1860048 | Not_Computer
Z-1242-2014 | 67618 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Part 475620 Intramedullary Bone Saw CAM Assembly |  |  | February 26, 2014 | 2014 | Biomet, Inc. | Incorrect labeling on specific lots of the Intramedullary Bone Saw.The Saw Blade Assembly was identified as CAM Assembly. | Labeling False and Misleading | The firm, BIOMET, sent an "URGENT MEDICAL DEVICE SAFETY ALERT" letter dated February 26, 2014 via Fed Ex overnight to domestic consignees. Australia and GSCC were notified by email and also received a hard copy letter via mail. The letter described the product, problem and actions to be taken. The letter stated that this action requires the immediate location and discontinued use of the product and its return to Biomet. The Consignees were were instructed to immediately locate, discontinue use of the product and remove the identified device(s) listed; carefully follow the instructions on the FAX Back Response Form and reply within three (3) business days; complete and return Form via Fax to: 574-372-1683 prior to return of product, and use priority carrier for your shipment. If you have further distributed this product, you MUST notify hospital personnel of this action via enclosed "Dear Risk/Recall Manager Notice", and note that you are charged with the location and return of these products. Shipping address is Return Goods, 56 E. Bell Drive Warsaw, IN 46582. Questions related to the notice should be directed to 574-372-1570, Monday through Friday, 8 a.m. to 5 p.m. |  | 1 | 1 | Worldwide Distribution: US (nationwide) to states of: NY, TX, MI, and M; to countries of: Australia and The Netherlands. | Terminated | July 07, 2015 | 496 | 17 | Not_Computer
Z-0191-2015 | 69464 | No | FTL | Surgical mesh. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | BardÂż PerFix Light Plug |  |  | October 08, 2014 | 2014 | Davol, Inc., Subs. C. R. Bard, Inc. | The patient record peel-off label product code and size did not match the other labeling and actual product in the packaging for the Bard Small PerFix Light Plug. | Error in labeling | Davol Inc. (Subsidiary of C.R. Bard, Inc)sent Customer Notification letters on Friday, October 10, 2014, via FedEx next day delivery with signature confirmation to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Action required:Examine your inventory and identify any product subject to this communication. A sample of the product labeling for the lot has been enclosed to assist in product code and lot number identification. The enclosure also depicts the supplied label as well as an image of the label with the correct information.If you have any remaining inventory, a copy of this communication can be kept with the units for future medical record documentation purposes.If you have already used some of the product lot, you may amend any patient record(s) based upon the content of this communication.Please complete and return the accompanying Effectiveness Check Form attached to this letter regardless of whether or not you have any remaining units of the affected product lot.If you may have further distributed or transferred this product, please identify the respective organizations and notify them of this product communication. Your notification to these organizations may be enhanced by including a copy of this product communication letter.This product communication is being made with the knowledge of the Food and Drug Administration. We apologize for any inconvenience this may have caused your organization. Please contact the Davol Customer Service Department at 1-800-556-6275 or C.R. Bard Medical Services &amp; Support Department at 1-800-562-0027 if you have any questions. |  | 78 | 78 units | US Distribution including the states of CA, NM, OH, DE, MN, UT, MD and MO. | Terminated | December 23, 2015 | 441 | 78 | Not_Computer
Z-1696-2014 | 68124 | No | GCJ | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | BardÂż Ventralight" ST Mesh with Echo PS" 8"x10" Reorder Number: 5955810 |  |  | April 24, 2014 | 2014 | Davol, Inc., Subs. C. R. Bard, Inc. | Pouch holding sterile inflation assembly and inflation adapter may be open or have a weak seal. | Packaging process control | Davol, Inc., Subs. C. R. Bard, Inc.sent an Urgent Medical Device Recall letter dated April 28, 2014, to all affected consignees. The letter identified the product, the problem, and the action to be taken by the consignee. Consignees were instructed to examine their inventory and quarantine product subject to the recall, contact Davol Customer Service at 1-800-556-6275 or C.R. Bard's Medical Services and Support Department at 1-800-562-0027 for instructions on how to return the product to Davol, Consignees were also instructed to complete and return the recall effectiveness check form, and notify their customers if the product was further distributed.For questions regarding this recall call 1-800-556-6275. |  |  |  | Worldwide Distribution - USA (nationwide) and Internationally to AUSTRIA, BELGIUM, CANADA, CYPRUS, ESTONIA, FINLAND, FRANCE,GERMANY, GREECE, IRELAND, ITALY, LUXEMBOURG, NETHERLANDS, NORWAY, POLAND, PORTUGAL, SOUTH AFRICA, SPAIN, SWEDEN, SWITZERLAND, and UNITED KINGDOM | Terminated | December 23, 2015 | 608 |  | Not_Computer
Z-0719-2014 | 66907 | No | FGX | Manual gastroenterology-urology surgical instrument and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | TWISTER PLUS 22mm Rotatable |  |  | November 20, 2013 | 2013 | Boston Scientific Corporation | Sterility of device may be compromised due to lack of pouch seal integrity | Packaging process control | Boston Scientific issued Customer Notification letters to US accounts via Federal Express Priority mail on November 20, 2013. The notification instructs the account to check their inventory, remove any affected product from their inventory, and return it to Boston Scientific. The notification also requests the completion of a Reply Verification Tracking Form. |  |  | 0 eaches | Worldwide Distribution-USA (nationwide) and the countries of Canada, Austria, Australia, New Zealand, France, Spain, Portugal, Italy, Switzerland, Denmark, and Netherlands. | Terminated | February 24, 2015 | 461 | 5338 | Not_Computer
Z-1710-2014 | 68090 | No | NJT | Immunohistochemistry reagents and kits. | Hematology | Pathology | 510(K) Exempt | Liquid Mouse Monoclonal Antibody CD15 (NCL-L-CD15), a liquid tissue culture supernatant containing 15 mM sodium azide as a preservative. For in vitro diagnostic use. |  |  | January 10, 2014 | 2014 | Leica Microsystems, Inc. | Product does not function as intended up to the expiry date on the product labeling. There is a link between the age of the product and staining intensity. | Incorrect or no expiration date | Leica sent an Urgent Field Safety Correction Notice letter dated January 10, 2014 was sent ot all affected customers. The notice included instructions for recipients to: 1) destroy any unused or partially used NCL-L-CD15 (lots 6015073 and 6017936) reagent or confirm that the products are no longer in stock; 2) indicate on the attached Field Safety Correction Notice Acknowledgement Fom which action was taken; and, 3) contact the local Leica represetnative immediately if a replacement is required. Customers with any concerns about this notice can contact Leica via e-mail at LMGRA@leica-microsystems.com. |  | 145 | 145 units | Worldwide Distribution - US Nationwide in the states of CA, CT FL, IL, MO, MS, NC, NY, PA, and the countries of Argentina, Australia, Brazil, Colombia, Italy, Malaysia, New Zealand, Nigeria, Poland, Romania, Russia, Saudi Arabia, Serbia, South Africa, Sweden, Taiwan, Tunisia, Turkey, and United Kingdom | Open, Classified |  |  | 145 | Not_Computer
Z-1233-2014 | 67194 | No | LYC | Bone grafting material. | Dental | Dental | 510(k) | ENDOBON XENOGRAFT GRANULES 0.5 ML (Bovine bone substitute) |  |  | October 03, 2013 | 2013 | Biomet 3i, LLC | Individual pieces of a batch of Biomet's Endobon Xenograft Granules 0.5 ML (Bovine Bone substitute) may have an incorrect label with the wrong batch number. | Error in labeling | Biomet 3i sent a Medical Advisory Notice dated October 3, 2013 via mail, facsimile or telephone to all affected customers. The notice identified the product, the problem, and the action to be taken by the customers. The customers were instructed to do the following: 1) Review their inventor for any incorrectly labeled products and return them to their local Biomet distributor. 2) Provide this information to persons in their organization that uses or order these products. 3) Ensure that a copy of this letter is provided to any other organization to which the products maybe have been transferred. Customer were also instructed to respond to the attached Response Form via fax at 561-514-6316 or email to: domesticcomplaints@biomet.com. For any questions or concerns, call Biomet 3i Customer Service at 1-800-342-5454.For questions regarding this recall call 561-776-6706. |  | 704 | 704 | Worldwide Distribution - USA including CA, DC, FL, LA, MA, MO, NC, NJ, NY, TX, VA, WA and Hawaii and Internatially to Austria, Belgium, Canada, Switzerland, Germany, Spain, France, Great Britain (UK), Ireland, Mexico, Poland, Portugal, and Singapore. | Terminated | December 11, 2014 | 434 | 704 | Not_Computer
Z-0174-2015 | 69454 | No | DZE | Endosseous dental implant. | Dental | Dental | 510(k) | Straumann Bone Level Implant 04.1 mm RC, SLActiveÂż 8mm, RoxolidÂż, Loxim |  |  | September 24, 2014 | 2014 | Straumann USA, LLC | Product transfer piece was fitted with a Narrow Connect (NC) instead of a Regular Connect (RC) | Packaging process control | Straumann USA sent a Field Corrective Action Notice letter dated September 24, 2014 and delivered via United Parcel Service. The letter identified the affected product, problem and actions to be taken. Straumann is requesting the return of unused implants. Returns and Customer Confirmation form is to be returned. Contact firm 781-747- 2514. |  | 111 | 111 units | US Nationwide Distribution | Terminated | June 26, 2017 | 1006 | 111 | Not_Computer
Z-1720-2014 | 68039 | No | KTT | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | SBi RingFix Olive Wires |  |  | April 07, 2014 | 2014 | Small Bone Innovations, Inc. | Small Bone Innovations (SBI) is recalling non-sterile RingFIX Olive Wires labeled as item #451-3004 Drill tip 1.8 mm diameter 400mm olive k-wire packaged with item #451-3005 Bayonet tip 1.8mm diameter 400mm olive k-wire inside. | Packaging process control | SBi sent an Urgent Medical Device Recall dated April 7, 2014, to affected customers (sales consultants, distributors, and end users). The letter identified the affected product, problem and actions to be taken. For questions call 215-428-1791 ext 248. |  | 44 | 44 | US Nationwide Distribution in the states of IN, NC, WA, NH, and CT. | Terminated | October 20, 2014 | 196 | 44 | Not_Computer
Z-1226-2014 | 67282 | No | CIG | Bilirubin (total or direct) test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | VITROS Chemistry Products White Reference Slides, White Correction Factor Slide (WCF), DT Slides, 25 Slides, Part Number J02315; Made in USA by Ortho- Clinical Diagnostics, Inc. 100 Indigo Creek Drive, Rochester, NY 14626 --- This product is a part/ tool used when performing the correction factors adjustment on the VITROS Chemistry Systems. It does not have a 510(k) number. |  |  | January 09, 2014 | 2014 | Ortho-Clinical Diagnostics | An incorrect assay value (0.3785) was printed on the product carton for the 460 nm wavelength (filter) that may cause a delay in the ability to obtain and report TBIL, Bu, Bc, and derived tests results. The correct assay value is 0.8658. | Error in labeling | Ortho Clinical Diagnostics sent an Urgent Distributor Product Correction Notification letter dated January 9, 2014, via FedEx overnight courier to all affected customers. The letter notified users of the issue and the correct 460 nm value. Foreign affitliates were notified on January 9, 2014 via e-mail. The letter advised users to inspect their current inventory of VITROS White Correction Slides to determine if Lot 9052-0045-6358 is present. If there is remaining inventory of VITROS White correction Slides, Lot 9052-0045-6358, the remaining inventory may be discarded or the user may continue to use the product with the updated 460nm value provided. If VITROS White Correction Slides, Lot 9052-0045-6358 have been previously used to perform the reflectometer correction factors adjustment , then this procedure should be repeated using the corrected 460 nm values or an alternate lot of VITROS White Correction Slides. Customers were instructed to complete and return the Confirmation of Receipt Form and identify all customers and any other branches who were shipped the affected lot from their facility and send a copy of the customer communication. Customers were also instructed to forward the information if they have distributed the affected product outside of their facility. Customers with questions were instructed to call Customer Technical Services.For questions regarding this recall call 908-218-8776. |  | 64 | 64 units worldwide | Worldwide Distribution - USA (nationwide) Puerto Rico, and Internationally to Australia, Brazil, Canada, Chile, China, Colombia, England, France, Germany, India, Italy, Japan, Mexico, Republic of Panama, Singapore, Spain, and Venezuela. | Terminated | July 07, 2015 | 544 | 64 | Not_Computer
Z-0177-2015 | 69512 | No | DRF | Electrode recording catheter or electrode recording probe. | Cardiovascular | Cardiovascular | 510(k) | St. Jude Medical, Fixed Curve Response 6F Electrophysiology Catheter, CSL, Model 401353 |  |  | October 14, 2014 | 2014 | St Jude Medical | St. Jude Medical is performing a voluntary recall of batch 4671318, model number 401353 of the Response Fixed Curve DECA 6F 65cm CSL 2/8 Catheter. Specifically, the firm has identified that every device packaged within batch 4671318 contains the incorrect catheter (Supreme FC QUAD 5F catheter) within the packaging. To date, no serious injuries have occurred as a result of this issue | Labeling mix-ups | St. Jude Medical sent an "Urgent Medical Device Recall Notice" letter dated October 7, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Advised customers to discontinue the use of the recalled device and inform them that a St. Jude Medical sales representative will contact them in the near future to facilitate removal and reconciliation of the recalled device. For questions they can contact their St. Jude Medical Sales Representative. For further questions please call (651) 756-2295. |  | 26 | 26 | US Distribution including the states of CA, GA, NJ, NY, OR and PA. | Terminated | December 10, 2014 | 57 | 26 | Not_Computer
Z-1268-2014 | 67599 | No | NJT | Immunohistochemistry reagents and kits. | Hematology | Pathology | 510(K) Exempt | AMS URIC ACID For the In vitro quantitative determination of Uric Acid in serum |  |  | February 11, 2014 | 2014 | AMS Diagnostics, LLC | Marketing the devices outside 510(k) requirements | No Marketing Application | AMS Diagnostics sent a Recall Notification letter dated February 11, 2014 to all affected customers. The letter instructed the customers to: 1) Review their inventory for the lots listed and if they currently have them, to discard and notify the firm for a replacement. 2) If they don't have the lots listed, please check the box, sign, date and return the letter via fax at (954) 217-0046 or email to dj@amsdiagnostics.com. Customers with questions were instructed to call 954-217-0040 ext. 302.For questions regarding this recall call 954-217-0040. |  | 13 | 13 kits | Nationwide Distribution including FL, GA, MD, CA, NC, MA, SC, WY, and TX. | Open, Classified |  |  | 296 | Not_Computer
Z-0207-2015 | 69500 | No | HHW | Vaginal pessary. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | CooperSurgical MILEX PESSARY KITS-Ring w/Support Folding Pessary- #1 |  |  | September 30, 2014 | 2014 | CooperSurgical, Inc. | Incorrect expiration date printed on the MILEX PESSARY KIT packaging | Incorrect or no expiration date | CooperSurgical notified customers a Recall Notification dated September 29, 2014, via UPS . The letter identified the product problem and the action needed to be taken by the customer. CooperSurgical's will replace any of the affected and provided further instruction on the disposition of the product at the time of replacement. If you have any further questions, contact 203.601.5200.. |  | 17 | 17 units | US Distibution in states of:OH, CA, WI, WA, FL, MN, NH, RI, MA, IL, MI, IA, and AL | Terminated | May 28, 2015 | 240 | 23 | Not_Computer
Z-1270-2014 | 67607 | No | DJG | Opiate test system. | Toxicology | Toxicology | 510(k) | Single Analyte DAU Calibrator (5 mL) |  |  | February 18, 2014 | 2014 | Lin-Zhi International Inc | Customer reported that the Synchron NBUP cutoff calibrator (10 ng/mL) could not be qualified through quality control testing. The cutoff calibrator was recovering at a concentration close to the NBUP Control Level 1 (7 ng/mL). | Packaging | Lin-Zhi International sent a Recall letter dated February 18, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Letters acknowledge that Lin Zhi was notified by the customer that the product could not be qualified and that this was confirmed by Lin-Zhi internal testing.Letters request that the distributor notify all customers of the situation and ask them to destroy any remaining product from the affected lots. the root cause is identified and letters also state that a corrective action has been implemented to prevent future problems.Questions may be directed to Lin Zhi International at 408-732-3856. |  | 7200 | 120 units, 60 of each lot number | US Distribution including CA. | Terminated | April 25, 2014 | 66 | 7200 | Not_Computer
Z-1210-2014 | 67686 | No | FMQ | Protective restraint. | General Hospital | General Hospital | 510(K) Exempt | DeRoyal(R) LIMB HOLDERS, COTTON ADJUSTABLE, UNIVERSAL, REF M2017 |  |  | February 21, 2014 | 2014 | DeRoyal Industries Inc | Moldy smell in product | Packaging | DeRoyal initiated an Urgent Voluntary Recall dated February 21, 2014 to all affected customers. The letter identified the affected product, problem and actions to be taken. DeRoyal has asked that products be destroyed, and that notice of destruction be sent back. For questions contact your DeRoyal Sales Representative or DeRoyal Customer Service at 1-800-251-9864. |  | 8962 | 8,962 pairs | Worldwide Distribution - USA Nationwide in the states of FL, NC, IA, AL, IL, TX and the countries of Argentina. | Terminated | May 08, 2015 | 441 | 8962 | Not_Computer
Z-1664-2014 | 68281 | No | HMX | Manual ophthalmic surgical instrument. | Ophthalmic | Ophthalmic | 510(K) Exempt | Visitec Anterior Chamber Cannula, .40x22mm (27Gx7/8in) |  |  | May 06, 2014 | 2014 | Beaver-Visitec International Inc. | Contains a sharp tip needle when the product should contain a blunt tip cannula | Packaging process control | Beaver Visitec (BVI) sent a Product Recall Notice dated May 6, 2014, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to immediately return any un-opened product of the affected lot so that BVI can replace the product or issue credit to their company. Customers were asked to fax or email the customer response form to 866-906-4304 or claimsUS@beaver-visitec.com as acknowledgement of the letter. Customers with questions should call 781-906-7986.For questions regarding this recall call 781-906-7950. |  | 9970 | 9970 units | Nationwide Distribution including AL, CT, IN, NC, NY, TX, and VA. | Terminated | September 05, 2014 | 122 | 9970 | Not_Computer
Z-1663-2014 | 68221 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | MAGNETOM Aera magnetic resonance system. |  |  | April 24, 2014 | 2014 | Siemens Medical Solutions USA, Inc | Siemens discovered a processing error for the MAGNETOM Aera magnetic resonance system that may have resulted in some users not receiving Operator Manuals. | Labeling design | A customer safety advisory notice, dated April 24, 2014 was sent to notify direct accounts of the delivery of user manuals. |  | 5 | 5 | Distributed in the states of MN, CA, CO, TX, and IA. | Terminated | January 08, 2015 | 259 | 5 | Not_Computer
Z-0716-2014 | 67142 | No | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Locking Cortical Screw |  |  | December 19, 2013 | 2013 | Biomet, Inc. | 3.5mmx38mm Locking Cortical Screw, implants are being recalled since an investigation identifiedthat a 3.5mmx28mm Cortical Locking Screw is labeled as a 3.5mmx38mm Cortical Locking Screw. | Labeling mix-ups | URGENT MEDICAL DEVICE RECALL notices were sent to all consignees on December 19, 2013. Distributors who have affected hospitals in their regions were notified and provided with a copy of the notice that went to each hospital. Consignees are being contacted on an individual basis in an attempt to reconcile the product 100%. Fax back forms were provided to all consignees in order for Biomet to track returns. Non-responders will be contacted on an individual basis. |  | 52 | 52 | US Distribution including states of: AZ, CA, GA, IL, KY, MI, NC, NH, NY, OH, PA, SD, TX, UT, and VA. | Terminated | July 07, 2015 | 565 | 52 | Not_Computer
Z-1208-2014 | 67627 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | REF 623421 RS-ROM Noiles Rotating Hinge, Femoral Rotating Hinge Cemented (with Pin), Right , X-Small, Sterile R |  |  | February 24, 2014 | 2014 | DePuy Orthopaedics, Inc. | DePuy Orthopaedics has identified the potential for holes to develop in the inner and outer flexible pouches that form the sterile barrier for S-ROM Noiles Rotating Hinge Femur with Pin devices. The outer carton and shrink wrap are intact. A package redesign is underway to resolve this issue. S-ROM Noiles Rotating Hinge Femur with Pin devices will remain in US Distributors inventory during the pa | Packaging | DePuy Orthopaedics sent an Urgent Information - Recall Notice dated February 24, 2014, to all affected customers. The notice identified the product, reason for the recall and potential risks associated with using the recalled product. The notice includes response cards for the US Distributors to verify they have read, understood, and are adhering to the directions given in the field communication and for hospitals where consignment inventory is removed a hospital letter is available, as needed. No hospital signatures are required. Distributors are asked to send completed distributor cards to DePuy. Product related questions should be addressed to DePuy Orthopaedics, .Sales Consultant. Clinical questions should be addressed to DePuy Scientific Information Office 1-888-554-2482; recall questions to 574-371-4917. |  | 132 | 132 | Worldwide Distribution - USA (nationwide) and Internationally to Canada, Australia, Austria, China, Czech Republic, Denmark, Dubai / UAE, Estonia, France, Germany, Hungary, Iberia, India, Ireland, Israel, Italy, Malaysia, Netherlands, New Zealand, Norway, Pakistan, Poland, Portugal, Russia, Singapore, Slovenia, Spain, South Africa, Sweden, Switzerland, Turkey, United Kingdom, and Vietnam. | Terminated | March 03, 2015 | 372 | 790 | Not_Computer
Z-1748-2014 | 68081 | formulations, paraffin, all | KEO | General purpose reagent. | Pathology | Pathology | 510(K) Exempt | Tissue-Tek PARAFORM PROC/EMB/MEDIUM, FORMULA 3 - 8x1Kg. |  |  | April 16, 2014 | 2014 | Sakura Finetek USA Inc | This Lot has variations in melting temperature. | Packaging process control | Customer notifications sent via e-mail letter.The letter titled "PRODUCT RECALL" dated 04/17/2014 informed customers (of Lot 263922), informed customers of the recall by providing the following information: product name, code, lot number, reason for recall, requested actions, contact information and an "FIELD CORRECTION EFFECTIVENESS CHECK" form. |  | 115 | 115 | Nationwide Distribution. | Terminated | September 02, 2014 | 139 | 115 | Not_Computer
Z-0153-2015 | 69392 | No | DZE | Endosseous dental implant. | Dental | Dental | 510(k) | Spline Twist Implant, HA 3.25mmD X 15mmL; Part No. 2135. |  |  | October 01, 2014 | 2014 | Zimmer Dental Inc | Zimmer Dental is recalling Tapered Screw-Vent Implants because the incorrect diameter size was printed on the inner vial. | Labeling Change Control | Customer Notification letters dated 10/6/14 were sent to all customers who purchased the Tapered Screw Vent Implant and the Spline Twist Implant because the incorrect diameter size was printed on the inner vial. The letters informs the customers of the problems identified and the actions to be taken. Customers are instructed to complete the attached Business Reply Form and fax it to (574) 372-4265 or email to corporatequality.postmarket@zimmer.com. Customers with questions are instructed to contact Zimmer Dental Customer Service at (800) 854-7019 from 7am-5pm, Monday-Friday. |  | 1044 | 1,044 units | Worldwide Distribution-US (nationwide) and the countries of Poland, Canada, Belgium, Italy, France, Israel, Morocco, Spain, Costa Rica, and Japan. | Terminated | April 28, 2015 | 209 | 1044 | Not_Computer
Z-1641-2014 | 67991 | Yes | BTR | Tracheal tube. | Anesthesiology | Anesthesiology | 510(k) | Portex Low Dead Space Connector with Sideport, 3.5mm |  |  | April 10, 2014 | 2014 | Smiths Medical ASD, Inc. | One lot of 3.0mm sized Connectors, Product Reorder No. 10035-05PS, Lot No. 2553426 were placed in packages labeled as size 3.5mm. | Packaging process control | Smiths Medical sent an Urgent Medical Device Recall Notice dated April 10, 2014, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to inspect their inventory and quarantine all affected products and to complete the attached Response Form and return it to Smith Medical within 7 days of receipt of the notice. Customers with questions were instructed to contact Smiths Medical's Customer Service Department at +1-800-258-5361. |  | 100 | 100 devices | US distribution in the state of North Carolina. | Terminated | November 29, 2016 | 964 | 100 | Not_Computer
Z-1652-2014 | 67956 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Siemens Linear Accelerator (LINAC) models. |  |  | April 01, 2014 | 2014 | Siemens Medical Solutions USA, Inc | Siemens Radiation Oncology became aware that customers may be using the Siemens Linear Accelerator in combination with stereotactic accessories which have not been validated as being compatible with Siemens LINAC models. | Labeling design | Siemen sent an Urgent Medical Device Correction letter to all affected customers. The letter identified the affected product, problem and actions to be taken. For questions contact your local Elekta representative. |  | 9 | 9 | US Nationwide Distribution in the states of AL, ME, PA, IL, and CA. | Terminated | March 16, 2015 | 349 | 9 | Not_Computer
Z-1300-2014 | 67567 | No | LJT | Subcutaneous, implanted, intravascular infusion port and catheter. | General Hospital | General Hospital | 510(k) | PowerPort Slim Implantable Port, indicated for patient therapies requiring repeated access to the vascular system. Product Codes: 1716000, 1716070 |  |  | February 27, 2014 | 2014 | Bard Access Systems | Labeling discrepancy.The label states that the implantable port contains silicone suture plugs, however approximately 23% of the units contained in the affected lots have a PowerPort Slim titanium implantable port without silicone suture plugs. | Packaging | Notified by Urgent Customer Notification Letter date February 13, 2014, customers were asked to immediately examine their inventory and identify any product subject to this notification. In addition, if they have further distributed this product, please identify customers and notify them at once of this notification. Customers may include a copy of this letter in your notification to customers. Customers may choose to use the involved product with the understanding that the kit may contain a PowerPortÂż Slim implantable port without suture plugs, or may return the applicable product for replacement or credit. |  | 777 | 777 | US Nationwide distribution including WA, KY, IL, and IA. | Terminated | September 05, 2014 | 190 | 777 | Not_Computer
Z-0114-2015 | 69256 | No | CAZ | Anesthesia conduction kit. | Anesthesiology | Anesthesiology | 510(k) | ASK-05501-AFH1, Lot Number 23F13K0969 Anesthesia Conduction Catheter, Epidural Anesthesia Kit |  |  | September 15, 2014 | 2014 | Arrow International Inc | Arrow International, Inc. is notifying customers that lot 23F13K0969 of product code ASK-05501-AFH1 contains an incorrect component. The lidstock indicates that the kit contains one Syringe, 10mL, Luer-Slip Loss of Resistance, Glass but contains one Syringe, 10mL, Luer-Lock Loss of Resistance, Glass. | Labeling mix-ups | Arrow International sent a Medical Device Advisory Notification dated September 15, 2014, to affected customers notifying them that lot 23F13K0969 of product code ASK-05501-AFH1 contains an incorrect component, and requesting they provide this Advisory Notice to all those who need to be aware of it within your organization. Questions were directed to local sales representative or Customer Service at 1-866-246-6990. |  | 300 | 300 | US Distribution in Utah only. | Terminated | December 18, 2015 | 459 | 300 | Not_Computer
Z-0232-2015 | 69579 | No | JWY | Antimicrobial susceptibility test powder. | Microbiology | Microbiology | 510(k) | EtestÂż Ceftaroline (CPT32) US B30. |  |  | October 16, 2014 | 2014 | Biomerieux | The products have a wrong expiration date on their labeling, 5 years instead of 2 years. | Labeling Change Control | Consignees of bioMÂżrieux, impacted by the identified issue, were notified of the FSCA and corrective action (FSCA notification) electronically on October 16th, 2014. Additionally, the bioMÂżrieux United States subsidiary notified US consignees of the FSCA notification via FedEx on October 23rd, 2014. |  | 652 | 652 devices | Worldwide Distribution-US (nationwide) including DC and the states of AL, AK, AR, AZ, CA, CO, DE, FL, GA, IA, IN, KY, MA, MD, MI, MO, NC, NE, NY, OH, OK, PA, SC, TN, TX, UT, VA, WA, WI, and the countries of Australia, Austria, China, Czech Republic, France, Germany, India, Italy, Japan, Netherlands, Poland, Portugal, Switzerland, Sweden, Turkey, United Kingdom, Egypt, Pakistan, Slovenia and Uruguay. | Open, Classified |  |  | 652 | Not_Computer
Z-2024-2014 | 68607 | No | GAR | Nonabsorbable polyamide surgical suture. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Dafilon Suture Black 10/0 (0.2) 15 cm DRM4 NS, non-sterile, non-absorbable polyamide Surgical Suture, additional label affixed on the box states the product is sterile, packaged in a cassette holding 12 sutures |  |  | June 13, 2014 | 2014 | Aesculap, Inc. | Aesculap Inc. (AIC (USA)) has initiated a voluntary recall of Dafilon Black 10/0 (0.2) 15 cm DRm4 non-sterile non-absorbable polyamide Surgical Suture due to a labeling issue. Box labels do state non-sterile; however there is an additional label that is affixed on the box that states the product is sterile. | Labeling design | A correction and removal notification, dated June 13, 2014, was sent to end users, sales reps, and distributors which described the product, problem, and actions to be taken. |  | 11 | 11 | Distributed in the states of IL, NY, GA, and RI. | Terminated | October 07, 2014 | 116 | 11 | Not_Computer
Z-2025-2014 | 68608 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | SmartSite Low Sorbing Infusion Set, Model Codes: 10014855A, 10015861A, 10061661A |  |  | June 17, 2014 | 2014 | CareFusion 303, Inc. | CareFusion is recalling the SmartSite Low Sorbing Infusion Set because they were assembled with a drip chamber that consist of a 15 micron filter. The Directions For Use did not indicate that the drip chamber has a 15 micron filter and it does not provide the instruction that the set should not be used with blood administration due to the potential for occluding the 15 micron filter in the drip c | Labeling Change Control | A recall letter is dated 6/25/14 and will be sent out to customers who purchased the SmartSite Low Sorbing Infusion Set because they were assembled with a drip chamber that consist of a 15 micron filter. The Directions For Use did not indicate that the drip chamber has a 15 micron filter and it does not provide the instruction that the set should not be used with blood administration due to the potential for occluding the 15 micron filter in the drip chamber. The letter informs the customers of the problems identified, potential risk, and actions to be taken. Customers are instructed to promptly complete and return the enclosed mandatory Customer Response Card to acknowledge the receipt of this communication and to expedite the corrective action process. |  | 5600 | 5,600 units | Distributed in the states of FL, NM, WA, WI, and TX. | Open, Classified |  |  | 5600 | Not_Computer
Z-1130-2014 | 67468 | No | LPH | Hip joint metal/polymer/metal semi-constrained porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | FMP X-alt Acetabular Liner |  |  | February 07, 2014 | 2014 | Djo Surgical | Packaging error -- two different types and sizes of acetabular liners used during hip replacement surgery were mistakenly packaged into each other's outer boxes after sterilization, resulting in the products being mislabeled. The units labeled as Model/Catalog Number 931-28-248 Lot/Serial Number 685F1034 on the outside and containing Model/Catalog Number 932-36-252, Lot# 728F1089 on the inside (1 | Packaging process control | DJO Global sent an Urgent Field Safety Notice dated February 7, 2014, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. All 10 devices with the outer label of 931-28-248, Lot#685F1034 were found in-house and were quarantined. Theagencies who have the affected product (devices labeled as932-36-252, Lot# 728F1089 containing 931-28-248, Lot#685F1034) were contacted to quarantine their devices. Of the affected devices, 2 of the lot of 10 have already been returned via product complaint.Customers were instructed to pass on the notice to all those who need to be aware within their organization or to any organization where the potentially affected devices have been transferred. Customers were instructed to place a replacement order and receive an RMA number to return all affected devices.. Customers with questions were instructed to call 512-834-6255. |  | 931282480 | 10 units of Model/Catalog Number 931-28-248 Lot/Serial Number(s) 685F1034 labeled asModel/Catalog Number 932-36-252, Lot# 728F1089 | Nationwide Distribution - USA including TX, LA, FL, CA, and AZ. | Terminated | October 01, 2014 | 236 | 931282480 | Not_Computer
Z-1122-2014 | 67545 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Triton FP Infusion Pump, Part Number 400000, packaged in boxes. |  |  | February 17, 2014 | 2014 | WalkMed Infusion, LLC | WalkMed Infusion is initiaing a recall on their Triton FP Infusion Pump, Part Number 400000 due to the label on the right side of the pump containing an incorrectinstruction. | Labeling design | WalkMed Infusion sent an Important Medical Device Recall letter dated February 14, 2014, to all affected customers. The letter identified the product the problem and the action customers need to take.Our records Indicate that you have purchased product affected by this recall (see the attached list).Our sales representative will contact you to determine a plan for label replacement for your units. Replacement of the label will include a simple process of removing the old label (PN 300063), replacing it with a corrected label(PN 400033) and documenting completion of the replacement. This action can be completed by you,your customers,our sales representative, or by returning the infusion pump to WalkMed Infusion.For further questions please call (720) 351-4934 or (720) 351-4965.Product will be corrected in the field when possible and supported by the customer. Otherwise, WalkMed Infusion employees with travel to the customer and complete the correction activities. |  | 797 | 797 | US Distribution including the states of IL, MA, KS, and, NJ. | Terminated | October 06, 2014 | 231 | 797 | Not_Computer
Z-0229-2015 | 69461 | No | NBC | B-type natriuretic peptide test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | ADVIA Centaur BNP Assay for in vitro diagnostic use in the quantitative determination of B-type Natriuretic Peptide (BNP) in human plasma using the ADVIA Centaur, Centaur XP, Centaur CP Systems; Catalog Numbers: |  |  | October 06, 2014 | 2014 | Siemens Healthcare Diagnostics, Inc | Siemens Healthcare Diagnostics has confirmed that approximately 2000 ReadyPacks from the ADVIA CentaurÂż Systems BNP reagent contains an ADVIA Centaur Systems FT4 human-readable side label. | Labeling mix-ups | A Customer Notification (CN) was sent to all affected Siemens Healthcare Diagnostics customers in the United States on October 03, 2014 to be delivered to customers on October 06, 2014 via Federal Express. An Urgent Field Safety Notice (UFSN) was sent out for distribution to all customers outside the United States on October 06, 2014. These notices inform the customer of the potential to receive an ADVIA Centaur BNP ReadyPack with a FT4 human-readable side label. Siemens instructs customers to inspect their inventory and discard any BNP Readypacks with a FT4 side label. In addition, Siemens informs customers the end labels of the affected ReadyPacks correctly identify the pack as BNP and Siemens has confirmed the reagent within these packs is BNP Reagent. Therefore valid BNP results will be generated when calibrating and running quality control or patient samples. Fax-back forms will be used to confirm that the customers received the notification. For those customers that do not return a fax-back form, a Siemens representative will follow up with those customers to ensure they received the Customer Notification and completed the required actions. |  | 7078 | 100 test - 4,600 kits; 500 test - 1,878 kits | Worldwide Distribution - US Nationwide including Puerto Rico; and countries of: Canada, Austria, Belgium, Czech Republic, Denmark, Estonia, Finland, France, Fren.Polynesia, Germany, Great Britain, Greece, Guadeloupe, Italy, Latvia, Lithuania, Netherlands, Norway, Poland, Portugal, Reunion, Russian Fed., Slovakia, Slovenia, Spain, Switzerland, Turkey, Unit.Arab Emir., AR, AU, BR, CL, CN, CR, IN, JP, KR, LK, MY, SG, TW, UY, VN, and ZA. | Terminated | July 22, 2015 | 289 | 7078 | Not_Computer
Z-1984-2014 | 68383 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | Synthes B37 Replacement Screws. M3.5 Screw for insertion handle guide/aiming arm, part numbers 3.113.025 and 03.113.026. Surgical instrument for use with Low Bend Medial Distal Tibia Plate. Orthopedic. |  |  | May 15, 2014 | 2014 | Synthes, Inc. | All lots of B37 Replacement Screws are being recalled due to potential mislabeling where the package may not contain the screws identified on the label. | Labeling design | An urgent medical device recall notice, dated May 29, 2014, was sent to end users and sales consultants to advise them of the issue and provide instructions for returning affected devices. |  | 44 | 44 | Distributed US nationwide, in DC and the states of KY, PA, IL, WA, AZ, CA, OK, FL, and KS. | Terminated | December 18, 2015 | 582 | 44 | Not_Computer
Z-0424-2015 | 69366 | No | HWA | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Trabecular Metal Tibial Impactor Replacement Pads. |  |  | October 14, 2014 | 2014 | Zimmer, Inc. | To provide clarifying instructions relating to the assembly/disassembly of the Trabecular Metal Tibial Impactor with/from TM tibialImpactor Pads due to complaints regarding Tibal Impactor pad fracture. | Labeling design | On 10/14/2014, URGENT MEDICAL DEVICE CORRECTION notifications were sent to the affected distributors and hospital Risk Managers. All distributors were notified via electronic mail, and all hospital Risk Managers and distributors with affected inventory were also notified via courier. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. For questions, call the customer call center at 1-877-946-2761. |  | 4961 | 4961 units | Worldwide Distribution - US Nationwide in the state of AL, CA, CO, FL, GA, HI, IA, ID, IL, IN, KS, KY, MA, MD, ME, MI, MN, MO, ND, NH, NJ, NV, NY, OH, OR, PA, SC, TX, VA, WI, NC, TN, OK, AZ, UT, AK, PR, DC. and the countries of: Germany, China, Canada, Australia, Japan, United Arab Emirates, Austria, Belgium, Switzerland, Germany, Denmark, Spain, Finland, France, United Kingdom, Greece, Iceland, Italy, Malta, Netherlands, Norway, Poland, Russian Fed., Saudi Arabia, Sweden, Slovakia, and South Africa. | Terminated | June 09, 2015 | 238 | 6126 | Not_Computer
Z-1614-2014 | 68161 | No | DYB | Catheter introducer. | Cardiovascular | Cardiovascular | 510(k) | Coaxial Micro-Introducer Kit. |  |  | April 28, 2014 | 2014 | Greatbatch Medical | Greatbatch identified that the documentation does not support the five-year shelf life of the Bard Access Division Coaxial Micro-Introducer Set. | Incorrect or no expiration date | Greatbatch sent an "Urgent Medical Device Recall" letter dated April 25, 2014 to affected customers via email and Fedex. The letter described the Reason for Recall, Risk to Health, and Instructions to Customers. Requested consignees to immediately place affected devices in quarantine and to complete and return the Field Correction Action Response Form. For questions they should contact their local Greatbatch representative or 1-763-951-8235, or e-mail FieldActionCenter@Greatbatch.com. |  | 6642 | 6,642 sets (66,420 devices) | US Nationwide Distribution in the state of UT. | Terminated | May 24, 2016 | 757 | 6642 | Not_Computer
Z-2002-2014 | 68415 | No | FRC | Sterilization process indicator. | General Hospital | General Hospital | 510(k) | 3M Attest Auto-reader. Model numbers 390, 390G,and 490. |  |  | June 09, 2014 | 2014 | 3M Company - Health Care Business | Labeling on units shipped prior to May 16th, 2014 did not contain the statement "This product contains dry natural rubber" as required by the United States Food and Drug Administration (FDA) for medical devices. The non-slip pads at the bottom of the unit contain dry natural rubber. One incident of an allergic reaction in a sensitized individual has been reported. | Labeling Change Control | Consignees were sent on 6/9/2014 a 3M "Urgent Medical Devices Correction" letter dated May 21, 2014. The letter described the problem and the product involved in the recall. The letter described what action is 3M taking and what action was required of the consignees. Consignees were sent an adhesive label and a response card. They requested consignees to adhere the label to a portion of the plastic covering each unit in a location where it is visible to the end-user. After that, they are requested to fill out and return the response card. |  | 5343 | 5343 (2621 USA, 2722 OUS) | Worldwide Distribution-USA (nationwide) including DC and Puerto Rico and the countries of AUSTRALIA, CANADA, CHILE , COLOMBIA, COSTA RICA, SOUTH AFRICA, BRAZIL, GERMANY, ECUADOR, DUBAI, JAPAN, HONG KONG , CHINA, KOREA, MALAYSIA, PANAMA, SINGAPORE, VATICAN, and THAILAND. | Terminated | June 03, 2015 | 359 | 5343 | Not_Computer
Z-1132-2014 | 67444 | No | NPJ | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Restoris Multicompartmental Knee (MCK) System is an implant system designed to be used with MAKO's Robotic Arm Interactive Orthopedic System (RIO). It is composed of a unicompartmental implant system (Restoris MCK Uni) and a patellofemoral implant system (Restoris MCK PF). The Restoris MCK Uni is designed for use when load bearing ROM is expected to be less than or equal to 155 degrees. In Restoris MCK combination where multi-compartmental areas are being treated, the Restoris MCK components were designed with 3 mm of gap between the components to ensure that the components do not interfere. |  |  | January 09, 2014 | 2014 | Mako Surgical Corporation | The product may be mislabeled and could result in the incorrect implant being used. | Error in labeling | MAKO Surgical Corporation sent an Important Medical Device Recall Notice letter dated January 9, 2014 to all affected customers. The letter identified the product, the problem, and the actions to be taken by the customer. The letter informed the customers that their MAKOplasty Sales Specialist will coordinate with them the return of the product that will be replaced at no charge. Customers were asked to sign and return the Acknowledgement form to MAKO Surgical using one of the following methods: Email, Fax or Mail. For questions, customers were instructed to call (954) 628-1721. For questions regarding this recall call 954-628-0607. |  | 41 | 41 units | Nationwide distribution including AZ, FL, HI, OH, MI, NC NH, WA, TX, WI, OK, CA, RI, LA, ID, IL, IN, and TN. | Terminated | August 14, 2014 | 217 | 41 | Not_Computer
Z-2018-2014 | 68244 | Yes | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | TerumoÂż Advanced Perfusion System 1 |  |  | June 17, 2014 | 2014 | Terumo Cardiovascular Systems Corporation | An evaluation of the APS1 Operator Manual found that existing instructions lack clarity and accuracy related to specific items. | Labeling Change Control | On June 17, 2014, TCVS sent URGENT MEDICAL DEVICE RECALL notification to consignees. Notification included reason for recall, potential hazard, correction, affected population, and customer instructions. Questions concerning the recall are directed to Terumo CVS Customer Service @ 800-521-2818 (M-F, 8am-6pm) |  | 1713 | 1713 units | Worldwide Distribution-USA (nationwide) including DC and the states of AL, AR, AZ, CA, CO, CT, DE, FL, GA, HI, IA, ID, IL, IN, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NJ, NM, NV, NY, OH, OK, OR, PA, SC, TN, TX, UT, VA, VT, WA, WI, and WV, and the countries of AUSTRALIA, BELGIUM, CANADA, CHILE, Hong Kong, Indonesia, Japan, Malaysia, Mexico, Philippines, Singapore, South Korea, Taiwan, Thailand, UNITED ARAB EMIRATES (UAE), and Vietnam. | Terminated | September 22, 2014 | 97 | 1713 | Not_Computer
Z-0433-2015 | 69638 | No | NAY | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | 8MM,MARYLAND BIPOLAR FORCEPS; |  |  | November 03, 2014 | 2014 | Intuitive Surgical, Inc. | Deviations in reprocessing steps from those stated in the reprocessing instructions can cause surface degradation of the housing and/or accelerate mechanical wear of the instrument. | Labeling design | Urgent Medical Device Recall letters were sent to affected sites on November 3, 2014. Letters inform customers of the reason for recall, risk to health and the affected instrument. A table of all of the affected IS4000 instruments affected is attached to the letter indicating the part number and product name for each. Letters also inform customers that all instruments shipped since July 2014 are manufactured with a new material and not subject to recall.Customers are to take the following actions:1. Ensure that all affected personnel are aware of the information.2. Identify affected product in their possession3. Contact ISI rep for replacement.4. Complete and return the attached Acknowledgement From.Until replacement instruments are received, the old instruments may be used provided they do not show signs of degradation and that the reprocessing instructions are followed. Questions to 800-876-1310 Option 3 ( 6 am to 5 pm PST) or ics@intusurg.com. |  | 59 | 5-9 units - total all devices | US Distribution -- FL, TX, NY, NV, OR, IN, UT, and MA. | Terminated | February 03, 2015 | 92 | 59 | Not_Computer
Z-1611-2014 | 67932 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | DPM 6 Monitor in use with a Multi Parameter Module (MPM). Patient monitor used for monitoring, displaying, reviewing, storing, and alarming of multiple physiological parameters including ECG, arrhythmia detection, ST segment analysis, heart rate (HR), respiration rate (RESP), temperature, pulse oxygen saturation (SpO2), pulse rate (PR), non-invasive blood pressure (NIBP), invasive blood pressure (IBP), carbon dioxide (CO2), anesthetic gas (AG), impedance cardiograph (ICG), bispectral index (BIS), cardiac output (C.O.) and respiration mechanics (RM) |  |  | March 05, 2014 | 2014 | Mindray DS USA, Inc. dba Mindray North America | The MPM included a cardiac ST/Arrhythmia analysis feature which is not cleared for marketing. | No Marketing Application | Mindray DS Usa Inc. contacted their sole customer via phone on March 5, 2014. The affected products will receive software upgrade. |  | 4 | 4 units | US Nationwide Distribution in the state of OH | Terminated | April 27, 2015 | 418 | 4 | Not_Computer
Z-2017-2014 | 68493 | No | BTT | Respiratory gas humidifier. | Anesthesiology | Anesthesiology | 510(k) | Humidifier Adaptor, 040 Shelf Pak, French, Teleflex Medical, respiratory gas humidifier adaptor. |  |  | June 05, 2014 | 2014 | Teleflex Medical | The seals of the adaptor packaging may be creased which may potentially affect packaging integrity. | Packaging change control | Consignees were notified by letter on 06/06/2014. |  | 7758650 | 7,758,650 ea. (total for all types listed) | Worldwide Distribution-USA (nationwide) and the countries of Belgium, Canada, China, Germany, Great Britain, France, India, Malaysia, Mexico, Philippines, Singapore, Thailand and Zambia. | Terminated | July 02, 2015 | 392 | 7758650 | Not_Computer
Z-0100-2015 | 69226 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | GYRUS ACMI, REF 3345 -Everest BiCOAG MOLly Forceps 5mm/45cm |  |  | August 27, 2014 | 2014 | Gyrus Medical, Inc | Due to an anomaly in the packaging process, it is possible that a defective seal could be present that may allow a breach of the package's sterile barrier and may compromise the sterility of the product. | Packaging process control | Olympus sent a Medical Device Recall letter dated August 26, 2014, via FEDEX to all affected consignees. The letter described the issue and the product involved in the recall. It also described the action that the consignee needs to take. Consignees were instructed to cease further use of the device, quarantine it and return it to the manufacturer.Consigness were instructed to complete and return by fax (484-896-7128) the questionnaire enclosed. If the product was further distributed consignees should notify their customers of the product recall. For additional information they can call 484-896-5688.For questions regarding this recall call 484-896-5688. |  | 5560 | 5520 US, 40 OUS | Worldwide Distribution - USA including US: AL, AZ, CA, FL, GA, HI, IL, IN, IA, KY, MA,MI,MS, NJ, NY, NC, OH, PA, SC, TN, UT, and WA.Internationally to ITALY, SPAIN, UNITED KINGDOM, and CHINA. | Terminated | September 15, 2015 | 384 | 5560 | Not_Computer
Z-0111-2015 | 69420 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Symmetric Total Knee Femoral Components |  |  | September 30, 2014 | 2014 | Signal Medical Corporation | Packaging of a device was compromised | Packaging | Signal Medical sent a Notice letter dated September 30, 2014 to customers. Customers were instructed to carefully inspect devices to ensure package integrity before using. If packaging is compromised it should not be used. Return the affected device and packaging to Signal Medical Corporation. Replacement devices with new packaging configuration is shipped to each location with the old packaging configuration in the event that the packaging is found to be compromised. For questions contact Signal Medical Corp. at 1-800-246-6324. |  | 41 | 41 | Worldwide Distribution - US (Nationwide) in the state of: MO and internationally to: Panama. | Terminated | March 14, 2016 | 531 | 41 | Not_Computer
Z-1946-2014 | 68592 | No | MQK | Assisted reproduction labware. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | Name: Pasteur Pipets, used by the embryologist in the IVF Laboratory. |  |  | June 20, 2014 | 2014 | ORIGIO, Inc | Identification of packaging breach of the nylon pouch sterilebarrier for the 3-pack configuration | Package design/selection | Field Safety Corrective Action / Urgent Product Recall letter dated June, 2014. |  | 134086680 | Total number of Boxes Distributed 32,760. US 4,093 / Canada 1,156 / EU &amp; Other 27,511 | Worldwide Distribution: US Nationwide, Canada, Australia, Barbados, Chile, China, Colombia, Denmark, Ecuador, France, Germany, Great Britain, Greece, India, Italy, Japan, Jordan, Malaysia, Netherlands, Peru, Russia, Singapore, South Africa, Spain, Taiwan, Turkey, and United Arab Emirates. | Terminated | July 20, 2015 | 395 | 134086680 | Not_Computer
Z-1129-2014 | 67421 | No | GEY | Surgical instrument motors and accessories/attachments. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Punch Hair Matic |  |  | February 18, 2013 | 2013 | MEDICAMAT S.A. | Medicamat is initiating a field correction due to misbranded marketing literature for the Punch Hair Matic SAFER. | Labeling Change Control | MEDICAMAT sent an Advisory Notice dated May 7, 2013, to the affected customer. The letter identified the product the problem and the action needed to be taken by the customer.The notification advises customers of the issue (misbranded brochure and website). The letter advises the distributor to delete the misbranded language from their communication, marketing, or other information provided to clients. They are also advised to ask their clients change their marketing if necessary and provided corrective labeling.For further questions please call +33146562404.. |  | 71 | 71 units | US Distribution including TX. | Terminated | March 10, 2014 | 385 | 71 | Not_Computer
Z-1133-2014 | 67475 | No | KZP | Bracket adhesive resin and tooth conditioner. | Dental | Dental | 510(k) | BencoDental tartar &amp; stain remover ultrasonic cleaning solution |  |  | February 10, 2014 | 2014 | Heraeus Kulzer, LLC. | As the result of a complaint to Heraeus Kulzer's customer service department, Heraues Kulzer is recalling Benco Tartar and Stain Solution part #1010-451, lot #13121434 because it has incorrect labeling on the immediate container. While the outer carton correctly states "Benco Tartar and Stain", the immediate gallon container in which the solution is contained incorrectly states "Benco Temporary C | Labeling mix-ups | Heraeus Kulzer, notified their one consignee of the recall February 4, 2014 via phone and then by official electronic notification. On February 7, 2014, an Important Communication Product Recall letter was also mailed to Benco direct accounts. On 2/10/14, Heraeus Kulzer also called Benco customers notifying them of the recall. The notifications explained the issue, identified the affected product and actions to be taken. For questions call 574-293-5444 |  | 812 | 812 units (gallons) | USA Nationwide Distribution in the states AL, AR, AZ, CA, CO, FL, GA, IL, IN, KS, KY, LA, MI, MN, MO, MS, NC, ND, NE, NM, NV, OH, OK, PA, SC, TN, TX, UT, VA, WA, WI, WV | Terminated | June 18, 2014 | 128 | 812 | Not_Computer
Z-0093-2015 | 69380 | No | MWI | Cardiac monitor (including cardiotachometer and rate alarm). | Cardiovascular | Cardiovascular | 510(k) | McKesson Cardiology Hemo. |  |  | June 25, 2014 | 2014 | Mckesson Information Solutions LLC | McKesson Cardiology Hemo front-end parameter boxes must undergo annualmaintenance testing to establish the proper functioning of the parameter boxes,and to identify in a timely manner whether a parameter box is out of calibration.McKesson has identified that its communication to customers regarding therequired maintenance tests was insufficient. While each unit has an attachedsticker i | Labeling design | Consignees were notified via Field Safety Notice on 06/25/2014. |  | 625 | 625 devices | Worldwide Distribution-USA (nationwide) and the countries of Australia, Botswana, Canada, Germany, New Zealand, Netherlands, Pakistan, Portugal, South Africa, Turkey and The United Kingdom. | Open, Classified |  |  | 625 | Not_Computer
Z-2034-2014 | 68390 | No | LPH | Hip joint metal/polymer/metal semi-constrained porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Zimmer M/L Taper Reduced Neck Standard &amp; Extended Offset, HIP JOINT PROSTHESIS, Sizes:4, 5, 6, 7.5, 9, 10, 11 &amp; 12.5. |  |  | May 23, 2014 | 2014 | Zimmer, Inc. | Black residue observed in the sterile pouch of the M/L Taper and Kinectiv stems. | Packaging | Zimmer sent an Medical Device Recall Letter on June 4, 2014, to their Distributors. On June 9, 2014, Zimmer sent Medical Device Recall Letters to the physicians. The letter identified the product the problem and the action needed to be taken by the customer.Customers were instructed:1. Locate all products from the affected lots and quarantine them immediately, if not previously done so per the May 23, 2014 Notice of Suspension of Loaner Pool. 2. Confirm the physical count of all affected product in loaner pool in your territory and complete the Confirmation of Loaner Pool Inventory and Return Certification Form (Attachment 2), and email a completed copy of Attachment 2 to corporatequality.postmarket@zimmer.com.3. Return the recalled product along with the completed Confirmation of Loaner Pool Inventory and Return Certification Form (Attachment 2). 4. Identify additional customers that may have used this product in the past 12 months by completing the spreadsheet template provided to you via email or confirm that Zimmer has identified all current customers in your territory.5. If after reviewing this notification you have further questions or concerns please call the customer call center at 1-877-946-2761 between 8:00 am and 5:00pm EST. |  | 852 | 852 units | Worldwide Distribution - USA (nationwide) and the countries of Australia, Canada, China, Europe, India, Japan, Korea and Singapore. | Terminated | March 06, 2015 | 287 | 2654 | Not_Computer
Z-0865-2014 | 64795 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | The Centricity Laboratory System is intended to be an information system designed to support the clinical and administrative activities associated with the provision and utilization of clinical laboratory services and facilities, e.g., the storing and delivering of analytical results. It is a specially designed data program application (software), which is supplied for installation in existing mainframe or decentralized computers or a computer network. This software system will not be the software directly used to run, steer, or control any specific laboratory analyzer or equipment. It is intended for use in laboratory facilities, including a central laboratory or in a multiple laboratory environment servicing satellite laboratories, reference laboratories, clinics, etc. Centricity Laboratory System is not intended for direct patient contact. |  |  | November 20, 2012 | 2012 | GE Healthcare It | GE Healthcare is aware of a potential safety issue associated with the use of GE Centricity Laboratory Core Lab calculated results in two scenarios. In Scenario 1: Calculated results greater than 6 digits plus a decimal are reported as 0 instead of TOO BIG when significant figures are not defined for the item. In Scenario 2: Calculated results greater than 6 digits plus a decimal are trunc | Software design | An URGENT MEDICAL DEVICE CORRECTION letter date March 12, 2013 was sent to customers. The letter requests that customers review all calculated results to determine if either of the scenarios could occur at their site. If a customer determines that either of the scenarios proposed in the letter could occur at their site, the customer can contact their GE service representative to request or install the software update. Customers with questions can contact the GE Healthcare IT Helpdesk at 1-847-589-8496 or tgg.nahelpdesk@med.ge.com. |  | 7 | 7 units | Worldwide Distribution - US (Nationwide) to states of: MN, GA, and CA; and International to: England. | Open, Classified |  |  | 7 | Computer
Z-0810-2014 | 66996 | No | MWI | Cardiac monitor (including cardiotachometer and rate alarm). | Cardiovascular | Cardiovascular | 510(k) | Medrad Veris MR Monitor units |  |  | November 21, 2013 | 2013 | Medrad Mr Inc | The main board, P/N 301641, installed in some Medrad Veris MR Monitor units may be faulty and could lead to unexpected shutdown of the system while in use, resulting in the loss of vital signs information from the monitor. | Device Design | The firm, Bayer HealthCare, sent a "MEDICAL DEVICE FIELD CORRECTIVE ACTION" letter dated November 21, 2013 via certified mail to its users/customers. The letter described the product, problem and actions to be taken. The firm stated that, "You may choose to continue to use your Veris units but you MUST be vigilant about this potential issue. Should a shutdown occur, the system can be quickly restarted by pressing the POWER button. If thisproblem does occur, we ask that you report this immediately to Bayer Customer Support at 1-877-229-3767." The customers were instructed to verify the serial number on your Medrad Veris unit(s). If this problem does occur, report this immediately to Bayer Customer Support at 1-877-229-3767. The customers were also instructed to fill out and fax back the attached form. Bayer Service Representative will then be in touch to schedule an appointment to come and replace your main board, at no cost to you.Bayer also informed their customers that based on the maximum inventory that can be produced by the supplier, they anticipate that this Field Corrective Action will take up to six months.If you have questions, please contact our Customer Support team at 1-877-229-3767. |  | 922 | 922 units | Worldwide Distribution: US (nationwide) and Internationally to: Canada. | Terminated | March 30, 2015 | 494 | 922 | Computer
Z-0085-2015 | 69181 | No | KPE | I.V. container. | General Hospital | General Hospital | 510(k) | Baxter 2B8013 Empty INTRA VIA Container with PVC Ports (Sterile, nonpyrogenic fluid path) ,500 ml. Admixture Products and Accessories |  |  | September 16, 2014 | 2014 | Baxter Healthcare Corp | Particulate matter found inside the fluid path. | Equipment maintenance | Baxter issued a letter to their consignees stating the following:Baxter is requesting that you take the following actions:Action to be taken if product was purchased directly from Baxter.1. Locate and remove all affected product from all locations in your facility(the product code and lot numbers can be found on the individualproduct package or shipping carton).2. Contact Baxter Healthcare Center for Service to arrange for return andcredit. The Center for Service can be reached at 888-229-0001between the hours of 7:00 am and 6:00 pm Central Time. Please haveyour Baxter eight-digit ship-to account number ready when calling.3. Complete the enclosed customer reply form , and return it to Baxter byeither tax or scanned e-mail.4. If you distributed any of the affected products to other facilities, pleaseconduct a recall with your end-user customers in accordance with yourcustomary procedures. Baxter distributed product 2B8011 lotU R 13D 15112 to customers between 4/26/2013 and 6/20/2013. Baxterdistributed product 2B8013 lot UR13K1 4095 to customers between 11 /27/2013 and 3/ 10/2014.Action to be taken if product was received from an alternate source.1. Locate and remove all of the affected products from all locations inyour facility (the product code and lot numbers can be found onthe individual product package or shipping carton).2. Contact Baxter Healthcare Center for Service to arrange for returnand credit. The Center for Service can be reached at 888-229-0001 between the hours of 7:00 am and 6:00 pm Central Time.Please have your ship-to account number ready when calling.3. Follow your supplier's reply and recall process. Please do not return the customer reply form to Baxter. |  | 20448 | 20448 | Nationwide Distribution including Puerto Rico, and the country of Canada. | Terminated | June 21, 2017 | 1009 | 114672 | Computer
Z-0089-2015 | 69262 | No | JXG | Central nervous system fluid shunt and components. | Neurology | Neurology | 510(k) | Codman EDS 3 Drainage System with Ventricular Catheter (External Drainage System) |  |  | September 15, 2014 | 2014 | Codman &amp; Shurtleff, Inc. | tubing within the system that drains CSF may leak or disconnect from the joints and may result in over- or under-drainage of CSF from the ventricular system or introduction of air into the ventricular system (pneumocephalus) | Device Design | Codman Neuro sent an Urgent Notice dated September 12, 2014, to all affected customers via overnight mailing and OUS customers notified via email the week of September 15, 2014.Accounts were requested to isolate all inventory of the affected product and return to Stericycle at the following address:Stericycle 2670 Executive Drive, Suite A Attn: Event #8894 Indianapolis, IN 46241For patients currently being managed with the EDS 3 System, the system should be replaced immediately. In the case where no substitute drainage system is immediately available, the EDS 3 System may continue to be used until an alternative product can be obtained.If you have any questions or concerns regarding this notification, please contact your local Codman Neuro Representative or Scientific and Medical Affairs at SciMedAffairs@its.jnj.com or (866) 685-7325.For questions regarding this recall call 508-880-8000. |  | 66120 | 66,120 units | Worldwide Distribution - USA (nationwide) and Internationally to:United Arab EmiratesArgentinaAustriaAustraliaArubaBelgiumBahrainBoliviaBrazilBahamasCanadaSwitzerlandChilePeople's Republic of ChinaColumbiaCosta RicaCyprusCzech RepublicGermanyAlgeriaEstoniaEgyptSpainFinlandFranceGabonGreat BritainFrench GuianaGreeceGuatemalaHong KongCroatiaHungaryIrelandIsraelIndiaIranItalyJordanKenyaKuwaitKazakhstanLebanonLithuaniaLuxembourgLatviaMauritiusMexicoMalaysiaMozambiqueNew CaledoniaNigeriaNetherlandsNorwayNew ZealandOmanPanamaPeruPolandPalestinePortugalReunionRomaniaSerbiaRussian FederationRwandaSaudi ArabiaSwedenSingaporeSloveniaSlovakiaTunisiaTurkeyTrinadad And TobagoTaiwanUkraineUruguayVenezuelaMayotteSouth Africa | Terminated | August 11, 2017 | 1061 | 160823 | Computer
Z-0733-2014 | 66202 | No | CJY | Calcium test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Consult DiagnosticsÂż Eon Calcium Reagent For In-Vitro Diagnostic Use |  |  | August 29, 2013 | 2013 | Vital Diagnostics, Inc. | High positive calcium bias on plasma sample versus results on serum | Device Design | Vital Diagnostics issued a Recall Notice dated August 29, 2013, to all affected customers via certified mail. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed not to use plasma for the determination of total calcium when using the affected product. The use of serum is acceptable, thus the product can still be used to determine total calcium with this sample type. Vital Diagnostics released a corection to the package insert for the Envoy Calcium Reagent. Customers with the affected product in their laboratory, were instructed to discard the insert packed with the product and replace it with the insert enclosed with this bulletin. If product was transferred to another location, the bulletin should be forwarded to that location. The package insert was revised indicating to only use Serum samples when running Calcium. Vital Diagnostics issued a revised (expanded) notice dated September 19, 2013, to include the Consult DiagnosticsÂż Eon Calcium Reagent (PN: 1429) requesting users to discontinue using the single vial Eon Calcium kit (PN: 1429). Customers with questions were instructed to call Vital Diagnostics Technical Support at 1-855-354-8324..For questions regarding this recall call 401-642-8400. |  | 213 | 213 units | Nationwide Distribution | Terminated | August 26, 2015 | 727 | 608 | Computer
Z-0045-2015 | 69257 | Yes | DRO | External transcutaneous cardiac pacemaker (noninvasive). | Cardiovascular | Cardiovascular | 510(k) | Covidien Kendall Adult Multi-Function Defibrillation Electrodes |  |  | September 18, 2014 | 2014 | Covidien LLC | Covidien Medi-Trace Cadence and Kendall Defibrillation Electrodes Not Compatible for use with Philips model FR3 or FRx AEDs | Device Design | Covidien notified customers by letter via Federal Express on 9/18/14. The letter requests that they immediately verify that they are not stocking Covidien electrodes for use with Philips FR3 or FRx AED units. A customer reply acknowledgement form is provided to allow Covidien to track the effectiveness of the notification. Covidien is adding a label to all levels of packaging that alerts the user that the product is not compatible with the FR3 and FRx AED units. Distribution partners notified to address product in the distribution chain that is available to be relabeled. For questions customers should call (800) 962-9888, option 8, then extension 2500.For questions regarding this recall call 508-261-8000. |  | 31590 | 31,590 pairs | Worldwide Distribtion-USA (nationwide) including Puerto Rico and the countries of Canada Greece Israel Belgium Finland France Germany Great Britain Ireland Italy Norway Poland Portugal Spain Sweden Switzerland Brazil Chile Philippines Taiwan Australia New Zealand. | Terminated | May 24, 2017 | 979 | 644460 | Computer
Z-0633-2014 | 66919 | No | MWI | Cardiac monitor (including cardiotachometer and rate alarm). | Cardiovascular | Cardiovascular | 510(k) | The Propaq LT Series monitors. Model numbers 802LTAN, 802LT0N, and 802LTRN. |  |  | November 21, 2013 | 2013 | Welch Allyn Protocol, Inc | Welch Allyn will update Propaq 802 Series Vital Signs Monitor, models are 802LTAN, 802LTON and 802LTR with software because the fault conditions relating to hardware or software failure or electrical noise caused by proximal defibrillation activity could, in rare instances, potentially interrupt device operations. | Device Design | Welch Allyn began sending out the Urgent Medical Device Correction letter, dated December 03, 2013, to consignees on December 10, 2013. The letter identified the product, the problem, and the action to be taken by the customer. Customer Actions: In the unlikely event that operation of your Propaq LT is interrupted and the display becomes blank or white, pressing the power (on/off) button will restart the device and quickly bring it back to full operational capacity.This notice should be passed on to all those who need to be aware within your organization, or to any organization where the potentially affected devices have been transferred.Welch Allyn is in the process of developing a software installation service tool that will enable installation of the software update by your biomed, IT staff or other technically qualified personnel.Upon release, the service tool will be made available to you via a download from Welch Allyn's website at www.welchallyn.com/PropaqltUpgrade or other media.In the meantime and on an on-going basis, the software update will be included if you schedule routine or other service with Welch Allyn or an authorized service provider.Customers with questions were instructed to call Welch Allyn Technical Support at 1-800-535-6663.For questions regarding this recall call 503-530-7935. |  | 6003 | 6003 units total (1715 units in the US and 4288 outside the US) | Worldwide Distribution - USA (nationwide) and Internationally to United Arab Emirates, Argentina, Australia, Belgium, Brazil, Canada, Switzerland, Chile, China, Columbia, Costa Rica, Czech Republic, Germany, Denmark, Ecuador, Spain, France, United Kingdom, Greece, Hong Kong,Taiwan, Ireland, Italy, Japan, Korea, Republic of, Kuwait, Lithuania, Luxemborg, Mexico, Malaysia, Mozambique, Namibia, Netherlands, Poland, Portugal, Qatar, Russia, Saudi Arabia, Singapore, Sierra Leone, El Salvador, Thailand, Turkey, Uganda, Vietnam, Kosovo, and South Africa. | Terminated | February 17, 2015 | 453 | 6003 | Computer
Z-2569-2014 | 68704 | No | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | Artis zee and Artis zeego systems. |  |  | June 05, 2014 | 2014 | Siemens Medical Solutions USA, Inc | There is a potential problem with Artis zee and Artis zeego systems running software version VC21B and being used in conjunction with the Large Display, in that under certain circumstances, the release of radiation can become blocked unnecessarily. | Software design | Siemens sent an "Important Customer Safety Notice" dated August 21, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Please forward this Safety information to other organizations that could be affected by this measure. If the device has been sold and therefore no longer in your possession, please forward this safety notice to the new owner. We would also request that you inform us of the identity of the device's new owner where possible.Please observe this Customer Safety Notice and comply with the corresponding measures until the update has been fully completed. For further questions please call (610) 219-6300. |  | 206 | 206 | Nationwide Distribution | Terminated | April 21, 2015 | 320 | 206 | Computer
Z-0164-2015 | 69430 | No | JIX | Calibrator. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Siemens Calibrator E; SMN 10309079 (2 pack), SMN 10321075 (6 pack), SMN 10335532 (6 pack Ref) |  |  | September 26, 2014 | 2014 | Siemens Healthcare Diagnostics, Inc | Positive bias for Calibrator E lot kits ending in 38 and 39 with the ADVIA Centaur Systems Cortisol and Progesterone assays compared to the respective master curves. | Device Design | An Urgent Medical Device Correction (UMDC) was sent to all affected Siemens Healthcare Diagnostics customers in the United States on September 26, 2014. An Urgent Field Safety Notice (UFSN) was sent out for distribution to all customers outside the United States on September 26, 2014. These notices inform the customer of the positive bias with ADVIA Centaur Systems Cortisol and Progesterone when calibrated with Calibrator E kit lots ending in 38 and 39.Customers may continue use of the affected lots as the clinical utility of the assay is not impacted. The communications also provide information on the performance of the assays when they use Calibrator E kit lots ending in 40. The notices should be reviewed with the facility's laboratory or Medical director.All affected domestic customers were sent a hard copy of the Urgent Medical Device Correction via Federal Express. Fax-back forms will be used to confirm that the customers received the UMDC. For those customers that do not return a fax-back form, a Siemens representative will follow up with those customers to ensure they received the Urgent Medical Device Correction and completed the required actions.All affected customers outside the US will be provided a copy of the Urgent Field Safety Notice via hard copy, e-mail, and/or fax as determined by each country's local regulations and procedures. Each customer will be required to return a completed confirmation fax-back to Siemens. For those customers that do not return a fax-back form, a Siemens representative will follow up with those customers to ensure they received the Urgent Field Safety Notice and completed the required actions. |  | 6 | Total = 19,352 (13,881 - 2 pack, 4878 - 6 pack, 593 - 6 pack REF) | Nationwide including Puerto Rico. Worldwide: Afghanistan, Armenia, Austria, Bahrain, Belgium, Bulgaria, Burkina Faso, Canary Islands, Croatia, Cyprus, Czech Republic, Denmark, Estonia, Finland, France, Georgia, Germany, Great Britain, Greece, Hungary, Iran, Iraq, Ireland, Italy, Jordan, Kazakhstan, Latvia, Lebanon, Libya, Lithuania, Luxembourg, Macedonia, Malta, Morocco, Netherlands, Norway, Oman, Pakistan, Poland, Portugal, Qatar, Romania, Russian Fed., Saudi Arabia, Serbia, Slovakia, Spain, Sweden, Switzerland, Turkey, Unit.Arab Emir., Vatican, Yemen, AR, AU, BD, BR, CA , CH, CL, CN, CO, DO, EC, EG, GT, HK, ID, IE, IL, IN, JP, KR, LY, MM, MX, MY, NI, NP, NZ, PE, PH, PY, SA, SG, SV, TH, TW, UY, VN, ZA | Terminated | January 29, 2016 | 490 | 6 | Computer
Z-2644-2014 | 69032 | No | NHT | Microorganism differentiation and identification device. | Microbiology | Microbiology | 510(k) | JBAIDS Instrument System Software version 3.5.0.72. |  |  | August 13, 2014 | 2014 | Biofire Defense | BioFire Defense, LLC is initiating a field action for JBAIDS Software version 3.5.0.72 because the wrong version software may have been installed. | Software Manufacturing/Software Deployment | Consignees were notified on 8/13/14 via letter and telephone and were provider with instructions on how to install the correct software. |  | 43 | 43 Units w/ laptops | Distributed USA (nationwide) and Guam. | Terminated | October 24, 2016 | 803 | 43 | Computer
Z-0748-2014 | 65366 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Merge Mammo is a multi-modality, vendor-neutral digital mammography workstation that enables imaging centers to use a single workstation to display and read images from different vendors and different acquisition devices, including digital mammography, ultrasound, MRI and CT. |  |  | August 26, 2013 | 2013 | GE Healthcare It | There is a potential safety issue with the 8.0.2 version software of the Merge (Cedara) Mammo viewer. The Merge Mammo viewer can load and display images not in synch with the selected or highlighted images in the thumbnail navigator. This incorrect image display issue could result in an inaccurate reading on the Merge Mammo workstation. This could result in a missed intervention or mis-diagnosis. | Software design | GE Healthcare sent an URGENT MEDICAL DEVICE CORRECTION letter dated August 26, 2013, to all affected customers. The letter includes important safety instructions for healthcare providers to refer to that can minimize the chance that an incorrect image will be used in the diagnosis or treatment of a patient. The letter further states that a GE Healthcare service representative will contact the customer to schedule a time to install a software update to version 8.0.2 that will correct the software defect.. Customers with any concerns about this defect or recall are encouraged to contact their local GEHC-IT service representative or the GE Remote Online Center (ROC) Customer Care Center at 1-800-437-1171.For questions regarding this recall call 800-437-1171. |  | 28 | 28 units | Worldwide Distribution - USA including AK, AR, CA, IA, ID, IL, MI, MS, MT, NJ, NV, NY, OR, PA, SC, TN, WI and Internationally to Israel, Mexico, and Spain. | Open, Classified |  |  | 28 | Computer
Z-0473-2015 | 69793 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | IntelliSpace Portal DX/HX/EX model number: 881001 |  |  | November 09, 2014 | 2014 | Philips Medical Systems (Cleveland) Inc | Software defect. In certain circumstances, the application may display incorrect measurements of cardiac and aortic anatomy used to select and size the Transcatheter heart valve (THV) to be implanted . | Software design | Consignees were notified of the Field Safety Notice via certified mail. |  | 32 | 32 | Worldwide Distribution: US Distribution in states of: FL, NC, and VT; and the countries of Argentina, Austria, Belgium, Canada, Denmark, France, India, Italy, Japan, Netherlands, United Kingdom, and Vietnam. | Terminated | September 14, 2015 | 309 | 32 | Computer
Z-0731-2014 | 64724 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Centricity PACS System versions 3.X and higher and 4.X and higher |  |  | September 15, 2012 | 2012 | Ge Healthcare It | There is a potential loss of image(s) associated with the Centricity to Centricity (C2C) exam transfer module of Centricity PACS System versions 3.X and higher and 4.X and higher when transferring from the source to a destination server. When another process in the destination server attempts to access the same object or table, the transfer process of a particular image may be terminated. Once ter | Software design | GE Healthcare sent an Important Product Information letter dated September 15, 2012 to affected customers. The letter identified the affected product problem and actions to be taken. GE Healthcare will provide a patch to 3.2.X and 4.0 Systems to remediate the issue. For questions contact your local GEHC-IT service representative. |  | 748 | 748 units | Worldwide Distribution - US: Nationwide (AK, AL, AR, AZ, CA, CO, CT, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WV, WY); FOREIGN: Australia, Austria, Belgium, Bermuda, China, Denmark, Egypt, France, Germany, Hong Kong, India, Ireland, Israel, Italy, Japan, Korea, Kuwait, Malaysia, Malta, Netherlands, Portugal, Saudi Arabia, Singapore, South Africa, Spain, Sweden, Switzerland, Taiwan, Turkey, United Kingdom. | Open, Classified |  |  | 748 | Computer
Z-0791-2014 | 66311 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Brilliance iCT, Brilliance 64 and Brilliance Big Bore |  |  | March 07, 2011 | 2011 | Philips Medical Systems (Cleveland) Inc | The DoseRight feature suggest a mAs based on the measured patient size, a reference size and a reference mAs. When scanning large children, the suggested mAs may be higher than clinicians would expect. | Software design | Philips Medical Systems sent a Field Safety Notice letter dated February 25, 2011, to all affected customers.The letter identified the product the problem and the action needed to be taken by the customers.Your Corrective Action Plan (CAP) is approved and consists of:1. Notifying all affected users with Field Safety Notice 72800527/72800528 starting March 8, 2011. This Field Safety Notice has been reviewed and is approved.2. Following up with all responses by dispatching a Field Service engineer to each site to install the software update at no cost to the user.3. Tracking the responses from users by certified mail.For further questions please call (440) 483-7600. |  | 23 | 23 | Worldwide Distribution - US Distribution including the states of DC, IL, IN, MI, MO, NY, OH, OR, PA, TX and VT., and the countries of Australia, Belgium, China, Germany, India, Israel, Italy and The Netherlands. | Terminated | April 28, 2014 | 1148 | 23 | Computer
Z-0165-2015 | 69424 | No | HAB | Surgical instrument motors and accessories/attachments. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | UNIVERSAL Modular Electric/Battery Handpiece for surgery |  |  | October 06, 2014 | 2014 | Zimmer Surgical Inc | Non-injury complaints from Europe identified a malfunction that made it possible for a handpieceto start by itself when the power source is connected. | Device Design | Zimmer sent an Urgent Medical Device Correction letter via email to US distributors with traceable Fed Ex mailings to all consignees dated October, 2014. The letter identified the affected product, problem and actions to be taken. Customers were advised of the potential malfunction and the operating procedure for the handpiece per the Instructions for Use (IFU). Zimmer Surgical, as the importer, will provide notice to consignees in the US accordingly. In addition, units will receive an upgrade as part of the annual preventive maintenance per the Instructions for Use. |  | 828 | 828 | US Nationwide Distribution in the states of: AK, AL,AR, AZ,CA,CO,CT, FL, GA, HI, IA, IL, IN, KS, KY, LA, MA,ME, MI, MN, MO, NC, NE, NJ, NY, OH, OR, PA, RI, SC, TN, TX, VA, WA, WI, WV, and WY. | Terminated | April 15, 2015 | 191 | 828 | Computer
Z-0211-2015 | 69557 | Yes | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | Cardiohelp-I System |  |  | October 17, 2014 | 2014 | Maquet Cardiovascular Us Sales, Llc | Cardiohelp- i System may have a software issue that can potentially result in an erroneous display of a &quot;Battery Needs Service&quot; message after startup of the units when using either AC or DC power | Software Design Change | Maquet Inc. sent a Product Recall Urgent Medicall Device Field Correction letter to all of their customers via Federal Express. The letter identified the affected product, problem and actions to be taken. The letter stated that a software upgrade will be performed by an authorized Maquet Service Technician. A Maquet representative will contact customer to schedule an on-site service of the device. Any questions can be directed to the local Maquet representative or Maquet Customer Service at 1-888-627-8383 (option 2, option 2) Monday through Friday, between the hours of 8:00 am and 6:00 pm (EST). |  | 331 | 331 units | Worldwide Distribution - US Nationwide | Terminated | April 06, 2016 | 537 | 331 | Computer
Z-0861-2014 | 65179 | No | JXG | Central nervous system fluid shunt and components. | Neurology | Neurology | 510(k) | Codman Certas Programmable Valve In Line Valve only with SiphonGuard Device |  |  | May 16, 2013 | 2013 | Codman &amp; Shurtleff, Inc. | Codman Certas Programmable Valves used for hydrocephalus may not operate properly | Device Design | Codman issued an Urgent Notice dated May 16, 2013, to all affected customers. The letter identifed the affected product, problem and actions to be taken. Mitigation instructions are provided and if not followed or are not successful, either issue could lead to symptoms associated with under- or over-drainage. Units not implanted are to be returned. A U.S Acknowledgement Form and fax are requested to be completed and faxed to: 1-888-943-4193, or e-mail the form to codman8894@stericycle.com. If you have any questions on concerns in regards to this recall, please contact your Codman Representative or Stericycle at 1-888-943-2395. |  | 1481 | 1481 units | Worldwide Distribution - US Nationwide, Australia, Austria, Belgium, Brazil, Switzerland, Czech Republic, Germany, Spain, Finland, France, Great Britain, Israel, Italy, Netherlands, Norway, Portugal, Russia, Saudi Arabia, Sweden, Slovenia, Slovakia, Turkey, South Africa, and UK. | Terminated | June 02, 2015 | 747 | 5791 | Computer
Z-0853-2014 | 67325 | No | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | GE Healthcare Innova IGS 630. Biplane Cardiovascular and Interventional Imaging System. |  |  | November 15, 2013 | 2013 | GE Healthcare, LLC | GE Healthcare has recently become aware of a potential safety issue with respect to IGS 630 Imaging Systems.While performing fluoroscopy on a biplane fluoroscopy unit, there is a potential for loss of the x-ray imaging function when the user changes the size of the lateral FOV (Field of View) and releases the biplane footswitch pedal simultaneously. This could lead to a system lock up requiring | Software design | GE Healthcare sent an "Urgent Medical Device Correction" letter dated November 15. 2013. The letter was addressed to Hospital Administrators / Risk Mangers, Mangers of Radiology / Cardiology / Surgery Department and Radiologists / Cardiologists / Surgeons. The letter described the Safety Issue, Safety Instructions, Affected Product Details, and Product Correction. The letter stated that GE Healthcare will correct all affected systems at no cost. For questions call (In US 800-437-1171; for other countries contact your local GE Healthcare Service Representative). |  | 28 | 28 (12 US, 16 OUS). | Worldwide Distribution - US Nationwide in the states of CA, CO, FL, NV, NJ, PA, TX and the countries of CZECH REPUBLIC, FRANCE, JAPAN, REPUBLIC OF KOREA, SAUDI ARABIA, and TURKEY. | Terminated | March 28, 2014 | 133 | 28 | Computer
Z-0854-2014 | 66982 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | GE Healthcare CARESCAPE Monitor B650 and CARESCAPE Monitor B850. |  |  | December 27, 2013 | 2013 | GE Healthcare, LLC | GE Healthcare has recently become aware of potential safety issues associated with the CARESCAPE Monitor B850 and CARESCAPE Monitor B650. These included 25 issues related to NIBP (Non-InvasiVentilation Alarms:ve Blood Pressure) , ECG, Central Monitor, General, Tram, and Bed to Bed Issues. | Software design | GE Healthcare sent an Urgent Medical Device Correction letter between December 27, 2013 and June 27, 2013, to all affected customers. The letter described the Safety Issues and Instructions, Affected Product Details, Product Correction and Contact Information. Customers with questions were instructed to contact Technical Support at 1-800-558-7044 or their local Service Representative.For questions regarding this recall call 800-558-7044. |  | 40239 | 40,239 (17,182 units US, 23,147 units OUS). | Nationwide Distribution including DC and PR, except DE, SD, ND, and VT. OUS: Algeria, Argentina, Australia, Austria, Azerbaijan, Bahrain, Bangladesh, Belgium, Bermuda, Bolivia,Brazil, Canada, Chile, China, Colombia, Costa Rica, Croatia, Cyprus, Czech Republic ,Denmark, Ecuador, Egypt, El Salvador, Estonia ,Finland, France, Germany, Greece,Guam, Guatemala, Honduras, Hong Kong, Hungary, Iceland, India, Indonesia, Iraq, Ireland, Israel, Italy, Jamaica, Japan, Jordan, Kazakhstan, Kuwait,Latvia, Lebanon, Libyan Arab Jamahiriya, Lithuania, Malaysia, Malta, Mexico, Morocco,Namibia, Nepal, Netherlands, New Zealand, Nicaragua, Norway, Oman, Pakistan, Panama, Paraguay, Peru, Philippines, Poland, Portugal, Puerto Rico, Qatar, Romania, Russian Federation, Saudi Arabia,Singapore, Slovakia, Slovenia, South Africa, South Korea, Spain, Sri Lanka, Sweden, Switzerland, Syrian Arab Republic, Taiwan, Tanzania, United Republic of Tunisia, Turkey, United Arab Emirates, United Kingdom, Venezuela, Uruguay. | Open, Classified |  |  | 40239 | Computer
Z-0161-2015 | 58185 | Yes | DTR | Cardiopulmonary bypass heat exchanger. | Cardiovascular | Cardiovascular | 510(k) | Terumo Custom Cardioplegia Delivery Sets |  |  | March 01, 2011 | 2011 | Terumo Cardiovascular Systems Corp | Defect in the luer connector can compromise the seal on the female side of the connector and the compromised seal can introduce air in the table line | Device Design | Terumo Cardiovascualr notified accounts by letter dated 3/1/2011 advising users users that the luer connector to the table line in certain lots of myocardial protection sets may leak. Users should not continue to use affected myocardial protection sets in a manner that exposes the connector to negative pressure without taking correction or replacement activities as described. Firm provided mitigating instructions:Clamp the table line distal to the luer connector when ending an infusion of cardioplegia thus isolating the leak from negative pressure. Contact Terumo CVS Customer Service to report the defect and arrange return of the unit: 800-521-2818. |  | 24 | 24 units | Worldwide Distribution: US (nationwide) and the countries of Canada, Brazil, Hong Kong, Korea, and Thailand. | Terminated | November 06, 2014 | 1346 | 46190 | Computer
Z-0094-2015 | 69301 | No | JDI | Hip joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Continuum Acetabular System Trabecular Metal Shell with Multi Holes, Porous 68mm, Prosthesis, Hip, Semi-Constrained, Metal/Polymer |  |  | October 01, 2014 | 2014 | Zimmer, Inc. | The affected products are missing polar boss threads. | Equipment maintenance | Zimmer sent an "URGENT MEDICAL DEVICE RECALL- LOT SPECIFIC" notifications dated October 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Your Responsibilities1. Review the notification and ensure affected personnel are aware of the contents.2. If you find any product from the affected lot , quarantine the product and notify your Zimmer sales representative.3. Your Zimmer sales representative will remove the recalled product from your facility.4. For patients that previously had this product implanted, it is recommended that you continue your normal post operative follow up routine.5. If after reviewing this notification you have further questions or concerns please call the customer call center at1-800-348-2759 between 8:00 am and 5:00pm EST. |  | 19 | 19 | Worldwide Distribution: US (nationwide) in the states of : IA, CA, NC, AL, WA, FL, TN, MA, AZ, TX, UT, and NV; and Internationally to Japan. | Terminated | March 02, 2015 | 152 | 19 | Computer
Z-0851-2014 | 66585 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | User Defined Method Flex Assignment/Siemens Dimension Vista 500 or Dimension Vista 1500 System with software version 3.5.1 or lower |  |  | October 07, 2013 | 2013 | Siemens Healthcare Diagnostics, Inc. | When utilizing the Routine Inventory screen to enter a User Defined Method (EMPTY) Flex the system may assign the User Defined Method Flex to a different Flex that is currently in inventory on the system, and then use the incorrect Flex Cartridge to process the user defined method. | Software design | Siemens sent a Urgent Medical Device Correction letter dated October 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Please use the following actions if using Vista SW Versions 3.5.1 and below to prevent the issue from occurring:1. When assigning a User Defined (EMPTY) Flex, use the Advanced Inventory Screen per the Operators Guide.2. Verify that EMPTY Flex is no longer listed in the inventory before proceeding with any testing.3. If the EMPTY Flex is still listed in the inventory, contact your Siemens Customer Care Center - Technical Solutions for assistance.Please retain this letter with your laboratory records and forward this letter to those who may have received this product. Discuss the content of this letter with your Medical Director.We apologize for the inconvenience this situation has caused. If you have any questions, please contact your Siemens Customer Care Center - Technical Solutions at (800-441-9250) or your local Siemens technical support representative. |  | 2038 | 2038 | USA (nationwide) Distribution including Puerto Rico. | Terminated | August 27, 2014 | 324 | 2038 | Computer
Z-0152-2015 | 69425 | No | IWB | Radionuclide radiation therapy system. | Radiology | Radiology | 510(k) | Laksell Gamma Knife Perfexion |  |  | October 06, 2014 | 2014 | Elekta, Inc. | Elekta has become aware that the latches may be locked before they have been fully turned resulting in a poorly locked frame. | Device Design | Elekta sent an" Important Field Safety Notice" distributed on October 6, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. This Important Field Safety Notice describes the problem and possible work arounds for customers to follow until a permanent solution can be developed. They are:" Always assure that the plastic lever is only operated when at a right angle with the frame adapter." Do not force the plastic lever in place if it meets significant resistance when turned." Ensure that the plastic lever is completely flush with the frame adapter and that no angulations are present.Corrective Action #2 - Elekta will check and replace as necessary affected frame adapters in the field. This will be released at the end of October 2014, and Elekta will have six months to complete the checks and replacements for those products in the field. If you have any questions please contact your local Elekta office, or call (770) 300-9725. |  | 181 | 181 | Worldwide Distribution - USA (nationwide) including the states of AK, AZ, AR, CA, CO, CT, FL, GA, HI, IL, KY, LA, MA, MI, MN, MS, MO, MT, NJ, NM, NY, NC, OH, OR, PA, RI, SC, TX, TN, UT, VA, WA and WV, and the countries of Australia, Austria, Belgium, Brazil, Canada, China, Colombia, Czech Republic, Egypt, France, Germany, Greece, Hong Kong, India, Indonesia, Italy, Japan, Mexico, Morocco, Netherland, Norway, Poland, Portugal, Russia, Saudi Arabia, South Korea, Spain, Sweden, Switzerland, Taiwan, Turkey and United Kingdom. | Open, Classified |  |  | 181 | Computer
Z-0677-2014 | 67166 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | 14" (36 cm ) Appx 0.50 ml, Smallbore Ext Set w/MicroClaveÂż Clear, |  |  | December 24, 2013 | 2013 | ICU Medical, Inc. | ICU Medical is recalling certain lots of the NanoClave Needlefree Connector product line because of a small percentage of devices that may leak. | Equipment maintenance | icumedical sent an Urgent Medical Device Recall letter dated December 24, 2013, to all affected customers. The letter informed the customers of the problems identified and the actions to be taken. Customers were instructed to return completed verification forms. Customers with questions were instructed to contact ICU Customer Service Monday through Friday between the hours of 8:30AM and 4:00PM Pacific Time: (866) 829-9025 and select option 8, or email: customerservice@icumed.com.ICU Medical reported an update on January 6, 2014, an expansion of the recall to additional lots. Updated recall letter was sent out to all customers on additional lots on January 6, 2014 via UPS with tracking.For questions regarding this recall call 949-366-2183. |  | 100 | 100 units | Worldwide Distribution - USA including UT, TX, GA, WI, VT, NH, AL, OH, TN, IL, CA, KY,HI, NC, and Internationally to CA, BE, IN, KW, and AU. | Terminated | April 01, 2014 | 98 | 147975 | Computer
Z-0204-2015 | 69367 | walker, mechanical | ITJ | Mechanical walker. | Physical Medicine | Physical Medicine | 510(K) Exempt | Transport Rollator with Padded Seat; Large 8-inch wheels with wheel locks (indoor and outdoor use), fold down footrests, ergonomic handles, curved padded backrest and handy wire basket for storage. Item# 30192 Burgundy each/Item# 30194 Blue each.. |  |  | October 01, 2014 | 2014 | Roscoe Medical Inc | The caster may separate completely from the device frame causing injury to user. | Device Design | The firm, Roscoe Medical, issued a "URGENT RECALL NOTICE-Timely Response Required" Quarantine Communication to customers between 09/08/2014 and 09/10/2014. The notice identified affected products as 30194, serial number range MH140407092MH140407291, the problem and actions to be taken. The customers were instructed NOT to use this product and to destroy any affected product that may be in their inventory/control; immediately complete and return the attached Recall Response Form/Certificate of Destruction to Roscoe Medical- "Transport Rollator Recall" within 15 days of receipt of the recall notification via fax to (440) 572-4261 or email: corporate@roscoemedical.com, even if you do not have any affected product; and notify your customers immediately of this product recall.Should you have any questions, contact Roscoe Medical Customer Support at (800) 376-7263 Monday-Friday 8:00 am EST - 5:00 pm EST. |  | 831 | 831 units | Nationwide Distribution: US including states of: AZ, CA, CO, FL, GA, IA, ID, IL, IN, KY, LA&lt; MA, MD, MI, MN, MO, NC, NJ, NM, NV, NY, OH, PA, SC, TN, TX, VA, WA, WI,and WY. | Open, Classified |  |  | 831 | Computer
Z-0714-2014 | 67056 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Draeger Infinity Acute Care System Monitoring Solution |  |  | December 10, 2013 | 2013 | Draeger Medical Systems, Inc. | After a test in which a low SpO2 (yellow) medium grade alarm initiated in neo-natal mode and was subsequently paused, the high grade life threatening SpO2 alarm (red) did not activate when the SpO2 values were decreased to reach the red alarm limit. | Software design | An Urgent Medical Device Recall Letter, dated December 2013, was sent to Users. Once available Drager will update components of the IACS Monitoring Solution System. |  | 439 | 439 | US Distribution including the states of: MA, NH, IN, CA, TX, NY, OK, and IL. | Terminated | May 20, 2015 | 526 | 439 | Computer
Z-0747-2014 | 67115 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | GE Healthcare Carescape Patient Data Module |  |  | November 27, 2013 | 2013 | GE Healthcare, LLC | GE Healthcare has recently become aware of a potential safety issue due to intermittent failure of the spring loadedmounting latch of PDMs produced from April 2007 through December 2012. | Device Design | The firm, GE Healthcare, sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated November 26, 2013 to its consignees/customers. The letter was addressed to Healthcare Administrator / Risk Manger, Chief of Nursing and Director of Biomedical Engineering. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction and Contact Information. The consignees/customers were instructed to please ensure that all potential users in your facility are made aware of this safety notification and the recommended actions. GE Healthcare will provide a new latch at no charge once it is available. GE will contact you to arrange for installation of the latch. If you have any questions regarding this notification, please contact Technical Support at 1-800-558-7044, or your local Service Representative. |  | 43033 | 43,033 (25,681 USA; 17,352 OUS) | Worldwide Distribution: US (nationwide) including states of: AL, AZ, AR, CA, CO, CT, DC, FL, GA, ID, IL, IN, IA, KS,KY, LA, ME, MD, MA, MI, MN, MS, NO, NE, NV, NJ, NM, NY, NC, OH, OK,OR, PA, RI, SC, TN, TX, UT,VT, VA, WA, WI, and WY. Internationally to: Argentina, Australia, Austria, Bangladesh, Belgium, Brazil, Canada, Chile, China, Colombia, Costa Rica, Czech Republic, Denmark, Egypt,England, Estonia, Finland,France, Germany,Greece, Guatemala, Hungary,India, Indonesia, Ireland, Israel, Italy, Jamaica, Japan, Jordan, Korea, Kuwait, Lebanon, Lithuania, Malaysia, Mauritius, Mexico, Morocco, Nepal, Netherlands, New Zealand, Norway, Oman, Pakistan, Panama, Paraguay, Peru, Poland, Portugal, Qatar, Russia, Saudi Arabia, Singapore, Slovakia, South Africa, Spain, Sweden, Switzerland, Taiwan, Republic of China, Thailand, Trinidad and Tobago, Tunisia, Turkey, United Arab Emirates and United Kingdom. | Open, Classified |  |  | 43033 | Computer
Z-0900-2014 | 67298 | No | ITX | Diagnostic ultrasonic transducer. | Radiology | Radiology | 510(k) | Atlantis PV, 8.5F, 15 MHz Peripheral Imaging Catheter: |  |  | January 06, 2014 | 2014 | Boston Scientific Corporation | Reports of physicians experiencing inability to pass a 0.035" guidewire through the guidewire lumen and exit port at the (proximal) Y-manifold of the imaging catheter. | Equipment maintenance | Boston Scientific sent an Urgent Medical Device Recall Removal letter dated January 14, 2014 to all affected customers by overnight mail. The letter identified the affected product, problem and actions to be taken. Customers were instructed to complete an Account Reply verification Tracking Form in order to received a Return Goods Authorization number for product return. For questions contact your local Sales Representative. |  | 1068 | 1068 single units | US Nationwide Distribution and Mexico | Terminated | October 29, 2014 | 296 | 1068 | Computer
Z-0905-2014 | 66835 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Medtronic Advanced Energy Aquamantys3 Pump Generator; Product Catalog Number: 10-1357 (Generator assembly), 40-404-1(as shipped, including accessories) |  |  | November 12, 2013 | 2013 | Medtronic Advanced Energy, LLC | The internal protection circuitry on the electrosurgical output that prevents a patient from exposure to DC voltage when the device is activated is being compromised during normal use. | Device Design | Medtronic sent an Urgent Product Removal Notification on November 12, 2013, via Next Day FEDEX. The communication advises users to immediately stop using the affected AQM3 electrosurgical generators and quarantine them until they can be returned to Medtronic Advanced Energy. Medtronic field personnel will collect them in order to remove them from service and return them to Medtronic Advanced Energy. Customers with questions were instructed to contact Customer Service at 866-777-9400.For questions regarding this recall call 866-777-9400. |  | 242 | 242 devices | Nationwide Distribution including CA, MO, PA, NJ, NY, OH, VA, TX, NC, DE, KS, TN, LA, IL, MI, SC, WI, and FL. | Terminated | April 03, 2014 | 142 | 242 | Computer
Z-2653-2014 | 68909 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | VERO Linear Accelerator System, Model Number MHI-TM2000; Common Name: Medical Linear Accelerator. |  |  | July 05, 2014 | 2014 | MITSUBISHI HEAVY INDUSTRIES, LTD., HIROSHIMA MACHINERY WORK | Software Anomaly: If a user changes the calendar setting from Workday to Holiday or vice versa, that would erroneously change the status of treatment/fractions, which are completed and to become Treated (completed) status on the exact day when such change is made, to Untreated status. Furthermore, such treatment would be cloned and mistakenly added to the schedule as Untreated treatment | Software design | A Preliminary Customer Information Letter (dated 7/4/14) was emailed to customers on 7/05/14 from Mitsubishi's sales and service representative in Germany. The letter informed customers of the potential safety issue associated with the VERO Linear Accelerator and the recommended actions to be taken. A formal URGENT FIELD SAFETY NOTICE Letter (dated 7/15/14) was emailed to customers on 8/07/14 from Mitsubishi's sales and service representative in Germany. The letter instructed customers to ensure that all potential users in the facility are made aware of the safety notification and the recommended actions to be taken. |  | 3 | 3 units in use in the US | Distributed in the states of NY, FL, and TX. | Terminated | April 01, 2015 | 270 | 3 | Computer
Z-0463-2015 | 69542 | No | JJH | Clinical sample concentrator. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Abbott m2000sp, an automated system for performing sample preparation for nucleic acid testing. |  |  | October 23, 2014 | 2014 | Abbott Molecular | Some versions of Application Specifications (App Spec) are incompatible with m2000sp system software version 6.0 and 7.0. This issue will cause Error Code 9000 (An unexpected error has occurred: Software error) to be generated when the Available Reagent Packs screen is selected. | Software design (manufacturing process) | Field Correction Recall / Urgent Field Safety Notice letters dated September 23, 2014 were sent to all direct accounts (customers). The letters included instructions for customers to: 1) review the app spec files that are installed on the m2000sp system and delete any incompatible app specs; 2) if one or more app spec files were deleted, installation of new app spec file(s) is needed; and, 3) review this information with laboratory personnel and retain the letter for future reference. Customers with questions about this market action can contact Technical Support at 800-553-7042, Option 2, or the local Abbott Molecular representative. |  | 9 | 9 CD ROMs (US); 1,273 CD ROMs (Foreign) | Worldwide distribution. US nationwide (CA, CT, FL, NC, PA, SC, and DC), Algeria,Australia,Austria,Belgium,Bosnia and Herzegovina,Cameroon,Canada,Chad,Colombia,Costa Rica,Croatia,Denmark,Egypt,Ethiopia,Finland,France,Germany,Greece,Guinea,Honduras,Hong Kong,India,Indonesia,Ireland,Israel,Italy,Japan,Kenya,Korea,Lithuania,Mali,Mexico ,Netherlands,New Zealand,Nigeria,Norway,Pakistan,Paraguay,Poland,Portugal,Romania,Russia,Saudi Arabia,Senegal,Slovenia,South Africa,Spain,Sweden,Switzerland,Taiwan,Tanzania,Thailand,Turkey,Uganda,Ukraine,United Arab Emirates, United Kingdom,Vietnam | Terminated | March 08, 2016 | 502 | 9 | Computer
Z-0725-2014 | 67136 | No | IZL | Mobile x-ray system. | Radiology | Radiology | 510(k) | GE Optima XR220amx and Optima XR200amx with Digital Upgrade Mobile Radiographic Systems. Model numbers 5555000-5, 5555000-6, and 5421698. |  |  | November 07, 2013 | 2013 | GE Healthcare, LLC | There exists a software issue associated with patient selection from the worklist on the Optima XR200 with Digital Upgrade, and Optima XR220amx Mobile X-ray Systems. Intermittently when a large number of patient exams are retrieved from the HIS/RIS system, the exams when displayed on the mobile x-ray system are not properly displayed. If the user continues to attempt to select one of the exams w | Software design | GE sent an Urgent Medical Device Correction letter, dated 11/7/2013 to consignees. The letter described the issue, identified devices affected and stated that users should pay close attention to the patient exam/entry selected from the worklist prior to selecting the Start Exam button to ensure that the system did not update and select an incorrect exam. The user may also confirm during the exam that the appropriate patient information was selected via the display on the upper left portion. GE will fix the units with a software fix at no cost. |  | 1027 | 1027 | Worldwide Distribution-USA (nationwide) including Washington D.C., and Guam and the countries of ALBANIA, AUSTRALIA, BAHRAIN, BELGIUM, BRAZIL, CANADA, CHILE, CHINA , COLOMBIA, DENMARK, FINLAND, FRANCE, GERMANY, GREECE, HUNGARY, INDIA, IRAQ, IRELAND, ITALY, JAPAN, JORDAN, KOREA, REPUBLIC OF, KUWAIT, MALAYSIA, MEXICO, NEW ZEALAND, NIGERIA, NORWAY, POLAND, ROMANIA, RUSSIAN FEDERATION, SAUDI ARABIA, SINGAPORE, SOUTH AFRICA, SPAIN, SWEDEN, THAILAND, TRINIDAD AND TOBAGO, TURKEY, UNITED ARAB EMIRATES, and UNITED KINGDOM. | Terminated | May 06, 2014 | 180 | 1027 | Computer
Z-2652-2014 | 68113 | No | FSA | Non-AC-powered patient lift. | General Hospital | General Hospital | 510(K) Exempt | Tenor is a mobile passive lift, intended to be used on horizontal surfaces for lifting and transfer in hospitals, nursing homes or other healthcare facilities. Transfers with the Tenor are made indoors in the patient's room, in communal areas or in a bathroom. Tenor is meant to be used with the patient type who: sits in a wheelchair, has no capability to support themselves at all, cannot stand unsupported and is not able to bear weight (not even partially), is dependent on caregiver in most situations, is a passive resident, might be almost completely bed ridden, is often stiff and has contracted joints, and/or is totally dependent. The Tenor has a safe working load of 320 KG (705 lbs). |  |  | June 10, 2014 | 2014 | ARJOHUNTLEIGH POLSKA SP. ZO.O. | ArjoHuntleigh received reports where the Tenor Lift without Scale (spreader bar) bolt broke during use. There is a remote risk concerning the breakage of the spreader bar bolt connection, part # TEN.018. During use, the bolt is subjected to mechanical forces, which may lead to breakage. | Device Design | ArjoHuntleigh sent an URGENT - FIELD SAFETY NOTICE letter dated June 10, 2014, to all affected customers. The letter included instructions for direct accounts to: 1) ensure that caregivers and users of the Tenor patient transfer lift referenced above are made aware of the field safety notice; 2) ensure that a copy of the field safety notice is placed in the Tenor Lift Instructions for Usemanual for future reference and training; 3) complete and sign the enclosed Customer Response Form and return the form to the ArjoHuntleighRegulatory Department; 4) and to be advised that an ArjoHuntleigh Service Technician will contact the facility representative listed on the Customer Response Form to schedule service to replace your spreader bar, free of charge. Customers with questions about this recall can contact ArjoHuntleigh at 800-323-1245 x 57985 or via e-mail at AHNA-RA@arjohuntleigh.com. |  | 382 | 382 lifts total worldwide (160 US) | Worldwide Distribution - USA (nationwide) and Internationally to Australia, Bahrain, Canada, France, Germany, India, Ireland, Israel, Italy, Mexico, Netherlands, New Zealand, Poland, Romania, Saudi Arabia, South Africa, Spain, Sweden, United Kingdom, and Uruguay. | Open, Classified |  |  | 382 | Computer
Z-0471-2015 | 69748 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Cardiac Function Analysis Software |  |  | November 05, 2014 | 2014 | Toshiba American Medical Systems Inc | Potential problem with the cardiac function analysis software (CFA). Incorrect analysis results may be displayed in the Function Parameters for the Entire Heart displayed as analysis results of the CFA and in a Left-Ventricular Volume Curve generated based on some of those parameters. | Software design | Toshiba sent an Urgent Medical Device Correction letter dated November 5, 2014, to all affected customers. Customers were asked to refrain from using the function parameters for the entire heart and the left-ventricular volume curve until corrective actions have been taken.Customers were instructed to share this information with all users and reviewing radiologist as well as clinical engineering or Biomedical group at facility.Please complete and return the attached form and fax it to the toll free number at the top of the form 877-349-3054. This form can also be sent via email to raffairs@tams.com.Revised software to prevent occurrence of this problem will be installed on the system. When the new software becomes available the Toshiba service representative will contact for an appointment to install it on the system.Customers with questions were instructed to call 800-421-1968 or contact their local Toshiba Representative at 800-521-1968. |  | 1 | 1 unit within the U.S. only | US Nationwide Distribution to CA, MA, MD | Terminated | October 09, 2015 | 338 | 3 | Computer
Z-0196-2015 | 69472 | No | GDC | Operating tables and accessories and operating chairs and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | The 5085 is a general surgical table with high patient weight capacity, extended width capability, and lower minimal table top elevation. |  |  | October 10, 2014 | 2014 | Steris Corporation | The 5085 and 5085SRT stainless steel column shrouds may become damaged as a result of misalignment during extreme table articulation, excessive side force on the table column, or pressure on the base of the table. Severe damage to the shrouds may prevent the user from raising or lowering the surgical table. | Device Design | STERIS will mail Consignee Notification Letters to the affected Consignees via FedEx with tracking numbers for delivery confirmation. The Consignee Notification mailing will initiate the week of October 13, 2014. |  | 999 | 999 | US distribution: AL, AZ, AR, CA, CO, CN, DC, FL, GA, HI, ID, IL, IN, IA, KN, KY, LA, ME, MD, MI, MN, MS, MO, NB, NV, NY, NC, ND, OH, OK, OR, PN, RI, SC, SD, TN, TX, VA, WV, WI. Foriegn: Austrialia, Brazil, Canada, China, Ecuador, India, Korea, Mexico, Pnama, and Thailand. | Terminated | June 26, 2015 | 259 | 999 | Computer
Z-0077-2015 | 69234 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | VITROS 4600 Chemistry System (VITROS 5,1 FS System family member), System Product Code 6802445. |  |  | September 08, 2014 | 2014 | Ortho-Clinical Diagnostics | Ortho Clinical Diagnostics (OCD) identified an anomaly with Software Version 3.0 and below on the VITROS 4600 Chemistry Systems and VITROS Integrated Systems. Internal testing confirmed that when using calibrator barcode labels supplied with VITROS Chemistry Products Calibrator Kits 1, 2, 3, 4, 6, or 9, an unexpected assay calibration may occur if assay targets are unassigned (i.e., hidden). | Software design | On 9/08/14, URGENT PRODUCT CORRECTION NOTIFICATION Letter (CL2014-232, dated 9/8/14) and Confirmation of Receipt Form were sent via FedEx overnight courier or ORTHO PLUS e-Communications and/or US Postal Service Priority Mail (for PO Boxes only) to US consignees (OCD Direct, Drop-Ship and US Federal Government only) to inform them of the issue and advise them of the required actions. Foreign affiliates were informed of the issue via e-mail on 9/8/14 and instructed to notify their consignees of the issue and the required actions. |  | 232 | USA: 97; Foreign: 232 | Worldwide Distribution-USA (nationwide) including Puerto Rico and the countries of Australia, Brazil, Canada, Chile, China, Colombia, England, France, Germany, India, Italy, Japan, Mexico, Panama, Singapore, Spain and Venezuela. | Open, Classified |  |  | 1177 | Computer
Z-0181-2015 | 68723 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Philips Medical Systems Allura Xper FD20C Radiological Imaging with software version R8.2.O; System Code: 722028. |  |  | July 08, 2014 | 2014 | Philips Medical Systems, Inc. | Faulty Automatic Motion Controller (AMC), a problem in the Power On Self Test (POST) error handling was detected, can result in a hazardous movement of the C-arc. system. | Software design | Philips Medical Systems issued certified letters on 7/8/14 and receipt returned for tracking purposes. A mandatory Field Change Order with reference FC072200253 is being released that requires Philips field service engineers to install Software release R8.2.0.3. If you need any further information or support concerning this issue, contact your local Philips representative: SUPPORT 1-800-722-9377. |  | 22 | 22 systems | Distributed in the states of IL, and TN, and the countries of Argentina , Australia, Belgium, Denmark, Egypt, France, Germany, Hong Kong, Italy, Japan, Korea, Republic of Netherlands, Norway, Saudi Arabia, Switzerland, United Kingdom, and Viet Nam. | Open, Classified |  |  | 24 | Computer
Z-2649-2014 | 69067 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | AXIOM Luminos dRF system with SW VD10 or on Luminos dRF Max |  |  | August 12, 2014 | 2014 | Siemens Medical Solutions USA, Inc | It was determined that a potential malfunction may occur on the Luminos dRF system with SW VD10 or on Luminos dRF Max when the table side control is exposed to fluids. Fluid may infiltrate the tableside control panel and result in malfunction. In the very unlikely case of fluid infiltration a permanent short circuit within the control panel of the Luminos dRF / Luminos dRF Max will automatically d | Equipment maintenance | Siemens sent a Customer Safety Advisory Notice XP045/14/S, dated August 12, 2014, to direct accounts, which identified the product, problem, and action to be taken to avoid potential risk to patients and users. Customers were instructed to avoid extensive exposure of fluids to the tableside control panel. When cleaning the system, customers were instructed not to appoly excessive fluids, only use damp towels, in accordance with the product manual. Customers were asked to forward the safety notice to any of their customers and to inform Siemens of the identify of the device's new owner where possible.For questions regarding this recall call 610-219-6300. |  | 33 | 33 | Nationwide Distribution including IL VA TN PA TX IN MN NM GA MI FL CA NY ME OK MO NJ and AZ. | Terminated | May 21, 2015 | 282 | 33 | Computer
Z-0183-2015 | 69480 | No | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | Artis zeego systems, model number 10280959, with serial numbers 1000-1599 : Artis zeego is a dedicated angiography system developed for single and biplane diagnostic imaging and interventional procedures including, but not limited to, pediatric and obese patients. Procedures that can be performed include cardiac angiography, neuro angiography, general angiography, rotational angiography, multipurpose angiography and whole body radiographic/fluoroscopic procedures. |  |  | September 23, 2014 | 2014 | Siemens Medical Solutions USA, Inc | A cable inside the C-arm systems may have been routed in an inappropriate manner. Improper routing of the cable may result in increased wear over time. Without additional measures, the potential exists for a cable inside the C-arm to break which could result in restricted functionality or system failure and an ongoing procedure could be terminated. Additionally, a subgroup of affected C-arm systems will receive a replacement of the cable harness. | Device Design | Siemens sent an " Important Safety Information " letter dated September 23, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Siemens will, without charge, remedy the defect or bring the product into compliance with each applicable Federal standard in accordance with a plan to be approved by the Secretary of Health and Human Services.Our service organization will contact with you to schedule an appointment to perform the corrective action. Please feel free to contact our service organization for an earlier appointment.Please observe this safety notice and comply with the corresponding measure until the update has been fully completed. For further questions, please call (610) 219-6300. |  | 164 | 164 | Nationwide Distribution including Puerto Rico. | Terminated | May 08, 2017 | 958 | 164 | Computer
Z-2654-2014 | 69126 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | 4D Integrated Treatment Console (4DITC), versions 10 and 11. |  |  | August 18, 2014 | 2014 | Varian Medical Systems, Inc. | Due to user error, the 4DITC can improperly allow users to clear a safety interlock imposed by the 4DITC on behalf of Varian's Bar Code Conical Verification system from the BrainLab's ExacTract console when it is used with BCCV. | Software design | Varian issued an Urgent Medical Device Correction Notification letter dated August 19, 2014, to all affected sites. The letter informed users of the specific product and versions numbers as well as the issue. A list of third party systems that perform verification and communicate authorizations via ADI are included in Table 1 of the letter. The issue is only now to occur with the ExacTract positioning system in combination with a C-Series HE linac BCCV.User recommended actions are provided for users using auxiliary devices requiring ADI interface and for users performing SRS/SBRT.Varian has developed a fix for the issue and customer service representatives will be contacting all affected sites to schedule installation.A copy of the correction letter is to be kept with the most current labeling and all personnel who work in the radiotherapy department should be made aware of the content of the letter.The recall Return response form is to be completed and returned.Questions may be directed to 1-888-827-4265 for US and Canada. Contact information is also provided for other areas of the world. |  | 512 | 512 affected devices | Worldwide Distribution - USA (nationwide) and Internationally to Argentina, Brazil, Canada, Chili, China, Columbia, Czech Republic, France, Georgia, Germany, hong Kong, Hungary, India, Indonesia, Ireland, Israel, Italy, Japan, Latvia, Lithuania, Macau, Malaysia, Mexico, Netherland,s Portugal, Qutar, Russia, Singapore, Slovenia, South Korea, Spain, Sri Lanka, Switzerland, Taiwan, Thailand, Ukraine and United Kingdom. | Terminated | July 30, 2015 | 346 | 512 | Computer
Z-0871-2014 | 67243 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | ARIATELE TELEMETRY TRANSMITTER, Model 96281, with SpO2 Option C. |  |  | December 19, 2013 | 2013 | Spacelabs Healthcare, Llc | The AriaTele Model 96281 with SpO2 monitoring Option C is recalled because the AriaTele (transmitter) display and the Telemetry Central Station display may show a SpO2 value when the sensor is not connected to the patient or in certain sensor failed conditions. | Software design | On January 6, 2014, Spacelabs sent an "URGENT - MEDICAL DEVICE CORRECTION Spacelabs Healthcare AriaTele Model 96281 with SpO2 Option C" letter, dated 03 January 2014 to US consignees.On January 9, 2014, Spacelabs e-mailed the "URGENT - MEDICAL DEVICE CORRECTION Spacelabs Healthcare AriaTele Model 96281 with SpO2 Option C" letter (translated as necessary), dated 03 January 2014, to foreign consignees.The letter described product, problem and actions to be taken. The Consignees where instructed to follow the "Advise on what action the clinician takes" on the letter.Spacelabs Healthcare will contact you at the earliest possible date to schedule a convenient tie for Spacelabs to correct all of your facility's affected transmitters (software update) at no cost.If you have any questions about this corrective action program, please contact Spacelabs Healthcare at 1-800-522-7025 and select 2 for Technical Support. |  | 72933 | 1057 units (988 in the US and 69 outside of the US). | Worldwide Distribution: US (nationwide) including states of: AZ, CA, FL, GA, IA, IL, LA, MA, ME, MI, MO, MS, NY, PA, SC, TN, TX, VA, and Puerto Rico; and Internationally to: Canada, Chile, and Panama. | Terminated | June 12, 2014 | 175 | 72933 | Computer
Z-0623-2014 | 67050 | No | GET | Surgical instrument motors and accessories/attachments. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Aesculap Hi-Line XXS handpiece |  |  | December 09, 2013 | 2013 | Aesculap, Inc. | Certain lots of the Aesculap Hi Line XXS handpiece (part # GB790R) were recalled due to a possible failure of the handpiece. | Device Design | Aesculap sent an Important Correction &amp; Removal Notification letter dated December 9, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to quarantine all affected product and returned to AIC (USA). Customers were instructed that a Aesculap Sales Representative would contact them to schedule return of any affected inventory and replace their product. Customers were asked to complete the inventory sheet with the quantity being returned. If customers could not locate the product, they were to provide an explanation as to why the inventory would not be returned (discarded, etc.). Customers with questions were instructed to call 610-984-9414,For questions regarding this recall call 800-258-1946, ext 5067. |  | 14 | 14 | Worldwide Distribution - USA including TX, NY, PA, and TN. Internationally to Austria, Switzerland, Germany, Denmark, Spain, Finland, France, Hong Kong, Israel, Italy, Japan, Netherlands, Poland, Russian Federation, and Taiwan. | Terminated | December 02, 2014 | 358 | 14 | Computer
Z-0074-2015 | 69283 | No | NAY | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | EndoWrist Stapler 45 instrument for the da Vinci Si System |  |  | September 18, 2014 | 2014 | Intuitive Surgical, Inc. | Intuitive is issuing an urgent product notice to all users of the EndoWrist Stapler 45 to discontinue use due to reports of inability to remove the stapler from tissue. | Device Design | Intuitive Surgical sent an Urgent Product Notice letter dated September 19, 2014 to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to:-Immediately stop using the EndoWrist Stapler 45 instrument.-Quarantine the EndoWrist Stapler 45 instruments within your hospital. Keep the instruments quarantined until Intuitive Surgical sends you a subsequent communication with instructions regarding the quarantined product. -Ensure all affected personnel are fully informed of this notice. Forward this notice to your Risk Manager, OR Director, Purchasing Manager, Biomedical Engineering staff, and members of your medical staff who perform da Vinci Si surgery procedures using the EndoWrist Stapler 45 instrument.-Please retain a copy of this notice for your records with your da Vinci Si User Manual.If Âżyou Âżneed Âżfurther Âżinformation Âżor Âżsupport Âżconcerning Âżthis ÂżProduct ÂżNotice, Âżplease contact your Clinical Sales Representative or Intuitive Surgical Customer Service at 800-876-1310 Option Âż3 Âż(6 Âżam Âżto Âż5 Âżpm ÂżPST). Âż |  | 1034 | 1,034 Devices | US Nationwide Distribution. | Terminated | March 03, 2015 | 166 | 1034 | Computer
Z-0201-2015 | 69251 | No | GWM | Intracranial pressure monitoring device. | Neurology | Neurology | 510(k) | Integra Licox Pt02 Monitor, Rx Only, Manufacturer: Integra LifeSciences (Ireland) Limited, IDA Business and Technology Park, Sragh, Tullamore, County Offaly, Ireland, Distributed by: Integra LifeSciences Corporation, 311 Enterprise Drive, Plainsboro, NJ 08536 |  |  | August 18, 2014 | 2014 | Integra LifeSciences Corp. | Integra LifeSciences Corporation received 2 complaints that the USB port on the Licox Pt02 monitor does not consistently provide the user the ability to extract the Pt02 trend data according to the User's Manual for the device. | Device Design | Integra LifeSciences sent an "Urgent Voluntary Medical Device Correction" letter/Return Response Form dated August 18, 2014, to their affected customers via traceable courier service. The letter identified the product the problem and the action needed to be taken by the customer.Complete the attached form. If you do not have an affected monitor, check the box: I do not have any monitors with an affected Serial Number listed below.If you do have affected monitors, check the box: I do have monitors with an affected Serial Number listed below. Circle all the Serial Numbersfor the affected monitors you have and record the total quantity.Complete the other information on the form and return the form by email or fax as indicated. Keep a copy of the form for your records.When your form is received, Integra will schedule this service with your facility and provide a temporary loaner monitor to minimize any disruption this action may cause.Should you have any questions, please contact Neuro Customer Service at 1-855-532-1723. We apologize for any inconvenience this may cause and thank you for your cooperation in this effort. |  | 43 | 43 (LCX02) and 17 (LCX02R) | US Distribution including the states of : CA, DC, FL, IL, IN, MI, MN, NC, NY, PA and TX | Open, Classified |  |  | 43 | Computer
Z-2657-2014 | 69113 | No | ITI | Powered wheelchair. | Physical Medicine | Physical Medicine | 510(k) | TRx Vertical Positioner, Classification Name: Wheelchair Component - Power Positioning System. |  |  | August 28, 2014 | 2014 | Motion Concepts | Motion Concepts identified that there is a possibility of knee support failure occurring through fatigue of two attachment bolts in the Vertical Position System (TRx Vertical Positioner). | Device Design | URGENT RECALL NOTIFICATION Letters (dated 8/26/2014) were sent to customers on 8/26/14 via Federal Express. Motion Concepts is requesting that all customers that have purchased a Motion Concepts TRx VerticalPositioner system in between June 2011 - May 2014 to make arrangements for replacement of the knee support mount in accordance with the instructions enclosed in the letter. For questions, please contact us at 905-695-0134. |  | 59 | US: 59 units | US Distribution in states of: AR, CA, CT, DE, IA, ID, IL, IN, MI, MD, MO, NE, NY, OH, PA, TN, TX, and WA. | Terminated | June 03, 2015 | 279 | 59 | Computer
Z-0067-2015 | 69357 | No | FNL | AC-powered adjustable hospital bed. | General Hospital | General Hospital | 510(K) Exempt | AC Powered adjustable bed frame. |  |  | September 15, 2014 | 2014 | Primus Medical LLC | Firm received several complaints that the head deck actuator bracket had separated from the bed frame. | Device Design | Primus Medical will send a letter notifying all customers with PCB900 beds at their facility of the potential head bracket separation. |  | 1726 | 1726 | Worldwide Distribution-USA (nationwide) including the states of CA, DE, FL, GA, IL, IN, KS, KY, MA, MN,MS, NC, NM, NV, NY, OH, PA, TN, TX, and VA, and the country of Canada. | Terminated | August 10, 2017 | 1060 | 1726 | Computer
Z-0220-2015 | 69615 | No | MAX | Intervertebral body fusion device. | Orthopedic | Orthopedic | 510(k) | LANX, REF, PCR8XXX-XXX, LOT, RXONLY. Packaged within aluminum sterilization case. |  |  | March 09, 2011 | 2011 | Biomet Spine LLC. | Biomet Spine, LLC. announces voluntary recall of Durango Anchored ALIF Plate due to the set screw component potentially becoming dislodged from the plate assembly. | Device Design | Biomet notified customers by phone on March 9, 2011 to return the affected product. |  | 120 | 120 | US Nationwide Distribution in the states of California, Indiana, Louisiana, and Texas, | Terminated | January 07, 2015 | 1400 | 120 | Computer
Z-2663-2014 | 69074 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | ADVIA Centaur XP Immunoassay System, Catalog Numbers/Siemens Material Numbers (SMN) 10285219, 10316507, 10317207, 10317284, 10319668, 10320757, 10323213, 10324519, 10327135, 10327836, 10328940, 10329339, |  |  | August 13, 2014 | 2014 | Siemens Healthcare Diagnostics | The firm is conducting a field correction for the ADVIA Centaur and ADVIA Centaur XP Immunoassay systems, regarding manually loaded sample racks that are improperly loaded or pushed too far on the sample entry queue during normal operation of the ADVIA Centaur and ADVIA Centaur XP Immunoassay systems. Pushing the sample racks may cause misreads of Sample IDs (SID) if the sample rack loading instructions are not followed exactly as described in the ADVIA Centaur or ADVIA Centaur XP Operators Guide. | Software design (manufacturing process) | The firm, Siemens, sent "Urgent Medical Device Correction" Letters (dated 8/13/2014) to their consignees/customers informing them of the issue and actions to be taken. Field service personnel were sent a Support Bulletin describing the issue and instructing them how to deal with customer questions. The UFSN must be used by customers as an addendum to the system operating instructions until further notice. The consignees/customers were also instructed to review this letter with your Medical Director; complete and return the attached FIELD CORRECTION EFFECTIVENESS CHECK within 30 days via fax to the Customer Care Center at (302) 631-7597; retain the letter with your laboratory records and forward the letter to those who may have received this product.If you have any questions, please contact your Siemens Customer Care Center, your local Siemens technical support representative and/or call 914-524-2868. |  | 10515 | 10,515 (Domestic: 2771; Foreign: 7744) in total | Worldwide Distribution: US (nationwide) and countries of: Argentina, Austria, Australia, Bahrain, Bangladesh, Belgium, Belgrade, Brazil, Botswana, Canada, Chile, China, Colombia, Croatia, Czech Republic, Cyprus, Denmark, Egypt, Estonia, Finland, France, Germany, Georgia, Great Britain, Hungary, Indonesia, India, Ireland, Israel, Italy, Japan, Jordan, Korea, Kazakstan, Lebanon, Libya, Lithuania, Luxembourg, Macedonia, Malaysia, Malta, Martinique, Mexico, Myanmar, New Caledonia, Netherlands, Norway, New Zealand, Pakistan, French Polynesia, Philippines, Poland, Portugal, Qatar, Romania, Russia, Saudi Arabia, Serbia, Singapore, Slovenia, Slovakia, South Africa, Spain, Switzerland, Thailand, Turkey, United Arab Emirates, Vatican, Venezuela, Vietnam, Yemen, and Zimbabwe. | Terminated | May 02, 2017 | 993 | 10515 | Computer
Z-2664-2014 | 69045 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | The Brilliance 64 and Ingenuity CT are Computed Tomography X-Ray Systems intended to produce cross-sectional images of the body by computer reconstruction of x-ray transmission data taken at different angles and planes. These devices may include signal analysis and display equipment, patient, and equipment supports, components and accessories. |  |  | September 01, 2011 | 2011 | Philips Medical Systems (Cleveland) Inc | During a Field Test (a customer external evaluation period), the customer detected a problem in the CIRS 4.0 Beta 2 software. When using the CIRS 4.0 Beta 2 software, during the recon during ready stage of reconstruction, images may be overlapped with, or superimposed on other images. Philips has disabled this feature at all sites evaluating 4.0 Beta 2 software. Philips will issue a software up | Software design | The firm, Philips, sent a Field Safety Notice (FSN 72800534) entitled "URGENT-Medical Device Correction" dated September 1, 2011 via mail on September 6, 2011 to its customers. The notice described the product, problem and actions to be taken. The customers were instructed to review the following information with all members of your staff who need to be aware of the contents of this communication and retain a copy with the equipment Instructions for Use. A Field Service Engineer will visit each site to install the corrected software.If you need any further information or support concerning this issue, please contact your local Philips representative or local Philips Healthcare office. For North America and Canada contact the Customer Care Solutions Center (1-800-722-9377, option 5: Enter Site ID or follow the prompts). |  | 6 | 6 | Worldwide Distribution: US distribution in OH and countries of: Australia, China, Belgium, Israel, and Sweden. | Terminated | July 01, 2015 | 1399 | 6 | Computer
Z-0190-2015 | 69473 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | MX 16-slice CT Scanner System and MX 16-slice CT Scanner System Phase 11 can be used as a whole body computed tomography X-ray system featuring a continuously rotating X-ray tube and detector array with multi-slice capability up to 16 slices simultaneously. The acquired X-RAY transmission data is reconstructed by computer into cross-sectional images of the body from the same axial plane taken at different angles. The system is suitable for all patients. Access Dual CT system can be used as a Head and Whole Body Computed Tomography X-ray System featuring a continuously rotating X-ray tube and detector array with multi-slice capability. The acquired X-ray transmission data is reconstructed by computer into cross-sectional images of the body from the same axial plane taken at different angles. |  |  | October 09, 2014 | 2014 | Philips Medical Systems (Cleveland) Inc | &quot;The orientation of images is displayed incorrectly. When operator chooses gantry on the right side from operation console, and the patient position is HFDL, HFDR, FFDL or FFDR, the patient position picture and text is not consistent. &quot;L&quot; and &quot;R&quot; of scan images are reversed and this is not consistent with the patient actual position. &quot;When a surview which is longe | Software Design Change | Philips is informing the consignees of this issue through a Field Safety Notice (FSN) 72800615 on October 09, 2014. The US consignees will receive this notice via certified mail and Philips Key Market representatives will distribute the FSN internationally. Additionally, A Field Service Engineer will contact each site to schedule a time to install the software update. |  | 708 | 708 | Domestic Distribution: AR, CA, FL, GA, IL, IN, KY, MI, MS, NC, ND, NJ, NV, NY, OH, PA, PR, SC, TN, and VA. Foreign distribution: Albania, Algeria, Argentina, Australia, Austria, Bangladesh, Belgium, Brazil, Chile, China, Colombia, Congo, Ecuador, Egypt, France, Germany, Ghana, Greece, Gutamala, Hungray, India, Indonesia, Ireland, Italy, Japan, Jordan, Latvia, Lybyan Arab JHamahiri, Lithuania, Malawi, Mexico, Myanmar, Namibia, Nepal, Netherlands, Oman, Palestine, Paraguay, Philippines, Poland, Portugal, Romania,Russia, Rwanda, Saudi Arabia, Senegal, South Agrica, South ZKorea, Spain, Sri Lanka, Switzerland, Tunisia, Turkey, Uganda, UAE, Venezuela, Vietnam, Zambia. | Terminated | July 02, 2015 | 266 | 708 | Computer
Z-0068-2015 | 69295 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | eFilm Workstation 4.0 and 4.0.1. eFilm Workstation is one of the components of a PACS (Picture Archiving and Communications System). |  |  | August 28, 2014 | 2014 | Merge Healthcare, Inc. | There is an issue related to eFilm Workstation versions 4.0 and 4.0.1 when having multiple studies open and utilizing the thumbnail panel to select multiple series from multiple studies may lead to the selection of an incorrect study. | Device Design | Merge sent an Important Product Information Notice letter, dated August 22, 2014 to consignees. The letter stated the issue, identified affected product, provided instructions, and told of a product update expected to be available by September 15, 2014. Customers with questions should contact Customer Service at 877-741-5369 |  | 2630 | 2,630 | Worldwide Distribution - US Nationwide and the countries of : Argentina, Australia, Canada, Chile, France, Germany, Hungary, Iceland, Italy, Lithuania, Mexico, New Zealand, Romania, Russia, South Africa, Sweden, and Yemen. | Terminated | June 01, 2017 | 1008 | 2630 | Computer
Z-2666-2014 | 69051 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | The Brilliance CT Big Bore is a whole-body Computed Tomography X-ray system intended to produce cross-sectional images of the body by computer reconstruction of X-ray transmission data from the same axial plane taken at different angles. Model Numbers 728244 and 728243. |  |  | August 31, 2012 | 2012 | Philips Medical Systems (Cleveland) Inc | Philips Healthcare received a complaint stating the system opens e-stop while sitting idle causing all movements and scan to stop. The problem has only occurred one time when the scan was out of Ready for Scan or scanning mode. | Software design | A Customer Letter (FSN 72800580), Attachment #1, will be sent to all US and Canadian consignees via certified mail. Philips Key Market representatives will distribute the letter internationally. Additionally, a Field Service Engineer will contact each site to schedule time to implement the Field Change Order. |  | 257 | 257 | Nationwide Distribution incuding states of: AL, AR, AZ, CA, CO, CT, FL, GA, HI, IA, IL, IN, KS, KY, MA, MD, MI, MN, MO, MS, NC, ND, NH, NJ, NM, NY, OH, OK, OR, PA, TN, TX, VA, WA, and WI. | Terminated | July 01, 2015 | 1034 | 257 | Computer
Z-0151-2015 | 69400 | No | FNL | AC-powered adjustable hospital bed. | General Hospital | General Hospital | 510(K) Exempt | Joerns EasyCare Bed |  |  | September 24, 2014 | 2014 | Joerns Healthcare | The Easy-Care Care-Lock feature installed at the foot of the bed may potentially fail to stay engaged (locked) after the Care-Lock System has been engaged by stepping on the red (locked) pedal. | Device Design | Joerns sent an Urgent: Medical Device Correction letter dated September 24, 2014 to affected customers. The letter identified the affected product, problem and actions to be taken. The firm will provide the customers with a temporary solution kit to safegard the bed until a replacement Care-Lock assembly can be shipped and installed. Customers were instructed to complete the response form. For questions call 800-826-0270. |  | 3135 | 3135 | Worldwide Distribution: US (nationwide) and country of: Canada. | Terminated | February 01, 2016 | 495 | 3135 | Computer
Z-0737-2014 | 66788 | Yes | CAC | Autotransfusion apparatus. | Anesthesiology | Anesthesiology | 510(k) | Atrium Ocean Water Seal Chest Drain ATS (Autotransfusion) Blood Recovery System with stopcock Blood Recovery System; Product Code: 2050-300. |  |  | November 19, 2013 | 2013 | Atrium Medical Corporation | Chest Drain tubing of of the ATS Blood recovery may leak or disconnect | Device Design | Atrium Medical issued notification letters to direct consignees by UPS next day on 19 November 2013, and indirect consignees (distributors customers) by UPS next day or USPS priority (for customers with a PO Box) on 27 November 2013. The Letter explain the issue and the steps that should be taken. The customer will be offered replacement with unaffected product or credit. The firm describes Procedures for Emergency Use. The firm identifies potential product substitutions. These products are the same drainage product without the auto-transfusion line.International Distributors. International distributors provided with the Recall Authorization' package and the customer letter.Accounts are requested to complete the Recall Reply Form acknowledging receipt of the notification.Contact Atrium Medical Customer Service at 1-800-370-7899, Monday through Friday between 9:00 am to 5:00 pm. |  | 6684 | 6,684 units | Worldwide Distribution-USA (nationwide) and the countries of Australia, Canada, Austria, Ireland, Romania, South Africa, Colombia, Lebanon Netherlands, Jordan, Russia, and Taiwan. | Terminated | June 11, 2015 | 569 | 71990 | Computer
Z-2665-2014 | 69129 | [6 empty] | Siemens CentraLink Data Management System, software version 14. |  |  | August 19, 2014 | 2014 | Siemens Healthcare Diagnostics, Inc. | Siemens Healthcare Diagnostics has identified a task scheduler issue with the CentraLink&quot; Data Management System, software version 14. The issue may cause the software to stop executing commands, including uploading validated results to the LIS. The issue is related to an internal software timer that overflows after 24 days. Automated rules triggered by task schedulers or by actions in the CentraLink user interface may also be affected. As each customer configuration is unique, in addition to the delay in uploading of validated results, the specific consequences of discontinued command execution will vary. | Software design | An Urgent Medical Device Correction (UMDC #10819364) letter was issued on August 19, 2014 to all affected customers. The letter instructed customers as follows:Siemens Healthcare Diagnostics is conducting a field correction for the CentraLink" Data Management System. To prevent the issue from occurring, log out of user sessions and restart the CentraLink services and translators as soon as it is convenient, then periodically restart going forward. Any restart frequency that occurs more often than every 24 days is acceptable and will prevent the issue from occurring. Refer to the Additional Information section for detailed instructions. |  | 858 | 858 | Worldwide Distribution-USA (nationwide) including Puerto Rico and the countries of Argentina, Australia, Austria, Bangladesh, Belgium, Brazil, Canada, Chile, China, Colombia, Denmark, Egypt, Finland, France, Germany, Greece, Hungary, India, Ireland, Israel, Italy, Lebanon, Lithuania, Luxemburg, Malaysia, Mexico, Netherland, New Zealand Norway, Pakistan, Portugal, Peru, Qatar, Russia, Saudi Arabia, Singapore Slovakia, Slovenia, South Africa, South Korea, Spain, Sweden, Switzerland, Taiwan, Thailand United Arab Emirates, United Kingdom, Uruguay, Zambia. | Terminated | May 12, 2017 | 997 | 858 | Computer
Z-0168-2015 | 69583 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | COMPASS, Model No. SW Version 3.1, catalogue number CS10-100 (medical Linear accelerator) radiological |  |  | September 25, 2014 | 2014 | Iba Dosimetry Gmbh | Error in the software. A deviation between reconstructed and planned dose distribution may not be detected prior to treatment and this can result in an over or under-estimation of the pretreatment delivered dose. | Software design | The field safety notice including a confirmation of receipt was sent to the affected hospitals; by email and UPS letter on 09/25/2014. |  | 124 | 124 units | Worldwide Distribution-Distributed in the US including Guam and the countries of Germany, Austria, France, UK, Poland, Finland, The Netherlands, Spain, Turkey, Russia, and South Africa, Algeria, Mexico, Brazil, Venezuela, Columbia, Panama, China, Japan Philippines, Korea, India, Singapore, Thailand, Taiwan, Hong Kong, and Malaysia. | Terminated | January 26, 2015 | 123 | 124 | Computer
Z-0169-2015 | 69463 | No | GEX | Laser surgical instrument for use in general and plastic surgery and in dermatology. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Elite XPL (Pulsed Light) Handpiece for the Cynosure Elite MPX Laser System |  |  | October 09, 2014 | 2014 | Cynosure, Inc. | If the laser system is turned on and the user adjusts, installs or removes the XPL pulsed light handpiece, there is a potential risk of an electrical shock | Device Design | Cynosure sent an Urgent Medical Device Recall Notification dated October 9, 2014, to all affected customers. They were also notified by phone. The letter identified the product the problem and the action needed to be taken by the customer.There are two methods for returning the Elite XPL handpieces. If the handpiece is NOT currently connected to the laser console, the customer will be asked to return it to the Cynosure office. If the handpiece is not connected there is no risk.If the handpiece is currently connected to the laser console, the customer is instructed to immediately discontinue use of the laser system and to turn off the system immediately. They are also being instructed to leave the handpiece attached to the console and to not touch it. A Cynosure Field service engineer will then come to the office and remove the handpiece and return it to the Cynosure factory. In addition, the Cynosure Field Service engineer will install an interlock into the connector receptacle on the laser console. This interlock plug is an existing component that is used on systems that do not have an IPL handpiece in place.If your XPL pulsed light handpiece is not attached to the laser system please call Cynosure customer service (1-877-222-6822) to arrange for the handpiece to be returned and repaired. |  | 322 | 322 units | Worldwide Distribution - Nationwide Distribution and the countries of : Australia, Canada, France, Germany, Italy, Netherlands, Ukraine and UK. | Terminated | March 26, 2015 | 168 | 322 | Computer
Z-0105-2015 | 69348 | No | HGM | Perinatal monitoring system and accessories. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | Philips Avalon Monitors with software revision J.30.58: |  |  | October 03, 2014 | 2014 | Philips Medical Systems, Inc. | Philips Intellivue and Avalon Fetal Monitors in time-synchronized automatic/sequence mode, the NBP automatic measurement series is stopped | Software Design Change | Philips Healthcare issued on 10/3/14 the Urgent Medical Device Correction notification/Field Safety Notice. The letter informs customers of the issue, identifies details of the units affected, gives instructions on actions to be taken by the customer and identifies what action Philips plans to take to remedy the issue.The correction will consist of a software upgrade free of charge.A Philips Healthcare representative will contact customers with affected devices to arrange for service. Philips is asking customers to follow the Action to be Taken by Customer/User section of the Urgent Medical Device Correction notification/Field Safety Notice:During the interim period until the SW is upgraded please make sure, if the NBP repetition interval needs to be increased (e.g. from 5 to 15 minutes), to restart the automatic NBP measurement.Please review this information with all members of your staff who need to be aware of the contents of this communication. It is important to understand the implications of this communication. |  | 118 | 118 units | Worldwide Distribution-USA (nationwide) and the countries of AUSTRALIA AUSTRIA BAHRAIN BELGIUM CHINA COLOMBIA CZECH REPUBLIC DENMARK FINLAND FRANCE GERMANY HONG KONG INDIA IRELAND ISRAEL ITALY JAPAN KAZAKHSTAN NETHERLANDS NEW ZEALAND NICARAGUA NIGERIA NORWAY OMAN POLAND ROMANIA RUSSIA SINGAPORE SLOVENIA SOUTH AFRICA SOUTH KOREA SPAIN SWEDEN SWITZERLAND TAIWAN THAILAND TURKEY UKRAINE and UNITED KINGDOM. | Open, Classified |  |  | 7705 | Computer
Z-0773-2014 | 64740 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Imagecast PACS is an image management system intended to be used by trained professionals, including but not limited to physicians, nurses and medical technicians. The system is used with general purpose computing hardware to acquire, transmit, process and store images and data throughout a clinical environment. Data and images are acquired through DICOM compliant imaging devices and modalities. Lossy compressed mammographic images and digitized film screen images must not be reviewed for primary image and interpretations. Mammongraphic may only be interpreted using an FDA approved monitor that offers at least 5 Mpixel resolution and meets other technical specifications reviewed and accepted by the FDA. |  |  | March 04, 2011 | 2011 | Ge Healthcare It | In Imagecast PACS versions prior to 3.6.51.01, the estimated magnification factor (0018,1114) was not accounted for. These versions of Imagecast PACS look for values to be present in the following tags to calculate a measurement from pixels to millimeters (mm): (0020,0030) Pixel Spacing and, (0018,1164) Imager Pixel Spacing. Because the estimated magnification factor is not used, distance meas | Software design | Important Product Information Letters dated March 4, 2011 were sent to all customers that received the defective software. The letters advised customers to upgrade their PACS software version to Imagecast 3.6. The software upgrade can be requested from GE Support via the eService portal or by telephone at 1-800-437-1171. If the customer elects to not upgrade the software, instructions were provided for a new system calibration procedure. |  | 43 | 43 units | Nationwide Distribution: US including states of: AL, AZ, CT, FL, GA, IA, IN, LA, MA, MI, MN, MO, NC, NH, NJ, NY, OH, PA, TX, VA, WI, WV + District of Columbia. | Open, Classified |  |  | 43 | Computer
Z-2587-2014 | 68618 | No | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | AXIOM Artis or Artis zee system |  |  | June 05, 2014 | 2014 | Siemens Medical Solutions USA, Inc | There is a potential problem with the AXIOM Artis or Artis zee system in connection with the hand switch operated at the patient table which may result in uncontrolled release of an x-ray image. | Device Design | Siemens sent an Important Safety Information letter dated June 5, 2014, to affected customers. The letter described the affected product, problem, and actions to be taken. Customers were instructed to observe this safety notice and comply with the corresponding measures until the update has been fully completed. |  | 1067 | 1067 | US Nationwide Distribution. | Terminated | April 06, 2015 | 305 | 1067 | Computer
Z-0805-2014 | 67013 | No | LNX |  |  | General Hospital | 510(k) | ORCHESTRA/ORCHESTRA PLUS Programmer |  |  | November 20, 2013 | 2013 | Sorin Group Italia S.r.l. | Sorin has voluntarily issued a notification to physicians related to the overestimation of the residual longevity displayed by its programmer (ORCHESTRA or ORCHESTRA PLUS) during a follow-up exam of patients implanted with a REPLY or ESPRIT pacemaker. | Software design | On Nov 20 and Nov 25, 2013, the Dear DoctorLetter was sent by certified mail to 1,310 physicians who are following 13,003 patientsimplanted with a REPLU or ESPRIT pacemakers in the US. Sorin provided the following recommendations to the physicians: You should consider checking the battery impedance of the last follow-up exam. In casethe battery impedance is greater than or equal to 3.5kÂż, a follow-up visit must bescheduled within a maximum of 6 months from the last follow-up visit. When pacemaker operation is checked by the simple application of a magnet, a magnetrate less than 95 min-1 should trigger a follow-up exam in the pacemaker centre. As a general rule, a maximum of 6 month follow-up interval when the battery impedancebecomes greater than or equal to 3.5 kÂż. This recommendation should also be followedsubsequent to the installation of the new programmer software version. |  | 1718 | 1,718 | Nationwide. | Terminated | September 08, 2014 | 292 | 1718 | Computer
Z-0787-2014 | 67126 | No | JAI | Powered radiation therapy patient support assembly. | Radiology | Radiology | 510(k) | PRIMUS, ONCOR, ARTISTE Linac systems with component 550 TxT Treatment Table. |  |  | December 05, 2013 | 2013 | Siemens Medical Solutions USA, Inc | Update to correct the possibility for patients or operators to have their fingers trapped or pinched between the table top and the bottom frame of the 550 TxT Treatment Table. Combined with this update instruction is a second complaint issue related to an incident in which a cable inside of the TxT" Treatment Table has been damaged due to contacting a moving part. The resulting short-circuit cau | Device Design | The firm, Siemens, sent a "Medical Device Correction" letter to affected users/customers starting on December 5, 2013 via Federal express. The letter described the product, problem and actions to be taken. Siemens has arranged for modification of the affected 500 TxT Treatment Table. A service representative will modify the mounts of the TT-A, TT-D, TT-M, or TT-s table tops ( Pleases note that this modification only applies to the Siemens provided table tops TT-A, TT-D, TT-M, and TT-S.). In addition, the Siemens service representative will modify the cable routing inside of the patient table base. The customers were instructed to include this Field Safety Notice in your Digital Linear Accelerator System Owner manual chapter 'Safety Advisory Letters' where it should remain.If you have any questions, please contact the Senior Manager, Regulatory Affairs at 1-925-293-5442 or email: christine.dunbar@siemens.com. |  | 305 | 305 active devices. | Worldwide Distribution: US (nationwide) and internationally to: Canada, Italy, P.R. China, Argentina, Pakistan, South Africa, Egypt, Iran, Vietnam, Indonesia, Brazil, Hungary, Bulgaria, India, Spain, Switzerland, Trinidad, Tobago, Italy, Algeria, Croatia, Turkey, Japan, Germany, Portugal, Poland, France, Sri Lanka, Russian Federation, France, Lebanon, Slovakia, Ukraine, Pakistan, and Korean Republic. | Terminated | January 30, 2015 | 421 | 305 | Computer
Z-0107-2015 | 69260 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Mindray V21, Size: 515 mm X 335 mm X 685 mm, N.W.: 8 kg, G.W.: 11 kg, Qty:1 |  |  | July 31, 2014 | 2014 | Mindray DS USA, Inc. dba Mindray North America | Mindray DS USA Inc. initiated a voluntary field corrective action for V-Series Patient Monitor because there is an issue with the V-Series Drug Calculator function. | Software design | Mindray DS USA, Inc. sent a Recall Correction Action Letter to their customers via certified mail with return receipt on July 31, 2014. |  | 2247 | 2,247 units | Worldwide Distribution: US (nationwide) and internationally to: Austalia, Canada, Columbia, Germany, India, Indonesia, Ireland, Israel, Kuwait, Malaysia, New Zealand, Pakistan, Russian Federation, Saudi Arabia, South Africa, Sri Lanka, Thailand, Turkey, Ukraine, United Arab Emirates, United Kingdom, and Venezula. | Open, Classified |  |  | 2247 | Computer
Z-0786-2014 | 67153 | No | KPQ | Radiation therapy simulation system. | Radiology | Radiology | 510(k) | SIMVIEW NT and SIMVIEW NT with IDI: |  |  | December 11, 2013 | 2013 | Siemens Medical Solutions USA, Inc | In a very rare occurrence an intended movement of the SimView NT simulation system components could not be stopped with usual hand control or the collision detection functionality. | Device Design | Siemens sent an Urgent Field Safety Notice to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. What preventive measures can the user take?To stop the movement in this case, push the emergency button / stop button at control console or push the in-room emergency button. The emergency button/ stop button functionality is described in the User Manual.There is no preventive measure to avoid this situation. This letter is intended to bring this potential safety issue to your awareness and to ensure that your personnel are trained about the use of the emergency button / stop button inside the simulation room and on the control console. Always monitor any movement of the simulator components.The Update Instruction containing the Urgent Field Safety Notice will be delivered by the Siemens Customer Service Group, as per their procedureSiemens is working on a solution to mitigate the described situation. You will be informed as soon as the update distributing this solution is available.Please include this Field Safety Notice in your SimView NT Simulator System Owner Manual chapter Safety Advisory Letters where it should remain.In the interests of safety, we ask that you perform the described preventive measures and inform all affected personnel immediately.The relevant National Competent Authority has been informed of this Customer Safety Advisory Notice.We regret any inconvenience that this may cause, and we thank you in advance for your understanding. For further questions please call (925) 602-8157. |  | 65 | 65 active devices | Class II Recall - Worldwide Distribution - US Distribution including the states of California, Louisiana, Massachusetts, Pennsylvania, New York, Kentucky and Michigan., and the countries of Bulgaria, Germany, Belarus, Russian Fed., P.R. China, India, Republic Korea, Germany, Costa Rica, Argentina, Portugal, Greece,Bulgaria, Vietnam, Hungary, Italy, Palestinian Ter, Hungary, Jordan, EgyptPoland, Greece, Curacao, St Eus, Malaysia, Egypt, Malaysia, South AfricaTurkey, Thailand, Qatar and Iraq. | Terminated | March 31, 2014 | 110 | 65 | Computer
Z-2738-2014 | 69122 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | SPECTRUM Pump, Model No. 35700BAX. |  |  | September 03, 2014 | 2014 | Baxter Healthcare Corp. | One Service Technician may not have correctly serviced specific Sigma Spectrum Infusion Pumps according to established procedures during the time period of 5/5/2014 through 6/3/2014. | Equipment maintenance | Initial notification was initiated via telephone call by the Baxter Medina Service Center on 8/28/14 to all affected Sigma SPECTRUM Infusion Pump consignees. URGENT DEVICE CORRECTION Letters (dated 9/03/2014) were sent to the consignees via USPS first class mail on 9/03/14 formally informing them of the recall. The letters identified the affected product, the description of the issue, the hazard involved, as well as, the actions to be taken by customer. Customers are being instructed to immediately remove the pump from use and return the pump to Baxter for inspection. The firm will provide replacements. Customers are to contact Baxter Healthcare Medina at 800-356-3454 for technical questions regarding the letter. |  | 58 | USA: 56 units, Canada: 2 units | Worldwide Distribution -- USA and Canada. | Terminated | March 16, 2017 | 925 | 58 | Computer
Z-0222-2015 | 69531 | No | HQF | Ophthalmic laser. | Ophthalmic | Ophthalmic | 510(k) | MC-500 Multicolor Laser Photocoagulator using software version 2.20. |  |  | May 24, 2013 | 2013 | Nidek Inc | Retrospective review found that a Engineering Change Order initiated in May 2013 for a memory function problem was not reported as required. | Software design | Nidek Engineers visited all affected sites to install new software version.For questions regarding this recall call 510-353-7785. |  | 5 | 5 | US Distribution in the states of: CA, NE, and NY. | Terminated | November 17, 2014 | 542 | 5 | Computer
Z-0491-2015 | 69541 | No | LFL |  |  | General &amp; Plastic Surgery | 510(k) | Integra Selector Kit |  |  | September 29, 2014 | 2014 | Integra LifeSciences Corp. | Some specific lots of Selector Tubing may potentially leak irrigation fluid during use if not placed on the Selector handpiece in a specific manner. | Device Design | Integra sent an 'Urgent- Medical Device Correction' letters/acknowledgement forms dated 9/29/2014 by traceable courier service. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to review their inventory for the affected product, complete the attached form. Complete the other information on the form and return the form by email or fax.Customers with questions should contact Customer Service at 1-855-532-1723. |  | 6358 | 6,358 tubing kits | Worldwide Distribution - USA (nationwide) and Internationally to Algeria, Australia, Bulgaria, Canada, Germany, Great Britain, Greece, Hungary, Indonesia, Israel, Italy, Japan, Mexico, Portugal, Spain, Sweden, Thailand, Turkey, and Wales. | Terminated | May 27, 2015 | 240 | 6358 | Computer
Z-2732-2014 | 69132 | No | LTI |  |  | Gastroenterology/Urology | PMA | LAP-BAND AP SYSTEM AP Large with Access Port, Allergan, LAP-BAND AP SYSTEM Access Port I. Product was distributed for both the AP Large and AP Standard sizes. |  |  | August 15, 2014 | 2014 | Apollo Endosurgery Inc | Product was distributed past its expiration date. | Software Manufacturing/Software Deployment | Firm called its consignees and advised them to either return or destroy the devices which had not already been implanted. |  | 10 | 10 | Distributed in the states of California, Delaware, Indiana, Missouri, New Jersey, New York, and Oregon. | Terminated | February 01, 2016 | 535 | 10 | Computer
Z-0217-2015 | 69565 | No | DWF | Cardiopulmonary bypass vascular catheter, cannula, or tubing. | Cardiovascular | Cardiovascular | 510(k) | Edwards Lifesciences Vent Catheters, . Packaged sterile and non-pyrogenic in a sealed, peel-type pouch. Model Numbers: E061, E060, PE062, EM012 |  |  | October 08, 2014 | 2014 | Edwards Lifesciences, LLC | Edwards Lifesciences is conducting a field action on the Vent Catheters due to the potential for disruption of the white plastic coating on the stylet. | Device Design | Edward Lifesciences sent an Urgent Medical Device Safety Notice date October 8, 2014, via Fed-Ex to all affected customers. On 3/4/15, letters were sent to surgeons. As part of the customer instructions in the safely notice, customers were requested to forward copies of the safely alert to any organization where the potentially affected devices have been transferred. No product is being returned. The distribution of the device has not been stopped as a result of this field correction. No additional production lots will be released for distribution until the IFU has been revised. Customers with questions were instructed to call (800) 424-3278.For questions regarding this recall call 801-553-7531. |  | 157305 | 157,305 | Worldwide Distribution - USA including Alabama, Arkansas, Arizona, California, Colorado, Connecticut, Florida, Georgia, Hawaii, Iowa, Idaho, Illinois, Indiana, Kansas, Kentucky, Louisiana, Massachusetts, Maryland, Maine, Michigan, Missouri, Mississippi, Montana, North Carolina, Nebraska, New Hampshire, New Jersey, Nevada, New York, Ohio, Oklahoma, Oregon, Pennsylvania, Rhode Island, South Carolina, South Dakota, Tennessee, Texas, Utah, Virginia, Vermont, Washington, Wisconsin, West Virginia. US Territory: Puerto Rico and Internationally to Australia, Bangladesh, Brazil, Canada, Chile, China, Colombia, Hong Kong, India, Japan, Malaysia, Mexico, Netherlands, Singapore, South Korea, Sri Lanka, Taiwan, and Venezuela. | Open, Classified |  |  | 157305 | Computer
Z-0042-2015 | 69377 | No | MWI | Cardiac monitor (including cardiotachometer and rate alarm). | Cardiovascular | Cardiovascular | 510(k) | McKesson Cardiology" Hemo intended for complete physiological/hemodynamic monitoring, clinical data acquisition, medical image and data processing, and analytical assessment. |  |  | August 27, 2014 | 2014 | Mckesson Information Solutions LLC | Software Error: The McKesson Cardiology Hemo calculation section incorrectly converts the Hemoglobin value before it is utilized in the applicable formula calculations. | Software design (manufacturing process) | On August 27, 2014, McKesson initiated a correction via a Field Safety Notice to all affected domestic customers. McKesson Israel will be responsible for contacting foreign consignees. |  | 272 | 272 devices | Worldwide distribution: US (nationwide) including states of: AR, CA, FL, DE, GA, HI, IL, IN, KS, LA, MA, MD, MN, MO, MS, MT, NJ, NV, OK, PA, SC, TNTX, , UT, WA; and countries of: Australia, Canada, Israel and the United Kingdom. | Terminated | May 18, 2016 | 630 | 272 | Computer
Z-0137-2015 | 69296 | No | LZG | Infusion pump. | General Hospital | General Hospital | 510(k) | Animas Vibe Insulin Infusion Pump and System. |  |  | October 25, 2012 | 2012 | Animas Corporation | Animas determined that keypad wear over time was contributing to an increased complaint rate associated with damaged keypads on the Vibe Insulin Infusion Pump and System. Animas has implemented an improvement to the keypad related to the material used to reduce susceptibility to damage and increase durability. | Device Design | Animas initiated the notification to Distributors by letter, dated July 2012, and Patients and Health Care Professionals were notified by letter in September 2012. The delay in notifying patients and health care professionals was driven by the need to have sufficient inventory of replacement product on hand in order to meet anticipated demand before sending the letters. |  | 1938 | 1938 | Worldwide Distribution in the countries of France, Italy, Sweden, and UK.. | Terminated | June 29, 2015 | 977 | 1938 | Computer
Z-0221-2015 | 69656 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | MOSAIQ |  |  | November 03, 2014 | 2014 | Elekta, Inc. | A problem can exist in MOSAIQ resulting in the display of incorrect numeric data due to a dose rounding error on printed reports. | Software design | Elekta sent an Urgent Important Field Safety Notice letter to alert customers of the issue. This notice was distributed on November 3, 2014. A fix to the defect has been developed and has been released by Product Creation. For question call 408-830-8000. |  | 15 | 15 | US Nationwide Distribution in the states of AZ. CO, IL, MD, MO, NY, OK, OR, PA, RI, TX, VA | Terminated | July 27, 2015 | 266 | 15 | Computer
Z-0450-2015 | 69783 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | BrightView product code: 882478 |  |  | November 04, 2014 | 2014 | Philips Medical Systems (Cleveland) Inc | Software issues | Software design | Philips Healthcare sent an URGENT - Field Safety Notice Medical Device Correction letter dated November 4, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.This Field Safety Notice 88200485 is intended to inform you about:" what the issues are and under what circumstances they can occur" the actions that should be taken by the customer / user in order to prevent risks for patients or users" the actions planned by Philips to correct these issues.If you need any further information or support concerning these issues, please contact your local Philips representative:For North America contact the Customer Care Solutions Center (1-800-722-9377: follow the prompts).Philips apologizes for any inconveniences caused by this problem. |  | 85 | 85 | Worldwide Distribution - USA including the states of AR, CA, FL, GA, ID, LA, ME, MI, MN, MO, MS, NJ, NY, OH, OK, PA, and TX. Government Facilities: AR, FL, MN, OH, and OK. and the countries of: Canada, China, Denmark, Germany, Italy, South Korea, Spain, Sweden, Switzerland, Taiwan, and United Kingdom. | Open, Classified |  |  | 85 | Computer
Z-0784-2014 | 66846 | No | KRO | Knee joint femorotibial metal/polymer constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Zimmer Segmental Systems; Polyethylene Inserts, Size C. |  |  | December 19, 2013 | 2013 | Zimmer, Inc. | After evaluation of 13 complaints(0.4% rate) for hyper-extension of the Zimmer Segmental System, Zimmer is updating the related surgical technique and instructions for use to provide additional instructions and patient conditions that may place excessive loading on the polyethylene insert. | Device Design | Depth of recall" Notifications were sent to customers, surgeons and risk managers, which use the affected product. Distributors were also informed of the notifications sent to surgeons and risk managers. The customers were identified by performing an invoice search and distribution history for the Polyethylene Inserts (00-5850-012-95 and 00-5850-013-95) since product launch. Recall Communication" All distributors were notified via electronic mail. Distributors that have received affected inventory were also notified via courier. Hospital risk managers and surgeons were also notified via courier.On 12/19/13, hospital risk managers and surgeons were provided with a notification identifying the issue and their responsibilities. These responsibilities include ensuring the affected personnel are aware of the contents of the notification. On 12/19/13, Distributors were sent a notification identifying the issue and to provide a copy of the notification sent to surgeons and risk managers. The distributors will also be instructed as needed to assist with identifying a comprehensive list of surgeons. " Communications outside of the United States will occur approximately two weeks after the United States communications.Effectiveness ChecksZimmer will conduct effectiveness checks in the following way:"Surgeons/Hospitalso100% of the FedEx notifications will be accounted for.oSecond and third notification attempts will be made.oFailed delivery attempts following the third attempt will result in the account being deemed as nonresponsive. |  | 5305 | 5305 distribution events | Worldwide Distribution. USA nationwide and Germany, with further distribution into EU countries. | Terminated | February 24, 2015 | 432 | 5305 | Computer
Z-2733-2014 | 69139 | No | HQF | Ophthalmic laser. | Ophthalmic | Ophthalmic | 510(k) | Array LaserLink, Manufactured by Lumenis, |  |  | August 15, 2014 | 2014 | Lumenis, Inc. | Lumenis initiated a field-correction for the Array Laser Link&quot;, GA-0006700 (SN XXYYZZ) because the system may project an erroneous pattern display on the retina, which is different than the desired pattern. | Software design | Lumenis sent a Safety Advisory Notice dated July 15, 2014, to all affected consignees. The letter identified the product, the problem, and the action to be taken by the consignee. Consignees were instructed that Lumenis would contact consignees to schedule a product update, to be completed at a convenient time. For questions consignees should call 801-656-2549.For questions regarding this recall call 801-656-2690. |  | 16 | 16 | Worldwide Distribution - USA (nationwide) and Internationally to JAPAN, CHINA, and CANADA. | Terminated | October 21, 2014 | 67 | 16 | Computer
Z-0039-2015 | 69146 | No | OMP | Powered suction pump. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | RENASYS EZ/ RENASYS EZ Plus 250 mL canister with Solidifier |  |  | July 18, 2014 | 2014 | Smith &amp; Nephew, Inc. | RENASYS EZ Canisters exhibiting visible evidence of deformation of the inlet port on the canister lid. | Device Design | Smith &amp; Nephew sent an Urgent Medical Device Correction Notice dated July 15, 2014 to all US customers via Federal Express delivery. The letter identified the affected product, problem and actions to be taken. The firm instructed customers to discard the affected product and replace with a new one from an unaffected lot. For questions contact 727-399-3785. |  | 8650 | 19,499 = (10,849 (Item 66800058) &amp; 8,650 (Item 66800913)) | Worldwide Distribution - US Nationwide in the states of FL, NC, CA, MO, MA, TX, NY, OH, CO, VA, IL, NJ, TN, NV, MN including Puerto Rico and countries ofCanada, Mexico, Dubai, Argentina, Australia, Colombia, Germany, Hong Kong, Jordon, Malaysia, Thailand, Tunisia, New Zealand, South Africa, Shanghai, Brazil, Saudi Arabia and Lebanon. | Terminated | August 05, 2015 | 383 | 55661 | Computer
Z-0780-2014 | 67078 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | XiO RTP System |  |  | November 25, 2013 | 2013 | Elekta, Inc. | When the number of fractions is changed for a proton spot beam, the MU are recomputed and rounded if necessary, but the dose is not recalculated as it should be even after saving and re-opening the plan. | Software Design Change | An Important Safety Notice, LINXIO0017 Proton Sport Beams, Changes in Fractionation not Updating Dose, is in the process of being distributed to affected customers. The release date for this notice is November 25, 2013. The notice included a Confirmation of Receipt which was to be returned to Elekta.For further questions please call (404) 993-5886. |  | 13 | 13 | Worldwide Distribution - US Distribution including the states of IL, IN, LA, NJ, TX, WA and WI., and the countries of Belgium, Germany, Italy and Netherlands. | Terminated | May 13, 2014 | 169 | 13 | Computer
Z-0002-2015 | 69258 | No | KPR | Stationary x-ray system. | Radiology | Radiology | 510(k) | Siemens Ysio Max system |  |  | September 02, 2014 | 2014 | Siemens Medical Solutions USA, Inc | Two malfunctions posing potential risk to patients were identified with Siemens Ysio Max, AXIOM Luminos dRF Max, Uroskop Omnia Max, and Luminos Agile Max systems with a specific set of serial numbers that will cause lost images due to an automatic Radiology Information System (RIS) worklist update, thus causing examinations to be repeated. Image label may also be displayed incorrectly after images | Software design | Siemens sent a Safety Advisory Notice, dated September 2, 2014, to all affected customers. The Notice identified the product, the problem, and the action to be taken by the customer. Customers were instructed to disable the automatic RIS worklist update and perform this procedure manually. Customers should disable the Measurement Field Display and Collimator Display by setting the display time for both setting to 0 sec (if editable). Customers should forward the safety notice to the new owner and inform Siemens of the new owner where possible. Customers with questions should call 610-219-2119 or 610-219-4834.For questions regarding this recall call 610-219-6300. |  | 10 | 10 | Nationwide Distribution - US including UT, MO, CO, MA, MN, MS, OH, SC, MI, and NY. | Terminated | February 24, 2015 | 175 | 18 | Computer
Z-0007-2015 | 69094 | No | DIP | Diphenylhydantoin test system. | Toxicology | Toxicology | 510(k) | VITROS Chemistry Products PHYT (Phenytoin) Slides, REF/Product Code 829 8671, IVD. |  |  | August 20, 2014 | 2014 | Ortho-Clinical Diagnostics | Ortho Clinical Diagnostics identified a potential for biased results to be generated when using VITROS PHYT Slides, Lot 2613-0150-4913. Some positively biased results were identified during testing of a different VITROS PHYT Slide lot; the affected slides were not released for distribution. Since Lot 2613-0150-4913 was manufactured in the same event, there is the potential for slides from this lot to also be affected. | Software Manufacturing/Software Deployment | On 8/20/14, an URGENT PRODUCT CORRECTION NOTIFICATION Letter (CL2014-222, dated 8/20/14) was sent via FedEx overnight courier and/or ORTHO PLUS e-Communications and/or US Postal Service Priority Mail (for PO Boxes only) to US consignees (OCD Direct, Drop-Ship and US Federal Government) to inform them of the issue and advise them to discontinue use of the affected lot(s) and to inform their Laboratory Medical Director if they suspect that a previously reported result may have been affected. One foreign affiliate was notified on 8/20/14. One distributor was notified on 8/20/14. The customers were advised to: - Immediately discontinue using and discard VITROS Chemistry Products PHYT slide Lot 2613-0150-4913. Review VITROS PHYT results obtained using Lot 2613-0150-4913 and discuss any concerns they may have regarding previously reported results with their Laboratory Medical Director or with the requesting physician. Complete the attached Confirmation of Receipt form no later than 8/28/14. Forward the information in this notification if they have distributed this product outside of their facility facility. |  | 349 | Domestic: 339 sales units; Foreign: 10 sales units | Distributed USA (nationwide) and in Singapore. | Terminated | April 11, 2017 | 965 | 349 | Computer
Z-0529-2014 | 66833 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | EOS System X- ray beam |  |  | November 07, 2013 | 2013 | EOS Imaging | EOS imaging discovered during production internal testing that the X-ray beam collimation set up by the operator at the user interface of the acquisition workstation is different from the actual collimation ofthe X-ray beam operated by the EOS system during the X-ray acquisition. | Device Design | CDRH approves the following conditions for EOS Imaging SA:1. Notification of all dealers and purchasers is to be made within 15 working days of receipt of this letter in the manner specified in 21 CFR 1003.21 and 1003.22. This office and the Food and Drug Administration (FDA) district office coordinator notes below what are to be included in the notification. The draft letter provided on November 28, 2013, is deficient for the following reasons:a. It does not provide a statement that all repairs will be made withoutcharge (21 CFR 1003.21(3)).b. It does not provide instructions, with respect to the use of the product pending the correction of the defect (21 CFR 1003.21(1). Specifically, the letter should include directions for the user to determine if the collimation defect is currently present in their system, which would initiate immediate correction, rather than at the next semi-annual service.2. The customer notification letter will be sent in the manner required by 21CFR 1003.21(b).3. The repairs outlined are to be completed by July 1, 2014. All systems under service contracts (which includes semi-annual service) would be visited between receipt of this letter and July 1, 2014. All repairs, including those to systems which are not under service contracts, are also expected to be completed by July 1, 2014.EOS may proceed with implementation of the CAP (subject to the conditions noted above. The customer notification letter may be sent out immediately upon addressing the above deficiencies. However, it is suggested that you provide a copy of the revised letter to FDA for approval prior to issuance.For further questions please call (678) 564-5400. |  | 22 | 22 EOS systems are installed in US | Nationwide Distribution | Terminated | January 27, 2015 | 446 | 22 | Computer
Z-0839-2014 | 66742 | No | ILO | Water circulating hot or cold pack. | Physical Medicine | Physical Medicine | 510(K) Exempt | DeRoyal (R) Aquarius Hot/Cold Therapy Combo Unit, w/ Foot/Ankle Blanket NS, REF T656NS, Rx only, Manufacturer DeRoyal Industries, Inc. 200 DeBusk Lane, Powell, TN 37849. |  |  | September 17, 2013 | 2013 | DeRoyal Industries Inc | The unit displays an error code of hose kinking when no hose kink is present; the unit fails to monitor the temperature of the water in the blanket as intended and controlling it through the automatic temperature control. This failure could result in temperature injuries to the treatment site; and there have been two reports of electric shock. | Device Design | DeRoyal sent a "URGENT VOLUNTARY RECALL DEROYAL HOT/COLD THERAPY UNIT" letter dated September 17, 2013. The letter identified the product, problem, and actions to be taken by the customers. |  | 25 | 25 units | Nationwide Distribution. | Terminated | June 05, 2017 | 1357 | 1343 | Computer
Z-0832-2014 | 67306 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Brilliance iCT &amp; Brilliance iCT SP Computed Tomography X-Ray Systems, Philips Healthcare, Cleveland, OH |  |  | December 09, 2013 | 2013 | Philips Medical Systems (Cleveland) Inc | Patient images exhibited ring artifacts. | Software design | The firm, Philips Healthcare, sent an "URGENT - Medical Device Correction" letter dated December 9, 2013 to their customers. The letter described the product, problem and actions to be taken. The customers were instructed to do the following to avoid the occurrence of these specific ring artifacts:1. Disable Automatic Scan Time for all Exam Cards2. Leave Automatic Scan Time on during planning and prior to clicking Goto initiate the scan, click the acquisition series of the scan and disableAutomatic Scan Time, then proceed with the scan.3. Disable Edit Before Final Recon (Reconstruction) for all Exam CardsIf a ring artifact is present, clinical judgment should be used by thecustomer/user to determine if the images acquired in the dataset contain enough information to make a diagnosis.Philips will be releasing a Field Change Order (FCO) to correct the issue with software version release 4.1.1.19261. In addition, a Philips Field Service Engineer will contact you to schedule the software installation at your site.If you need any further information or support concerning this issue, pleasecontact your local Philips representative or local Philips Healthcare office. ForNorth America and Canada, contact the Customer Care Solutions Center (1-800-722-9377, Option 5: Enter Site ID or follow the prompts). |  | 7 | 7 units | Worldwide Distribution: US (nationwide) including states of: FL, IN, NY, OR and PA; and internationally to: Australia and Israel. | Terminated | May 29, 2014 | 171 | 7 | Computer
Z-0224-2015 | 69537 | No | HQF | Ophthalmic laser. | Ophthalmic | Ophthalmic | 510(k) | Nidek MC-500 Multicolor Laser Photocoagulator using software version 2.20. |  |  | October 24, 2012 | 2012 | Nidek Inc | When the MC-500 is turned on and coagulation is performed without changing the laser color, the delivery unit occasionally fails to emit the laser, or the power output is lower than the set program. | Device Design | Nidek Field Service engineers visited each affected location to upgrade software. |  | 21 | 21affected devices | US Distribution in the states of: NJ, SC, IL, MI, NY, CA, TX, OH, IN, UT, and MN. | Terminated | November 17, 2014 | 754 | 21 | Computer
Z-0038-2015 | 69104 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | RECLAIM Distal Reamer Extension |  |  | September 03, 2014 | 2014 | DePuy Orthopaedics, Inc. | Specific lots are being recalled due to the potential for the tabs to break, leaving fragments in the patient. | Device Design | On February 25, 2013, DePuy Orthopaedics, Inc. issued a device correction for the specific lots of the RECLAIM Distal Reamer Extension impacted by this recall. The letter stated the reason for the recall and swap out. A subsequent letter outlined a formal swap-out process of the affected reamer extensions as part of the device correction strategy. Newly designed devices are now available. Questions about device swap-out information may be directed to Kim Earle, Recall Coordinator, 574-371-4917; and Surgeon questions may be directed to DePuy's Scientific Information Office at 1-888-554-2482. |  | 875 | 875 units | Worldwide Distribution -- USA, including the states of ME, MD, PA, MA, FL, LA, IL, IA, OH, MN, TX, MT, WA, CA, HI, NV, AZ, KY, WI, CO, GA, MI, NC, VA, NY, TN, and IN; and, the countries of Canada, Australia, Austria, Belgium, Czech Republic, United Kingdom, France, Germany, Ireland, Israel, New Zealand, Norway, Poland, Slovenia, South Africa, Switzerland, and Chile. | Terminated | November 09, 2015 | 432 | 875 | Computer
Z-0801-2014 | 67168 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Computed Tomography X-Ray System (Brilliance CT Big Bore Oncology &amp; Brilliance CT Big Bore Radiology) Distributed by Philips Healthcare System, Cleveland, OH. |  |  | December 04, 2013 | 2013 | Philips Medical Systems (Cleveland) Inc | Machine Name in TumorLOC Basic Mode is obscured. When TumorLOC is in Basic Mode and Set Vertical Laser equals Yes, the Machine Name label and field are covered by the controls above it. The Machine Name cannot be viewed or changed. | Software design | On 12/4/2013 the firm sent Urgent Medical Device Corrections letters to their customers. |  | 32 | 32 units | Worldwide Distribution-USA (nationwide) including the states of AK, MD, NH, NV, OH, and OR, and the countries of Australia, China, France, India, Indonesia, Italy, Korea, Mexico, Poland, Romania, Russia, Spain, Switzerland, Taiwan, Thailand &amp; United Arab Emirates. | Terminated | July 01, 2015 | 574 | 32 | Computer
Z-0101-2015 | 69242 | No | BSB | Blood and plasma warming device. | Hematology | Hematology | 510(k) | Ultra Battery 1 Model: TA-UB1 NSN: 6515-01-592-5803 |  |  | July 25, 2014 | 2014 | Estill Medical Technologies, Inc | Estill has discovered that a fully discharged Ultra Battery 1 (TA-UB1) will reach 96.9% of its capacity during a 60 minute charging period. If you have a fully discharged Ultra Battery 1 (TA-UB1), in order to achieve a 100% of charge capacity you should charge it for a period of 90 minutes. In addition, the charging voltage range of 10-36V specified on the Ultra Battery 1 (TA-UB1) label and in t | Device Design | Thermal Angel sent a Notice of Corrections letter dated July 25, 2014, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to communicate the contents of the notice to all employees, consignees, consultants, or other parties affected by this correction. The firm also updated the labeling and instructions for use. For additional information cusotmers were instructed to contact Medical Technologies at (877)354-0286.The recalling firm issued a Notice of Corrections to its distributors and customers.For questions regarding this recall call 214-561-6001. |  | 6886 | 6,886 | Worldwide Distribution - USA (nationwide) and Internationally to Serbia. | Terminated | August 21, 2015 | 392 | 6886 | Computer
Z-0828-2014 | 67278 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Accuray CyberKnife Robotic Radiosurgery System; Accuray Incorporated Sunnyvale, CA. |  |  | January 13, 2014 | 2014 | Accuray Incorporated | Potential Safety issue with Synchrony Boom Arm Mounting Assembly - one complaint of mounting assembly detaching. | Device Design | Urgent Medical Device Correction notification letter were sent out beginning January 14, 2014 by trackable method. Letters describe potential problem and provide inspection instructions to be performed prior to use. |  | 44998 | 302 devices; 149 in US, 153 out of US. | Worldwide Distribution. - US (Nationwide) and Mexico; and Internationally to: Australia, Belgium, Canada, China, Colombia, Czech Republic, England, Finland, France, Germany, Greece, India, Ireland, Italy, Japan, Korea, Malaysia, Luxembourg, Netherlands, Poland, Russia, Saudia Arabia, Spain, Switzerland, Taiwan, Thailand, Turkey, Ukraine, United Kingdom, Venezuela, and Vietnam | Terminated | November 25, 2014 | 316 | 44998 | Computer
Z-0844-2014 | 67188 | No | KPR | Stationary x-ray system. | Radiology | Radiology | 510(k) | KODAK DirectView DR 7500 Dual Detector System, MODEL DR 7500, Manufactured by Carestream Health, Inc. Made in U.S.A. |  |  | December 26, 2013 | 2013 | Carestream Health, Inc. | Carestream Health, Inc. has recalled DR 7500 Dual Detector System utilizing Version 5.X Software due to a possible patient image display error. | Software Manufacturing/Software Deployment | On December 26, 2013 Carestream Health, Inc. distributed Urgent Medical Device Recall notification letters dated December 19, 2013 to their customers via Fed Ex. Carestream Health, Inc. recommends that once the daily detector calibration process is initiated on the DR 7500 System running Version 5.X software it should not be interrupted. The calibration process should be completed only when patients are not being imaged. A Carestream Health Service representative will contact all customers to install a software modiciation to Version 5.X software which will not allow the cancellation of detector calibration. Customers will questions can call 1-800-328-2910 7 days per week on a 24 hour basis. |  | 164 | 164 units Total (42 domestically &amp; 122 internationally) | worldwide Distribution - US (Nationwide) including states of: IN, PA, MD, MI, NY, GA, MN, OK, VA, TX, WI, CA, IL, TN, WA, NC, SC, MA, NJ, and KY; and Internationally to: Canada, New Zealand, South Korea, Australia, Hong Kong, Saudia Arabia, Guinea, Kuwait, Israel, Oman, South Africa, China, UK, Sweden, Poland, Denmark, Germany, Finland, Czech Republic, Italy, France, Portugal and Spain. | Open, Classified |  |  | 164 | Computer
Z-0011-2015 | 69092 | No | INM | Mechanical chair. | Physical Medicine | Physical Medicine | 510(K) Exempt | Nemschooff, Model Numbers - Serenity II:790-12, Serenity III: 791-12, Serenity III Exam: 791-12E, and Serenity IIS: 792-12. Designed to provide patient comfort during patient treatment sessions. The chair may include a manual control reclining mechanism. Customer configurable models are also available. The Serenity III Exam 791-12E was discontinued 11/1/2010. |  |  | January 18, 2011 | 2011 | Nemschoff Chairs, Inc. | Nemschoff Chairs, Inc. is recalling various model of the Serenity Treatment Chair because of potential elevated temperatures in the seat and back lumbar areas of the chair. | Device Design | The corrections have been performed via heat service kits being installed "on demand" to address heating element malfunctions starting on 1/18/2011. Consignee letter to all customers is pending HHE assessment. |  | 1868 | 1868 | Worldwide Distribution: USA (nationwide) including DC and the states of AK, AL, AR, AZ, CA, CO, CT, FL, GA, IA, IL ,IN, KS, KY, MA, MD, MI, MN, MO, MN, MT, NC, ND,NE, BH, NJ NM, NV, NY, OH, OK, OR, PA, PE, SC, SD, TN, TX, UT, VA, WA, WI, WV, WY, and the country of CANADA. | Terminated | July 24, 2015 | 1648 | 1868 | Computer
Z-0110-2015 | 69467 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Mortara Surveyor S12 and Surveyor S19 Patient Monitors. |  |  | October 07, 2014 | 2014 | Mortara Instrument, Inc | Martara Instrument, Inc. has recently became aware of a potential safety hazard involving our Surveyor rand S12 S19 Patient Monitors. A hazardous situation may be created when mounting these monitors using the 75mm VESA mounting hole pattern provided in the battery cover of these monitors. If the mounting screws provided by Mortara are not used and if the mounting screws used are longer than the p | Device Design | Mortar Instruments, Inc. sent an "Urgent Medical Device Correction" letter dated October 3, 2014. The letter was addressed to Health Care Administrator / Risk Manger and Director of Biomedical / Clinical Engineering.The letter described the problem and the product involved in the recall. Customers were advised to contact Mortara Technical Support at 1-888-667-8272 or TechSupport @mortara.com to schedule a service representative to update their Surveyor S12 and Surveyor S19 units or for additional questions they can contact their local Service Representative. The update will include replacement of the battery cover and replacement of the battery pack if mounting screw damage is evident. International consignees received similar dated letter except it provided instructions detailing the correction steps and requested customers to complete and return the Response Acknowledgement and Receipt Form. For questions they can contact Technical Support at +1.414.354.1600 or TechSupport@mortara.com, or contact your local Service Representative. |  | 624 | 624 | Worldwide Distribution: US (Nationwide) in the states of: AZ, AR, CA, DC, FL, ID, IL IN, KY, LA, ME, MD, MA, MI, MN, MO, MT, NV, NJ, NY, NC, OK PA, SC, TN, TX, VA, and WI; and internationally to: AUSTRALIA, BANGLADESH, CANADA, CHILE, CROATIA, EGYPT, GERMANY, ITALY, JORDAN, LEBANON, MALAYSIA, MOROCCO, OMAN, PHILIPPINES, POLAND, ROMANIA, SAUDI ARABIA, THAILAND, THE NETHERLANDS, UNITED KINGDOM. | Terminated | February 13, 2015 | 129 | 624 | Computer
Z-0800-2014 | 67030 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Siemens ADVIA 1200 and ADVIA 1650 Chemistry Systems. |  |  | December 17, 2013 | 2013 | Siemens Healthcare Diagnostics, Inc. | An overflow flag //// is displayed in the result column for a test when the algorithm for the test cannot calculate a result or when the concentration cannot be determined because it is outside the absorbance limits, and it is possible for the user to misinterpret or edit the result. | Software design | The distribution of the Urgent Medical Device Correction letter scheduled to begin on December 17, 2013.Âżx Field service personnel will be notified of the Urgent Medical Device Correction Letter via the Siemens Knowledge Base, as well as information regarding placing labels on the customer system to remind the customer that ///// is not a reportable result.Âżx The Customers were instructed to contact Siemens Healthcare Diagnostics Technical SolutionCenter if they have technical questions regarding this issue.Âżx Customers are requested to complete and return the Effectiveness Check questionnaire attached to the Urgent Medical Device Correction. |  | 1563 | 1563 | Worldwide Distribution. USA nationwide and the country of Canada. | Terminated | August 26, 2014 | 252 | 1563 | Computer
Z-0802-2014 | 66591 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Civco Belly Board, Reusable non-sterile carbon fiber belly board with cushion, REF 125012, packaged 1 per box. |  |  | October 25, 2013 | 2013 | Civco Medical Instruments Inc | Scales adhered to the product may be misaligned from the zero reference such that one side will not match the other. | Device Design | Civco Medical Instruments Inc sent a FIELD SAFETY NOTIFICATION letter dated October 28, 2013 via UPS to their customers. The notification identified the affected device, problem, and field safety corrective action. The letter describes how customers can check to see whether the scales on their device are out of alignment. The letter states CIVCO will rework the system to correct the difference. For questions contact sales representative at CIVCO Medical Solutions, Radiation Oncology at 800-842-8688 or 712-737-8688. |  | 138 | 138 belly boards | Worldwide Distribution - USA Nationwide in the states of AL, AR, AZ, CA, CO, CT, FL, GA, IL, KS, MA, MD, MI, MN, MT, NC, NJ, NY, OH, OK, PA, SC, TN, TX, VA, WA, WI, and WV. Foreign distribution was made to Brazil, Canada, Chile, Czech Republic, France, Germany, Iraq, Israel, Italy, Japan, Malaysia, Saudi Arabia, Slovenia, South Korea, Spain, Turkey, and United Arab Emirates. | Open, Classified |  |  | 140 | Computer
Z-0831-2014 | 66197 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Brilliance iCT and Brilliance iCT SP |  |  | July 16, 2013 | 2013 | Philips Medical Systems (Cleveland) Inc | An artifact that resembles thrombus may appear on the image. | Device Design | Philips sent an Field Safety Notice dated July 15, 2013, to all US customers via certified mail. Market representatives will distribute the letter internationally. The letter identified the product the problem and the action needed to be taken by the customers. Additionally, a Field Service Engineer will contact each site to schedule a time to implement the Field Change Order.If you need any further information or support concerning this issue, please contact your local Philips representative or local Philips Healthcare office.For North America and Canada, contact the Customer Care Solutions Center (1"800-722"9377,Option 5: Enter Site 10 or follow the prompts).This notice has been reported to the appropriate Regulatory AgencyPhilips apologizes for any inconveniences caused by this problem |  | 506 | 506 | Worldwide Distribution - USA (nationwide) including Puerto Rico and the states AL, AR, AZ, CA, DC, DE, FL, GA, IA, IL, IN, KY, MA, MD, MN, MO, MS, NC, NV, NY, OH, OK, OR, PA, SC, TX, VT, WA, and WV., and the countries of Argentina, Australia, Austria, Azerbaijan, Bangladesh, Belgium, Brazil, Chile, China, Czech Republic, Denmark, Egypt, France, Germany, Greece, Hungary, Iceland, India, Indonesia, Iraq, Ireland, Israel, Italy, Japan, Jordan, Lebanon, Netherlands, New Zealand, Norway, Oman, Panama, Philippines, Poland, Qatar, Romania, Russia, Saudi Arabia, Singapore, South Korea, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, United Arab Emirates, United Kingdom, and Uzbekistan. | Terminated | September 19, 2014 | 430 | 506 | Computer
Z-0127-2015 | 69334 | No | LCP | Glycosylated hemoglobin assay. | Hematology | Hematology | 510(k) | Dimension VistaÂż HbA1c |  |  | August 14, 2014 | 2014 | Siemens Healthcare Diagnostics, Inc. | In March, Siemens Healthcare Diagnostics issued an Urgent Medical Device Correction for decreased recovery of patient samples with the Dimension Vista&reg; CSA &amp; CSAE Flex&reg; reagent cartridges when whole blood samples are run from Small Sample Containers (SSC). Siemens has confirmed that when whole blood samples are run from SSC, HbA1c values may exhibit a percent bias ranging from - 8.2% t | Device Design | An Urgent Medical Device Correction, dated August 2014, was provided to all Dimension VistaÂż HbA1c customers who received the affected lots to notify them of the issue, the potential risk to health, and actions to be taken by customers. |  | 20043 | 20043 | Worldwide Distribution: US (nationwide) and country of: Canada. | Terminated | August 10, 2015 | 361 | 20043 | Computer
Z-0235-2015 | 69601 | No | KPQ | Radiation therapy simulation system. | Radiology | Radiology | 510(k) | Siemens SimView NT and SimView NT with IDI simulation system components: |  |  | October 10, 2014 | 2014 | Siemens Medical Solutions USA, Inc | Siemens became aware of a very rare occurrence where an intended movement of the SimView NT simulation system components could not be stopped with usual hand control commands or by use of the collision detection functionality.It is possible that simulation system components could collide with patient and/or system equipment. | Device Design | Customers were initially notified about the issue by letter on January 7, 2014 (this action was reported as Z-0786-2014 which has been terminated). Siemens promised to notify them when a fix became available. Customer letter, dated November 7, 2014, was sent to notify end users that a fix has become available and Siemens would contact them to schedule an appointment for the correction. |  | 6 | 6 | Worldwide Distribution - US Nationwide in the states of PA, KY, LA, MA, MI, and CA; and countries of: Argentina, Belarus, Bulgaria, Costa Rico, P.R. China, Curacao, Egypt, Repulic Korea, Germany, Greece, Hungary, Italy, Iraq, Jordan, Malaysia, Palestanian Ter, Poland, Thailand, Turkey, and Vietnam. | Terminated | June 29, 2017 | 993 | 6 | Computer
Z-0020-2015 | 68846 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Proteus 235 |  |  | March 27, 2014 | 2014 | Ion Beam Applications S.A. | While the Proteus 235 is in the pencil beam scanning (PBS) treatment mode, on rare occasions, Low Level Radio Frequency (LLRF) pulsed dee voltage (i.e., voltage applied between two electrodes in the cyclotron to accelerate protons) may generate a hot spot on a pencil beam position within a tumor, generating a dose that might exceed the prescribed dose on a single voxel. | Device Design | A User Notice was sent to direct accounts in/about March/April 2014 to notify them of the issue and provide instructions to avoid the identified potential risk to patients. |  | 8 | 8 | Worldwide Distribution-USA (nationwide) including the states of PA, NJ, IL, WA, and TN, and the countries of Germany, Italy, and Czech Republic. | Terminated | December 10, 2014 | 258 | 8 | Computer
Z-0216-2015 | 69487 | No | IOR | Mechanical wheelchair. | Physical Medicine | Physical Medicine | 510(k) | Transport Chair and Rollator in 1, Model HE21630011 Burgundy and HE21630011 Blue. |  |  | October 10, 2014 | 2014 | Silver Star Brands | Silver Star Brands is recalling the Transport Chair and Rollator in One because they have received several reports of the wheel assembly failing resulting in damaged wheels.. | Device Design | Consignees were sent on 10/10/2014 an Easy Comforts "Urgent: Rollator Recall" letter dated October 10, 2014. The letter described the problem and the product involved in the recall. Advised consignees to stop using the product immediately. The consignees can contact the Product specialist at Easy Comforts at 1-888-327-2390 Monday though Friday from 7:30 am to 3:30 pm central time. |  | 129 | 129 | Nationwide Distribution including the states of AL, AZ, AR, CA, CO, CT, FL, GA, HI, IL, IN, IA, KS, LA, MD, MI, MO, MT, NE, NV,NH, NM, NY, NC, OH,OK, and PA. | Open, Classified |  |  | 129 | Computer
Z-2708-2014 | 69191 | No | KPA | Automated slide stainer. | Pathology | Pathology | 510(K) Exempt | Dako Autostainer Link 48 with software version Dako Link 4.0.3, Model No. AS480 |  |  | August 29, 2014 | 2014 | Dako North America Inc. | Dako is recalling the Dako Link 4.0.3 software on the Dako Autostainer Link 48 because Dako has determined that an upgrade to this version of the software may cause premature rejection of reagent bottles or inadequate aspiration of reagent if the instrument has also had a preventative maintenance or waste pump replacement, and the baseplate is not completely aligned. | Software change control | Dako North America sent an Urgent Medical Device Correction letter to all customers who were upgraded to the Dako Link 4.0.3 software that is used on the Dako Autostainer Link 48 and who also had preventative maintenance or a waste pump replacement. The letter informs the customers of the problems identified and the actions to be taken. Customers are instructed that Dako will contact customers to set up an appointment for a service visit in order to ensure that the instruments are running safely or reliably. Customers with questions are instructed to contact Dako. |  | 191 | 191 units total (57 units in US) | Worldwide Distribution - US Nationwide in the states of AT, AU, BE, CA, CN, DE, DK, ES, Fl, FR, GB, IE, IT, NL, NO, SE, US, DZ, GR, IS, OM, SG, UY. | Terminated | March 03, 2015 | 186 | 191 | Computer
Z-0857-2014 | 67199 | Yes | DWA | Cardiopulmonary bypass pump speed control. | Cardiovascular | Cardiovascular | 510(k) | CP5 centrifugal pump system |  |  | December 04, 2013 | 2013 | Sorin Group Deutschland GmbH | Sorin is initiating a field correction on S5/C5 Heart-lung machine and CP5 centrifugal pump system due to reports of pump speed control knob failures resulting in no response. | Device Design | Sarin Group US distributed a Field Safety Notice dated December 5, 2013 to all affected customers and distributors. The letter informed users of the potential for shaft encoder failures and notified them that they would be contacted by the Sorin Service Group to arrange an appointment to repair the units onsite. Customers and distributors were also provided with a Response form to confirm they had received, read and understood the Field Notice. They were instructed to return the completed form to assist in monitoring the effectiveness of the communication. For questions call 303-467-6527 or 303-467-6503. |  | 75 | 75 | Worldwide Distribution - US (Nationwide) and PR; and Internationally to:BE, ES, FI, GB, NL, SE, AR, AT, AU, CA, CH, CN, CR, EG, HK, ID, JO, JP, KR, KW, LB, LY, MM, PL, RO, RS, RU, SG. TH, TN, TR, TW, UA, VN, and ZA. | Terminated | December 28, 2015 | 754 | 398 | Computer
Z-0611-2014 | 67108 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Puritan Bennett 840 ventilator systems software part number 4-070212-85, revision AB-AG. |  |  | December 16, 2013 | 2013 | Covidien | Covidien is conducting a voluntary field corrective action on Puritan Bennett 840 ventilator systems with certain software revisions in response to customer reports of ventilator diagnostic code XB0069 in which the device stops mechanical ventilation due to a software error. | Software design | Covidien sent an Urgent Medical Device Voluntary Field Correction letter to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to continue using their Puritan Bennett 840 ventilators until it is feasible to install the software update discussed in the letter. Covidien has developed a software solution to address this issue. Upon availability, there will be several ways to install the software update. To initiate the process for updating your software and to select the method that best suits your facility, please visit our software update management portal at www.PB840technicalupdate.com and follow the instructions to register your ventilators. If you need further assistance contact our Technical Support Department @ 1-800-255-6774, option 4, then option 1.For questions regarding this recall call 800-255-6774. |  | 56295 | 16,923 US, 39,372 OUS | Worldwide Distribution - USA (nationwide) and Internationally to Austria, Estonia, Belgium, Spain, Italy, Poland, Bulgaria, Finland, Portugal, Switzerland, France, Lithuania, Romania, Cyprus, Great Britain, Czech Republic, Greece, Latvia, Slovenia, Germany, Hungary, Slovak Republic, Ireland, Netherlands, Turkey, Croatia, Andorra, Albania, Algeria, Bahrain, Bosnia and Herzegovina, Egypt, Georgia, Iran, Iraq, Israel, Jordan, Kenya, Kuwait, Lebanon, Libya, Macedonia, Montenegro, Morocco, Russia, Saudi Arabia, Serbia, South Africa, Syria, Sudan, Tunisia, UAE, Uganda, Puerto Rico, Mexico, Columbia, Chile, Brazil, Argentina, Uruguay, Panama, Costa Rica, Peru, China, Korea, Japan, Taiwan, Hong Kong, Malaysia, Singapore, Sri Lanka, Australia, New Zealand. | Terminated | November 29, 2016 | 1079 | 56295 | Computer
Z-0060-2015 | 69207 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | GE Healthcare Carescape Patient Data Module |  |  | September 19, 2014 | 2014 | GE Healthcare, LLC | GE Healthcare has recently become aware of a potential safety issue to the ECG calculations following a disconnect/reconnect cycle with the Patient Data Module, when used with the Carescape Bx50 monitors.If the PDM is quickly disconnected and reconnected (a less than 10 second cycle) from a Carescape Bx50 monitor following a Patient Discharge or during active monitoring, the ECG waveform an | Software design | Consignees were sent on 9/19/2014 a GE Healthcare "Urgent Medical Device Correction" letter Ref# 336107 dated September 18, 2014.The letter was addressed to Director of Biomedical Engineering, Chief of Nursing and Healthcare Administrator / Risk Manager. The letter described the Safety Issue, Safety Instructions , Affected Product Details, Product Correction and Contact Information. For questions they can contact Technical Support at 1-800-558-70444 or your local Service Representative. |  | 1691517372 | 68,452 (39,505 units US, 24,711 units OUS, 4,236 units unidentified/unknown) | Worldwide Distribution: US (nationwide) including DC, PR, Guantanamo except DE, HI, ND, SD &amp; WY; and countries of: SOUTH AFRICA, KOREA, UNITED KINGDOM, UNITED ARAB EMIRATES, TURKEY, TUNISIA, TRINIDAD TOBAGO, THAILAND, TANZANIA, TAIWAN, SWITZERLAND, SWEDEN, SPAIN, SLOVAKIA, SINGAPORE, SAUDI ARABIA, RUSSIA, QATAR, BRASIL, PORTUGAL, POLAND, PHILIPPINES, PERU, PANAMA, PARAGUAY, OMAN, NEW ZEALAND, NORWAY, NETHERLAND, NEPAL, MYANMAR, MEXICO, MOROCCO, MALDOVA, MAURITIUS, MALAYSIA, LITHUANIA, LEBANON, BAHRAIN, KUWAIT, REPUBLIC OF KOREA, JAPAN, JORDAN, JAMAICA, ITALY, ISRAEL, IRELAND, INDONESIA, INDIA, ICELAND, HUNGARY, HONG KONG, GUATEMALA, GREECE, GERMANY, FRANCE, FINLAND, ESTONIA, EGYPT,ECUADOR, DENMARK, CECH REPUBLIC, COSTA RICA, COLOMBIA,, CHINA, CANADA, BULGARIA, BELGIUM BANGLADESH, AUSTRIA, AUSTRALIA, and ARGENTINA. | Open, Classified |  |  | 1691517372 | Computer
Z-0823-2014 | 66694 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Siemens CentraLink" Data Management System V14.0.4, V14.0.5 or V14.0.8 |  |  | October 30, 2013 | 2013 | Siemens Healthcare Diagnostics, Inc. | Under extremely rare circumstances, a patient result that had been previously disabled may be released without proper review. (Disabled Test in CentraLink System or Aptio System May Become Enabled in CentraLink System.) | Software design | The firm, Siemens, sent an "Urgent Medical Device Correction" letter dated October 23, 2013 to its customers. The firm also sent a follow-up letter dated October 30, 2013. The letter describes the product, problem and actions to be taken. The customers were instructed Until a service visit is scheduled, do the following: Instruct the laboratory staff to recognize this sequence of events and Ensure that whenever the operator disables a method on an Aptio system or a CentraLinksystem and replenishes reagent for that method, the operator then confirms that the method remains disabled in the CentraLink system or disables it again. The operator must allow up to five minutes to verify that the test remains disabled. In addition, the customers were instructed tocomplete and return the FIELD CORRECTION EFFECTIVENESS CHECK response form via fax to Siemens Healthcare Diagnostics Customer Care Center at 800-441-9250, and to retain this letter with your laboratory records, and forward this letter to those who may have received this product. After a service visit is performed and an interim solution is implemented, the laboratory staff will be trained on this update to understand the minor but necessary changes in the workflow.If you have any questions, please contact your Siemens Customer Care Center at 800-441-9250 or your local Siemens representative. |  | 40 | 40 | US Distribution in the states of: MN, NJ, and NV. | Terminated | December 15, 2014 | 411 | 40 | Computer
Z-0485-2015 | 69728 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Esprit V1000 Ventilator and Esprit/V200 Conversion Ventilator |  |  | October 29, 2014 | 2014 | Respironics California Inc | If the power supply fan mounting screws are installed incorrectly, the ends of the screws may touch the power supply printed circuit board assembly (PCBA) and short to ground. This may prevent the ventilator to switch to battery power if AC power is disrupted. | Device Design | Philips Healthcare sent an Urgent - Field Safety Notice to all affected consignees to inform them that Respironics is recalling the Esprit V1000 and Esprit/V200 Conversion ventilators because if the power supply fan mounting screws are installed incorrectly, the ends of the screws may touch the power supply printed circuit board assembly (PCBA) and short to ground. This may prevent the ventilator to switch to battery power if AC power is disrupted. A high urgency audible alarm will sound alerting medical personnel of the loss of power. In the absence of AC power, failure of battery power will cause loss of ventilator support, which may result in hypercarbia or hypoxemia. The potential issue affects ventilators manufactured prior to 12/31/04 due to design of the fan/shroud assembly. Esprit V1000 and Esprit/V200 Conversion Ventilators manufactured after that affected date are not affected. Customers with concerns were instructed to contact a Philips Respironics representative at (800) 722-9377. For questions regarding this recall call 760-918-7300. |  | 3943 | 3,943 units total (3,071 units in US) | Worldwide Distribution - USA (nationwide) and Internationally to Argentina, Australia, Brazil, Bulgaria, Canada, Chile, China, Colombia, Ecuador, Egypt, Greece, Hong Kong, India, Italy, Korea, Lebanon, Malaysia, Mexico, Pakistan, Panama, Peru, Philippines, Saudi Arabia, Singapore, South Africa, Spain, Taiwan, Thailand, Turkey, United Arab Emirates, United Kingdom, Venezuela, Vietnam, and Yemen. | Open, Classified |  |  | 3943 | Computer
Z-0782-2014 | 67123 | No | JEQ | Powered toothbrush. | Dental | Dental | 510(K) Exempt | UP &amp; UP kids' power toothbrush |  |  | December 18, 2013 | 2013 | Ranir Corporation | Ranir, LLC is recalling all lots of UP &amp; UP brand Kid's Power Toothbrush and TopCare brand Clifford Kids Power Toothbrush because the tufted disk may become detached from the power brush handle. If thetuft retention disk detaches from the handle, it is a potential choking hazard for children three years old and younger. | Device Design | Ranir, LLC phoned their two direct accounts notifying them of this recall on December 12, 2013. On December 18, 2013, Ranir, LLC sent an URGENT: MEDICAL DEVICE RECALL notification letter was sent to their direct accounts. The notification stated the affected product, issue, and instructions for responding to the notification, including instructions for return of the affected product. Questions are directed to Paula Bojsen (616)698-8880. |  | 601616 | 601,616 total | Distributed in MN and IL. | Terminated | August 20, 2015 | 610 | 601616 | Computer
Z-0483-2015 | 69627 | No | BZD | Noncontinuous ventilator (IPPB). | Anesthesiology | Anesthesiology | 510(k) | Cannulaide, Model 101, 102, 103. Respirator. |  |  | May 13, 2008 | 2008 | Beevers Manufacturing &amp; Supply, Inc. | Cannulaide recalled May 2008 due to improper seal of the pouch. The improper seal of the Cannulaide's package could potentially affect the device shelf life. | Equipment maintenance | Beevers Manufacturing &amp; Supply, Inc. (BMS) did contact the customers via letter and phone call in May 2008. Customers that didn't reply were further contacted via email on June 2008. Salter Labs provided BMS' recall report to FDA in October 2014.Salter Labs was unable to locate a recall letter issued by BMS to the affected customers. For further questions please call (503) 472-9055 |  | 16575 | 16,575 units total (15,625 units in US and 950 units outside US) | Worldwide Distribution - US Distribution including the states of LA, GA, OR, MA, NJ, FL and CA., and the countries of Canada, France, Japan, Netherlands and United Kingdom. | Terminated | November 28, 2014 | 2390 | 16575 | Computer
Z-2693-2014 | 68594 | No | FSA | Non-AC-powered patient lift. | General Hospital | General Hospital | 510(K) Exempt | Minstrel (with scale) |  |  | July 28, 2014 | 2014 | ARJOHUNTLEIGH POLSKA Sp. z.o.o. | ArjoHuntleigh received three reports where the bolt under the scale unscrewed from the scale. There is a remote risk concerning the bolt unscrewing from the scale. When the scale and bolt parts are not correctly serviced, or when they receive a sharp blow during or outside of use, it is possible the bolt holding the spreader bar to the scale could unscrew during subsequent use. As a result, the sp | Device Design | ARJOHUNTLEIGH sent an URGENT - FIELD SAFETY NOTICE letter dated July 28, 2014 to all customers. The letter identified the affected product, problem and actions to be taken. The letters provided customers with two options with how to address this recall: 1) because the devices are close to their 10-year lifetime, customers can return the devices for a trade-in discount that will be applied to the purchase of a new ArjoHuntleigh patient transfer lift; or 2) an ArjoHuntleigh Service Technician will install a new scale with added locking features to the existing unit free of charge. Customers were further instructed to: 1) ensure that all caregivers and users of the devices are given a copy of the letter; 2) place a copy of this letter with the device's instruction for use; and, 3) to complete and return the enclosed Customer Service Response form. For questions or assistance completing the Customer Response Form, please contact ArjoHuntleigh at 1-800-323-1245, ext. 57985 during office hours; Monday to Friday, 8:00am to 5:00pm (CST) or via email at AHNA-RA@arjohuntleigh.com. |  | 18 | 18 Lifts | Worldwide Distribution - US Nationwide in the state of MA and the countries of Australia, Germany, and Italy. | Open, Classified |  |  | 18 | Computer
Z-0063-2015 | 69270 | No | NAY | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | EndoWrist One Vessel Sealer used in conjunction with the da Vinci Si Surgical System IS3000. |  |  | September 17, 2014 | 2014 | Intuitive Surgical, Inc. | Intuitive Surgical is initiating a voluntary correction relating to the labeling and software associated with the EndoWrist. One Vessel Sealer for the da Vinci Si (IS3000) Surgical System. | Device Design | Urgent medical device correction letter was sent to all consignees on 9/22/14 by trackable mail system.Letters identified the issue and provides information for users should a error code be displayed. |  | 1483 | 1,483 devices | Worldwide Distribution-USA (nationwide) and the countries of Australia, Austria, Belgium, China, Chile, Colombia, Czech Republic, Denmark, Finland, France, Germany, Greece, Italy, Japan, Netherlands, Norway, South Korea, Spain, Sweden, Switzerland, Turkey, and United Kingdom. | Terminated | June 01, 2015 | 257 | 1483 | Computer
Z-0146-2015 | 69323 | No | LZG | Infusion pump. | General Hospital | General Hospital | 510(k) | ACCU-CHEK Spirit Combo insulin infusion pump sold as part of Accu-Chek Combo System |  |  | September 25, 2014 | 2014 | Roche Diagnostics Operations, Inc. | The insulin pump may lose time and date settings during a power interruption (such as a battery change) due to a faulty capacitor. If the capacitor fails, the time and date will return to default setting. Though the pump prompts customers to confirm the time and date, if the user does not recognize the time and date have returned to default, a shift of their basal rate time block could occur. | Equipment maintenance | Roche sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated September 25, 2014 via UPS Ground (receipt required) to the consignees. The letter included a description of the reason for recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. No product returns were expected. For questions, contact the ACCU-CHEK Customer Care Service Center at 1-800-688-4578. |  | 8116 | 8116 | US Nationwide Distribution in the states of CA, CT, IA, NV, OH, IN, WI, GA, MO, NY, VA, LA, ME, TX, NC, ND, SD, NJ, CO, AL, SC, IL, KY, FL, OR, MN, MD, MS, AR, WV, WA, TN, CT, ID, KS, MI, PA, UT, MA, RI, NM, OK, DE, AZ, NE, MT, NV, WY, AK, and NH. | Open, Classified |  |  | 8116 | Computer
Z-0062-2015 | 69294 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | MaxLock Extreme Mod-Foot Concave Reamer; Individually packaged in a 3 x 4 bag and heat sealed with label on the outside. |  |  | September 19, 2014 | 2014 | Orthohelix Surgical Designs Inc | Incorrect sizing was being laser etched on the MaxLock Extreme Mod-Foot Concave Reamer | Device Design | An Urgent - Medical Device Recall customer notification, dated September 19, 2014, was sent out via email and FEDEx to all distributors with instructions on returning the product. The letter identified the affected product and the reason for the recall. Customers are to immediately cease distribution of the affected product, and complete and return the questionnaire. The letter also states that the recalled product will be replaced. Questions are to be directed to the Customer Service Department 1-888-494-7950. |  | 38 | 38 | USA Distribution -- in the states of WI, NM, TX, WA, MO, OH, UT, TN, MI, VA, NC, VA, and PA. | Terminated | October 31, 2016 | 773 | 38 | Computer
Z-0099-2015 | 66541 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | CRN Crono Syringe CRN20 with 19G x 1 "Steel Needle" |  |  | October 18, 2013 | 2013 | Intra Pump Infusion Systems | Product has the potential to have a clear debris on the tip of the needle. | Device Design | Cane S.p.A. Medical Technology sent an Urgent: Correction and Removal Action Required letter dated October 21, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Actions to be taken:a. Stop using the 19G x 11/2" 1.10 x 40mm steelneedle included in the CRONO Syringe blister pack. B. Safely discard the 19G x 11/2" 1.10 x 40mm steel needle into a Sharps container.c. Use a TRANSFER DEVICE (e.g.,Needleless adapter with female-to-female luer connector) or alternative device to transfer medications into the CRONO syringe.d.For intravenous infusions,we are recommending that a filter also be used during infusions.e. Complete the required 'Acknowledgment and Receipt Form'and fax back to our Distributor;Intra Pump Infusion Systems at 630-845-2768.Contact Information for our US Distributor: INTRA PUMP INFUSION SYSTEMS " Monday through Friday 8:00am to 5:00pm CST,TollFree:1-866-211-7867 " Direct:630-845-7867 " Fax:630-845-2768 www.lntrapump.com |  | 25440 | 25,440 pieces | Worldwide Distribution: US (nationwide) in states of CA, PA, MA, IL, OH and NY; and Internationally to Canada. | Terminated | February 01, 2016 | 836 | 25440 | Computer
Z-0142-2015 | 64774 | walker, mechanical | ITJ | Mechanical walker. | Physical Medicine | Physical Medicine | 510(K) Exempt | DMI Duro-Med Industries Rollator Transport Chair; Wheel Size: 7 1/4"; Overall Width: 23 1/2"; Seat Height: 21 1/4"; Seat Size: 14" x 12 3/4"; Handle Height: 31" - 36"; Width Between Handles: 18"; Weight: 19 Lbs; |  |  | June 05, 2014 | 2014 | Briggs Medical Service Company d.b.a. Mabis | There have been 5 customer reports of a 501-1018 series DMI Rollator Transport Chair collapsing during use. | Device Design | The firm, Briggs Healthcare, sent "URGENT PRODUCT RECALL" letters dated July 28, 2014 to all distributors, retailers and consumers. The letters included instructions for distributors and dealers to: 1) examine their inventory; 2) quarantine any affected products; 3) cease distribution of the affected products; and, 4) complete and return the attached Customer Recall Acknowledgement Form. Distributors and dealers who further distributed the products were also instructed to notify their customers of this recall by forwarding a copy of the enclosed consumer notification letter and its associated documents. Distributors and dealers with questions about this recall can contact Briggs at 800-342-4426 (Monday - Friday, 8:00 AM - 6:00 PM EDT) or by e-mail at RecallCoord@briggscorp.com. The consumer-directed letter dated July 28, 2014 included instructions for consumers to examine their rollator transport chair to determine whether or not it is subject to the recall. If the consumer has the affected product, they should discontinue using it immediately and contact Briggs at 800--342-4426 (Monday - Friday, 8:00 AM - 6:00 PM EDT) or by e-mail at RecallCoord@briggscorp.com to arrange for the return of the product. Consumers will be given the option of receiving a cash voucher or a replacement product in exchange for the recalled product. |  | 1869 | 1,869 units | Worldwide Distribution: US (nationwide) including states of:US: Nationwide (AL, AR, AZ, CA, CO, CT, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, WY) + District of Columbia + Puerto Rico; and country of : Canada.); FOREIGN: Canada | Open, Classified |  |  | 1869 | Computer
Z-2709-2014 | 69123 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Brilliance CT System 6/10/16/16P/40 configuration. No packaging, sold as single unit. |  |  | December 27, 2005 | 2005 | Philips Medical Systems (Cleveland) Inc | The system has a gap that can appear between the upper and lower tilt stand covers. The fingers may be squeezed if a person was to place them in this region while the gantry was tilting from an angle to zero tilt angle. | Device Design | Customers will be contacted by a field service technician with the field correction order. The firm will install the new labeling to the device. |  | 364 | 364 | Worldwide Distribution -- AZ, CO, GA, IA, IL, IN, KS, KY, MD, MI, MN, MS, NC, NY, OH, OK, TX, VT, and WI; and, countries of Singapore, New Zealand, India, Australia, Korea, Japan, Korea, China, India, Taiwan, Australia, Turkey, France, South Africa, Spain, Italy, Austria, Belgium, Switzerland, Israel, Netherlands, England, and Brazil. | Terminated | November 06, 2015 | 3601 | 364 | Computer
Z-0112-2015 | 69320 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Covidien Puritan Bennett 980 Ventilator, Rx ONLY. |  |  | October 01, 2014 | 2014 | Nellcor Puritan Bennett Inc. (dba Covidien LP) | A software issue may lead to ventilator inoperative situations. | Software design | The firm sent Urgent: Field Correction Action letters, dated October 3, 2014, to consignees. The letter identified the affected product and the reason for the recall. Customers were provided actions to be taken, as well as, important safety reminders. The letter informed customers that they may continue to use their ventilators pending the software correction as long as two gas sources are connected to the ventilator at all times. Customers are to complete the attached acknowledgement and receipt form and fax it to Covidien at the number provided. For further assistance, customers are to contact the Technical Support Department at 1-800-255-6774. |  | 324 | 324 units | Worldwide Distribution -- USA, including the states of CA, CO, FL, GA, KY, MA, MN, NC, NY, OH, OK, PA, SC, TN, TX, UT, WI; and, the countries of Canada, Mexico, Saudi Arabia, South Africa, and United Arab Emirates. | Terminated | May 02, 2016 | 579 | 324 | Computer
Z-2711-2014 | 69235 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | The Ingenuity CT, Ingenuity Core, and Ingenuity Core128 scanners are Computed Tomography X-ray systems intended to produce cross-sectional images of the body by computer reconstruction of X-ray transmission data taken at different angles and planes. This device may include signal analysis and display equipment, patient and equipment supports, components and accessories |  |  | September 11, 2014 | 2014 | Philips Medical Systems (Cleveland) Inc | Ingenuity CT system can mislabel the exposure that results in an image in the incorrect order that it was taken by the user. | Software design | Philips sent an Urgent - Field Safety Notice Medical Device Correction letter dated September 5, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Please review the following information with all members of your staff who need to be aware of the contents of this communication. It is important to understand the implications of this communication. Please retain a copy with the equipment Instruction for Use.Additionally, a Field Service Engineer will contact each site to schedule a time to install the software update.If you need any further information or support concerning this issue, please contact your local Philips representative or local Philips Healthcare office. For North America, contact the Customer Care Solutions Center(1.800.722-9377 and follow the prompts). |  | 79 | 79 | Worldwide Distribution - US including the states of : AZ, CA, CT, FL, GA,MA, MS,OH, PA, TX, and WY., and the countries of : Australia, Denmark, Finland, France, Germany, India, Ireland, Japan, Malaysia, Netherlands, New Zealand, Norway, Russia, Slovakia, South Africa, South Korea, Sweden, Switzerland and United Kingdom. | Open, Classified |  |  | 79 | Computer
Z-0486-2015 | 69602 | No | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | The Artis zee / zeego Angiography System is designed as a set of components that |  |  | September 11, 2013 | 2013 | Siemens Medical Solutions USA, Inc | There is a potential issue on running Artis systems running software VC1x software if a network problem arises, the function cannot be deactivated again by pressing the &quot;Block Radiation&quot; key on the touch screen control. If this behavior occurs, the system image generation function is not available for patient examinations without switching the system off and back on again manually. This | Device Design | Siemens sent an Important Customer Safety Notice in September 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. If the fault described above occurs, we urgently recommend switching the system OFF manually ("System shutdown") and then back ON again. The fault cannot generally be rectified sufficiently, either by executing the "System Restart" or "Application Restart" functions. A corresponding supplement to the Operator Manual is enclosed with this letter.The problem will be solved when the next Service Update is available (currently scheduled for 2014).We thank you for your cooperation in dealing with this safety notice and request that you promptly notify and instruct accordingly all the staff at your organization who have to be aware of this problem. Please also forward this safety information to other organizations that could be affected by this measure. Please observe this safety notice, and comply with the corresponding measures until the update has been fully completed.If this device has been sold and is therefore no longer in your possession, please forward this safety notice to the new owner. We would also request that you inform us of the identity of the device's new owner where possible.For further questions please call (610) 219-6300. |  | 1131 | 1131 | Worldwide Distribution - USA (nationwide) including Puerto Rico., and the countries of: United Arab Emirates, Armenia, Austria, Australia, Belgium, Bahrein ,Belarus, Switzerland, Chile, China, Colombia, Czech Republic,Germany, Denmark, Egypt, Spain, Finland, France, United Kingdom, Hong Kong, Hungary, Indonesia, Ireland, Israel, India, Italy, Jordan, Japan, Kenya, Korea, Kazakhstan, Malaysia, Netherlands, Norway, New Zealand, Philippines, Pakistan, Poland, Palestina, Portugal, Serbia, Russian Federation, Saudi Arabia, Sweden, Singapore, Slovakia, Thailand, Turkey, Taiwan, Ukraine, Vietnam and South Africa. | Terminated | March 30, 2015 | 565 | 1131 | Computer
Z-0825-2014 | 67226 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | ACUSON SC2000 |  |  | December 22, 2013 | 2013 | Siemens Medical Solutions USA, Inc. | The locking mechanism that is intended to hold the control panel in a fixed position can become loose and fail to lock the rotation of the control panel, making it unable to steer or control during transport. | Device Design | Customer Safety Advisory Letters were sent to all affected users on December 22, 2013. Letters advised of the issues and the actions they should take until a Siemens Field Engineer will visit and upgrade the control panel. |  | 271 | 271 units | Worldwide Distribution-USA (nationwide) and the countries of Australia, Austria, Canada, Ecuador, Finland, France, Germany, India, Ireland, Israel, Italy, Japan, Mexico, P.R. China, Philippines, Portugal, Republic Korea, Russian Fed. Saudi Arabia, Serbia, South Africa, Spain, Suriname, Sweden, Taiwan, Turkey, U.A.E., and United Kingdom. | Terminated | March 05, 2015 | 438 | 271 | Computer
Z-2717-2014 | 68329 | No | GFD | Surgical instrument motors and accessories/attachments. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | CelluTome Epidermal Harvesting System Part number CT-H25 or Part Number CT-H50 |  |  | May 19, 2014 | 2014 | KCI USA, Inc. | Blade can shift during shipping which makes the product ineffective. | Device Design | KCI sent an Urgent - Voluntary Medical Device Field Notification letter dated June 6, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.The firm is sending Field Notifications to affected customers instructing them to complete appropriate actions.Option I: Contact your local KCI Representative or KCI Customer Technical Support at 1-800-275-4524 and select Option 3 to arrange for return and replacement of the affected harvester(s).Option 2: If you prefer, you may re-set the blade after removal from the packaging by gently tapping the curved end of the harvester against the palm of your hand, as shown in Figure 3, until the blade is no longer visible in the harvester top plate, as shown in Figure 4. |  | 1331 | 1,331 units | US nationwide Distribution | Terminated | September 18, 2015 | 487 | 1331 | Computer
Z-0143-2015 | 69406 | No | DQK | Programmable diagnostic computer. | Cardiovascular | Cardiovascular | 510(k) | Merge Healthcare, Merge Hemo Programmable diagnostic computer. |  |  | September 29, 2014 | 2014 | Merge Healthcare, Inc. | It has been reported that during use, the SpO2 value displayed on the Hemo Monitor may not update to reflect changes in the patient's oxygen value. It is also possible that if using the pulse rate of the SpO2 finger clip to calculate the patient's heart rate, it too may be subject to displaying a stale value. | Software design | Consignees were e mailed on 5/29/14 a Merge "Important Alert" letter. The letter described the problem and the product involved in the recall. Informed consignees that they had the opportunity to replace the Nellcor SpO2 module with a Masimo SpO2 module free of charge. The letter also provided Clinical Information and the Next Steps to follow. This included to contact Merge Technical Support if they wanted to replace their Nellcor module. For questions, they can contact Lee Harrop, Director Solutions Management, Cardiology, 386-239-2224, Lee.Harrop@Merge.com.A second letter was sent Sept. 11, 2014 via e-mail. Consignees were made aware of a firmware patch available. |  | 52 | 52 | Nationwide Distribution including the states of AZ, CA, FL, IL, IA, MA, MO, OH, SC, TN, and WI. | Open, Classified |  |  | 52 | Computer
Z-0464-2015 | 69515 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | AlignRT- Intended for prescription use. The system is indicated for use during simulation, setup and stereotactic radiosurgery and precision radiotherapy for lesions, tumors and conditions anywhere in the body where radiation is indicated. |  |  | October 15, 2014 | 2014 | Vision Rt Inc | Potential failure of AlignRT to assert interlock. | Software in the Use Environment | The firm, visionrt. sent an "URGENT MEDICAL DEVICE CORRECTION NOTIFICATION-URGENT FIELD SAFETY NOTICE" dated October 15, 2014 to all affected customers on 21-Oct-2014. The notice describes the product, problem and actions to be taken. The customers were informed that they did not need to return their devices; they should pass the letter to all those who need to be aware of it within their organization; promptly inform Vision RT if they believe that patient harm occurred due this issue, and complete and return the acknowledgement form via email to service@visionrt.com.Vision RT is designing an upgrade that will resolve the issue, and will subsequently contact the customer in order to install the upgrade on their device.Should you have any queries on this letter, please do not hesitate to contact Vision RT by telephone on +44 20 83464300 (866 778-2379 from the US) or as per http://www.visionrt.com/contact/details. |  | 7 | 7 (Six Units distribited in the US and One in OUS.) | Worldwide Distribution: US distribution in states of: CA, MN, TX, and UT; and internationally to: Republic of Ireland. | Open, Classified |  |  | 7 | Computer
Z-0218-2015 | 69645 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | CyberKnife Robotic Radiosurgery System with the first generation IRIS Variable Aperture Collimator. |  |  | October 31, 2014 | 2014 | Accuray Incorporated | Software upgrade to correct potential safety issue related to CyberKnife System that occurs when upgrading the Treatment Delivery Software for the first generation Iris Variable Aperture Collimator. | Software design | Accuray sent an Urgent Safety Advisory Notification letter dated October 17, 2014, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to check the radiaiton field size (beam width) for each Iris setting to ensure that it matches the beam commissioning data in the MultiPlan System to within the stated tolerance of 0.2 mm. Customers with questions were instructed to contact Accuray Customer Support at 1-877-668-8667 or visit the Accuray Technical Solutions Center at www.accuray.com or call North America: +1.877.668.8667, Europe: +800.4141.9595, Hong Kong +852.2247.8688, Japan +81.3.6269.9556.For questions regarding this recall call 408-716-4600. |  | 84 | 84 devices subject to correction. | Worldwide Distribution - USA including CA, MS, FL, AZ, DC, TX, CO, CT, IL, WI, RI, NV, OK, MA, AL, NY, MO, NC, TN, PA, DE, NJ, MI, IN, KS, OH. Foreign distribution to Turkey, Switzerland, Spain, Saudi Arabia, Russia, Poland, Mexico and Korea, Italy, India, Greece, Germany, France, England, Canada, and Belgium. | Terminated | February 20, 2015 | 112 | 84 | Computer
Z-2715-2014 | 69084 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | ADVIA Centaur Immunoassay System (including refurbished units), |  |  | August 19, 2014 | 2014 | Siemens Healthcare Diagnostics | The firm has identified an issue with patient demographic information sent to the LIS from the ADVIA Centaur/ADVIA Centaur XP Immunoassay systems. Siemens has confirmed that under extremely rare circumstances patient demographic data from the previous order received from the LIS is merged with the next order. This issue can occur when the LIS data buffer on the ADVIA Centaur system becomes full | Software design | Siemens sent an Urgent Medical Device Correction Letter dated August 19, 2014, via Fed Ex to US customers and via email to regional countries for implementation outside the US (OUS). The letter identified the product the problem and the action needed to be taken by the customer.The letters inform the customers of the issue and specific actions to be taken if the instrument is interfaced to an LIS system that transmits patient demographics with each order. Field service personnel were sent a support bulletin describing the issue and instructing them how to deal with customer questions. Please complete and return the Field Correction Effectiveness Check attached to this letter within 30 days. If you have received any complaints of illness or adverse events associated with the products listed in Table 1, immediately contact your local Siemens Customer Care Center or your local Siemens technical support representative.Please fax this completed form to the Customer Care Center at (302) 631-7597. If you have any questions, contact your local Siemens technical support representative |  | 10515 | Domestic: 2771 units; Foreign: 7744 units | Worldwide Distribution including the United States., and the countries of Argentina, Austria, Australia, Bahrain, Bangladesh, Belgium, Belgrade, Brazil, Botswana, Canada, Chile, China, Colombia, Croatia, Czech Republic, Cyprus, Denmark, Egypt, Estonia, Finland, France, Germany, Georgia, Great Britain, Hungary, Indonesia, India, Ireland, Israel, Italy, Japan, Jordan, Korea, Kazakhstan, Lebanon, Libya, Lithuania, Luxembourg, Macedonia, Malaysia, Malta, Martinique, Mexico, Myanmar, New Caledonia, Netherlands, Norway, New Zealand, Pakistan, French Polynesia, Philippines, Poland, Portugal, Qatar, Romania, Russia, Saudi Arabia, Serbia, Singapore, Slovenia, Slovakia, South Africa, Spain, Switzerland, Thailand, Turkey, United Arab Emirates, Vatican, Venezuela, Vietnam, Yemen and Zimbabwe. | Open, Classified |  |  | 10515 | Computer
Z-2714-2014 | 69194 | No | CAI | Anesthesia breathing circuit. | Anesthesiology | Anesthesiology | 510(K) Exempt | Hudson RCIÂż Pediatric Anesthesia Breathing Circuits, Single Patient Use, Rx Only, |  |  | September 08, 2014 | 2014 | Teleflex Medical | The pediatric breathing circuits can crack prior to and during use. | Device Design | Teleflex Medical sent an Urgent Medical Device Recall dated September 8, 2014, to all affected customers. The letter identified the affected product, problem and actions to be taken. The letter instructed customers to immediately discontinue use, quarantine any products with the product codes and lot numbers provided in the notification, return any affected product, complete and return the enclosed Recall Acknowledgement Form and fax it to the attention of Customer Service. For questions contact your local sales representative or Customer Service at 1-866-246-6990. |  | 27176 | 27,176 units | Worldwide Distribution - US Nationwide in the states of AZ, CA, FL, IL, KY, OH, NJ, NY, TX, LA, OR, SD, WA, including Puerto Rico, and the countries of Belgium, Canada, Chile, China, Colombia, Korea, Mexico, Peru, Venezuela and Zaire. | Terminated | August 16, 2016 | 708 | 27176 | Computer
Z-0651-2014 | 66603 | No | MUH | Extraoral source x-ray system. | Dental | Radiology | 510(k) | The Vantage Digital Panoramic System delivers high-value features standard on every system: - Five pre-programmed exam settings, including bitewing, minimize exam set-up time and allow routine diagnostic exams to be performed extraorally for improved clinical efficiency and patient experience. - Progeny Clarity Enhanced Filter Set provides superior diagnostic imaging quality - Exclusive Focal Trough Alignment reminder assists in consistent patient positioning - Progeny Imaging software included for download on an unlimited number of operatory workstations with no additional costs - Exclusive VantageTrust remote support and training service offered at no charge for the first 60 days of use - Ceph upgradeable for expanded extraoral exam capabilities - The Progeny Vantage Panoramic System s proudly designed and manufactured in the USA; 2013 Midmark Corporation; 675 Heathrow Drive Lincolnshire, IL 60069; 847-415-9800 Toll-free 888-924-3800 |  |  | September 20, 2013 | 2013 | Midmark Corp dba Progeny Inc | An error was identified in the Real Time Controller (RTC) firmware versions 3.0, 3.1, 3.2, and 3.3 used in Vantage Panoramic X-Ray systems. The error may cause the column of the Vantage Panoramic X-Ray unit to continue moving vertically despite the release of the movement control button (i.e. failure to stop"). | Software design | An URGENT MEDICAL DEVICE FIELD CORRECTION letter dated September 20, 2013 was sent to all consignees. Letters sent to distributors include instructions on how to identify the defective units, to quarantine the machine(s), and to contact Midmark at 847-415-9739 or via- e-mail at abavuso@midmark.com. Midmark will either schedule an on-site field correction or arrange for the replacement of the affected device(s) depending on the customer's preference. Midmark further requested the distributors provide Midmark with the names, addresses, phone numbers, and e-mail addresses of any customers that purchased any defective units. Field corrections include the upgrading of RTC firmware to revision 3.4 on the defective units and will be performed as no additional charge by Midmark personnel. Letter sent to practitioners include instructions on how to identify the defective units, instructions on how to temporarily reduce the likelihood of the unintentional movement, and to contact Midmark at 847-415-9739 or via- e-mail at abavuso@midmark.com. Midmark will schedule an on-site field correction by upgrading the firmware to revision 3.4. |  | 58 | 58 units | Worldwide Distribution-USA (nationwide) including Puerto Rico and the states of AK, AL, CA, CO, CT, FL, GA, IL, KS, MA, NC, NH, NJ, OH, OK, PA, SC, TN, TX, VA, and WI, and the countries of Australia, Canada, and Russia. | Open, Classified |  |  | 58 | Computer
Z-0647-2014 | 67132 | No | CAW | Portable oxygen generator. | Anesthesiology | Anesthesiology | 510(k) | On Site Gas Systems POGS 33C Portable Oxygen Generation System |  |  | December 19, 2013 | 2013 | On Site Gas Systems Inc. | Heat exchanger hose becomes weak and may fail. | Device Design | On Site Gas Systems issued letter on 12/19/13 and stated replacement parts to retrofit the blue Federal Heater hose with stainless steel will be provided and the letter included the Technical Bulletin TB 12-19-13-1 . |  | 198 | 198 units | US Distribution: PA only. | Terminated | October 02, 2014 | 287 | 198 | Computer
Z-0840-2014 | 66917 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | GE Optima CT520 and Optima CT540 Computed Tomography (CT) systems. |  |  | October 09, 2013 | 2013 | GE Healthcare, LLC | There is an issue with the Manual Film Composer feature on some CT products. There is an opportunity, while following a specific workflow, to create a film image with one patient's images and another patient's name in the footer. This is also an issue when within Manual Film Composer, if a color image is selected for printing, another patient's black and white image appears in the preview screen | Software design | Consignees were sent on 10/9/2013 a GE Healthcare "Urgent Medical Device Correction" letter dated October 9, 2013. The letter was addressed to Hospital Administrators / Risk Managers, Radiology Directors and Radiologists. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction and Contact Information (Contact your local GE Healthcare Service Representative or the GE Healthcare Call Center : US or Canada : 800-437-1171 and Other countries: 262-896-2890). |  | 272 | 272 (29 US; 243 OUS) | Worldwide Distribution-USA (nationwide) including the states of AL, CO, FL, IN, KS, KY, MN, NC, OH, TX, and VA, and the countries of ALGERIA, AUSTRALIA, BELGIUM, BULGARIA, CANADA, CHINA, ABIDJAN COTEDIVOIRE EGYPT, FRANCE, GERMANY, GREAT BRITAIN, INDIA, IRAQ, ITALY, JAPAN, KENYA, LATVIA, LEBANON, MALAYSIA, MARTINIQUE,MOROCCO, NEW ZEALAND, NORWAY, PAKISTAN, POLAND, PORTUGAL ROMANIA, RUSSIAN FEDERATION, SAUDI ARABIA, SOUTH AFRICA, SPAIN, SWITZERLAND, THAILAND, TUNISIA, TURKEY, UNITED ARAB EMIRATES, and VIETNAM. | Terminated | June 19, 2014 | 253 | 272 | Computer
Z-2737-2014 | 69039 | No | GFC | Surgical instrument motors and accessories/attachments. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Persona The Personalized Knee System 2.5 mm Female Screw 48 mm Length Sterile Qty-2. Sterilized using irradiation, single use. |  |  | September 09, 2014 | 2014 | Zimmer, Inc. | The Persona 48mm X 2.5mm Female Screw is being recalled due to potential use of the screw in cortical bone, and corresponding stripping of the 2.5mm Persona Hex Driver. | Device Design | Zimmer sent an URGENT MEDICAL DEVICE RECALL notifications dated September 9, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.1. Review the notification and the updated portion of the Persona surgical technique provided in Attachment 1, and ensure affected personnel are aware of the contents.2. Locate and quarantine any affected product listed above and notify your Zimmer sales representative.3. Your Zimmer sales representative will remove the recalled product from your facility.4. If after reviewing this notification you have further questions or concerns please call the customer call center at1-877-946-2761 between 8:00 am and 5:00pm EST. |  | 16380 | 16,380 units | Nationwide Distribution including the states of : HI, CA, UT, NY, OH, IN, AL, FL, GA, WI, IA, MN, NE, SD, NJ, PA, ND, IL, MO, VA, MD, NC, MI, OR, WA, KS, KY, TN, WV, OK, CT, MA, ME, AZ, NM, TX, LA, ID, CO, MT, UT, NV and AK., and Internationally to Singapore, Japan, Germany and Canada. | Terminated | May 05, 2015 | 238 | 16380 | Computer
Z-0141-2015 | 69426 | No | JEA | Operating tables and accessories and operating chairs and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | TITAN Operating Room Table |  |  | October 09, 2014 | 2014 | Trumpf Medical Systems, Inc. | Unintended movement of the operating table during surgical cases due to the use of damaged TITAN remotes. | Device Design | TRUMPF sent an Urgent Medical Device Correction letter dated October 3, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Actions to be taken:Complete and return the attached Confirmation of Receipt form with serial numbers of the TITAN table(s) and the serial number(s) for any remote controls you have. We will then arrange to send replacement remotes, which have been redesigned to eliminate this mode of failure. You will then be required to dispose of your recalled remotes.If you have any questions concerning this request or the procedure as outlined, please contact TRUMPF Medical Systems, Inc. at 888-474-9359. |  | 184 | 184 units | Worldwide Disttribution: US (nationwide) including the states of AL, AZ, CT, FL, GA, IA, IL, IN, KS, MD, MI, MN, MO, NC, ND, NJ, NM, NV, NY, OH, OR, PA, SC, TX, VA, WA and WY: and Internationally to Canada. | Terminated | March 23, 2016 | 531 | 184 | Computer
Z-0079-2015 | 63897 | No | FSA | Non-AC-powered patient lift. | General Hospital | General Hospital | 510(K) Exempt | Medium 2-Hooks Spreader Bar for use with Loop Slings, sizes small, medium and large on the Maxi Sky 2 Ceiling Lift; Model 700-19425. |  |  | January 21, 2013 | 2013 | Arjo, Inc. dba ArjoHuntleigh | The upper portion of the hooks on the spreader bar are too flexible for holding the loops of the sling. The loop of the sling could potentially come out of the hooks if positioned in the upper portion of the hooks. | Device Design | URGENT FIELD CORRECTION RECALL letters dated January 21, 2013 were sent to all direct accounts (customers). The letters identified the potential safety related issue; included instructions on how to use the Medium Two Hooks Spreader Bars, and to complete and return the enclosed Customer Response form. Once the Customer Response form has been returned customers will be contacted by an ArjoHuntleigh service representative to make arrangements for the replacement of the existing spreader bars with spreader bars of a different design. Customers with any questions about this recall can contact ArjoHuntleigh by telephone at 800-323-1245, extension 57985, or by e-mail at usa.quality@arjohuntleigh.com. |  | 363 | 363 spreader bars | USA Distribution - AR, GA, MI, PA, and the District of Columbia. | Terminated | October 27, 2015 | 1009 | 363 | Computer
Z-1469-2014 | 67569 | No | HAW | Stereotaxic instrument. | Neurology | Neurology | 510(k) | Brainlab, Spine &amp; Trauma 3D 2.0, Navigation Software. An intraoperative image-guided localization system to enable minimally invasive surgery. Catalog Number: 22264 |  |  | November 18, 2013 | 2013 | Brainlab AG | The "fluoro match registration" function of the Navigation Software Spine &amp; Trauma 3D 2.0 allows the user to intraoperatively match e.g., CT data sets to the current patient anatomy visible on fluoroscopic images, used by the navigation software during spinal surgeries to display the position and orientation of instruments. For this registration function in combination with a digitally integrated | Software design | Brainlab sent a FIELD SAFETY NOTICE / PRODUCT NOTIFICATION letter dated November 18, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.The letter included instructions for the modification of C-arm configurations: 1) If possible, Use analog image transfer between C-arm and Brainlab Navigation System and to consult with the local Brainlab Customer Support Representative for this modification, and 2) If possible, adjust the image size of the C-arm to 8 bit and to follow the instructions provided with the C-arm and if required to consult the C-arm manufacturer. For measures when performing a Fluoro match registration, the following instructions are provided to reduce the probability that the algorithm results in an inaccurate match: 1) perform the preoperative CT scans according to the descriptions in the Brainlab scan protocol, 2) set the threshold so that a smooth bone surface is displayed and the quality of the chosen bone threshold influences the accuracy of the match result (refer to section "Bone Threshold" in the Software User Guide"), 3) use the Radiolucent Spine Reference Clamp instead of the Spine Reference X-Clamp, 4) make sure that no additional objects (e.g. wires, retractors, instruments, etc.) are within the view of the C-arm. Additional reminders in the letter include: 1) always verify registration accuracy by holding the pointer or instrument tip to at least three anatomical landmarks and verifying their position in the software, 2) verify that the registration is at the correct level on the patient and data set, and 3) accuracy must be checked on the treated bone surface. Brainlab will provide a software update with the issue solved to affected customers tentatively by January 2014. Customers with questions can contact the Customer Hotline at 800-597-5911 or via e-mail at us.support@brainlab.com. |  | 133 | 133 units | Worldwide Distribution - US Distribution including the states of AZ, CA,GA, IL, IN, LA, MA, MI, NY, NC, OH, PA, TN, TX and WI and the countries of Australia, Austria, Belgium, Brazil, Chile, Czech Republic, France, Germany, Italy, Japan, Malaysia, Netherlands, Norway, Philippines, Poland, Russian Federation, Singapore, South Africa, Spain, Sweden, Switzerland, Thailand, Turkey and United Kingdom. | Open, Classified |  |  | 133 | Computer
Z-0902-2014 | 67191 | No | JCX | Powered suction pump. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Neptune 2 Rover Ultra Waste Management System. |  |  | December 30, 2013 | 2013 | Stryker Instruments Div. of Stryker Corporation | The purpose of this correction is to upgrade the Neptune 2 Waste Management Systems in the field to the recently obtained 510(k) clearance specifications. | Device Design | On 12/30/13, Stryker sent an URGENT MEDICAL DEVICE UPDATE notification to consignees including government accounts and Sales Representatives, Stryker Canada and 3 in-house accounts and 1 account for lost rovers. Domestic and government consignees were sent the notification by mail FedEx overnight, and Sales Representatives and Stryker Canada were sent the notification by email. Notification described the affected product, instructions for the consignees, and upgrade information. Stryker also stated that a training kit and in-servicing will be provided to consignees. For questions regarding this recall, please contact Stryker Instruments: Monday-Friday 8am-5pm (EST)Neptune Customer Care Center 855-458-7441 or 269-389-2316strykerinstrumentsrecalls@stryker.com |  |  |  | Nationwide and Canada, Germany, Australia, New Zealand | Terminated | January 12, 2016 | 743 |  | Computer
Z-1221-2014 | 67598 | No | DWF | Cardiopulmonary bypass vascular catheter, cannula, or tubing. | Cardiovascular | Cardiovascular | 510(k) | Sarns Malleable Dual-stage Venous Return Cannulae, 34/46 Fr with 1/2" connector, 15" (38 cm) long |  |  | March 03, 2014 | 2014 | Terumo Cardiovascular Systems Corporation | During in-process inspection, Terumo Cardiovascular Systems (Terumo CVS) identified exposed basket wire on certain product lots of Sarns Malleable Dual-stage Venous Return Cannulae and Sarns Dual-stage Venous Return Cannulae. | Device Design | Terumo sent an Urgent Medical Device Recall letter dated March 3, 2014, to all affected consignees. The letter included the reason for removal, potential hazards, and customer instructions. Terumo CVS advised users to stop using the affected devices and return any unused product on hand. Questions or concerns were directed to Terumo CVS Customer Service: 1-800-521-2818 M-F, 8am-6pm.For questions regarding this recall call 734-741-6173. |  |  |  | Worldwide Distribution - USA including CA, CT, FL, GA, IL, IN, LA, MI, MN, MO, NC, NY, PA, TN, TX, and Internationallly to United Arab Emirates (UAE), Taiwan, Thailand, Malaysia, Belgium, Japan, Mexico, and Canada. | Terminated | August 12, 2014 | 162 |  | Computer
Z-1195-2014 | 67447 | No | NQI | Immunomagnetic circulating cancer cell selection and enumeration system. | Immunology | Pathology | 510(k) | CELLTRACKS ANALYZER IIÂż (Product Code 9555) |  |  | January 08, 2014 | 2014 | Veridex, LLC | Janssen Diagnostics, LLC received complaints of camera failures on the CELLTRACKS ANALYZER IIÂż. | Device Design | Janssen Diagnostics, LLC sent an Important Product Correction Notification letter dated January 8, 2014 to all affected customers via FedEx overnight courier letter and Foreign Affiliate Consignees by email on January 8, 2014. The letter identified the affected product, problem and actions to be taken. Customers were instructed to complete and return the Confirmation of Receipt Response form attached. For questions contact Customer Technical Service at 1-877-837-4339. |  | 19 | 19 units | Worldwide Distribution - USA Nationwide | Terminated | July 06, 2017 | 1275 | 19 | Computer
Z-1746-2014 | 68349 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | IntelliVue MX40 Wearable Patient Monitor in use with the Philips IntelliVue Information Center iX (PIIC iX); 865350, 865351, 865352 |  |  | May 22, 2014 | 2014 | Philips Medical Systems, Inc. | IntelliVue MX40 (part numbers 865350, 865352) with software revisions B.00 or B.01 may not resume alarming when the alarms are paused and the pause timer ends. | Software design | On May 22, 2014, each affected customer was sent an Urgent Medical Device Correction notification/Field Safety Notice. The letter explains the problem, explains the action to be taken by the consumer, and the action planned by Philips. The Urgent Medical Device Correction notification/Field Safety Notice informs customers of the issue, identifies details of the units affected, gives instructions on actions to be taken by the customer and identifies what action Philips plans to take to remedy the issue.The correction will consist of a free of charge software upgrade. A Philips Healthcare representative will contact customers with affected devices to arrange for service. Philips is asking customers to follow the Action to be Taken by Customer/User section of the Urgent Medical Device Correction notification/Field Safety Notice:The following two items must be configured when using MX40 with the IntelliVue Information Center iX: 1.Utilize the default state of Pause Alarms at Yellow only, which disables the ability to Pause/Suspend all Red level alarmsUtilize the default state of Remote Suspend Alarms as Off, which disables the use of the Pause Alarms feature at the MX40 device |  | 42600 | 42,600 | United States: NationwideForeign Countries:Germany, Austria, Finland, Sweden, Australia, Qatar, Canada, Netherlands, New Zealand, Portugal, United Kingdom, Spain, Belgium, France, Switzerland, Norway, Czech Republic, Poland, Italy, Singapore, Denmark, Iceland, India, Saudi Arabia, Ireland, Oman, Bahrain, Japan, Israel, Malaysia, South Africa, Turkey, Thailand, Latvia, Bulgaria, Indonesia, Slovenia, United Arab Emirates, Hong Kong, Aruba, Chile, Lebanon, Korea, Kuwait, Taiwan, Argentina, Hungary, Luxembourg, Mexico, Gabon | Open, Classified |  |  | 42600 | Computer
Z-1197-2014 | 67625 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | XiO RPT System |  |  | February 17, 2014 | 2014 | Elekta, Inc. | Dose and Monitor Unit values are not computed correctly when Elekta Motorized Wedges are used with Dynamic Conformal arcs. | Software design | Elekta, Inc. issued an Important Safety Notice (LINXi000019) issued February 14, 29, 2014. The Safety Notice identified the product, the problem, and the action to be taken by the customer. All affected customers of the issue globally by February 28, 2014. Provide a fix to the defect in an expeditious timeframe. It is mandatory for all customers to apply the fix. Version 5.00.01 will be released for the fix in March 2014. Follow up with the affected customers for confirmation receipt of the LIN, and confirmation of installation of patch for the fix by November 30, 2014.For questions regarding this recall call 770-670-2422. |  | 1776 | 1776 | USA including AL, AK, AZ, CA, CO, CT, DE, DC, FL, GA, IL, IN, KS, KY, LA, ME, MD, MA, MI, MN, MS, MO, MT, NE, NV, NJ, NY, NC, ND, OH, OK, PA, RI, SC, TN, TX, VA, WA, WV, WI, WY, Puerto Rico and Internationally to Albania, Algeria, Argentina, Australia, Austria, Bahamas, Barbados, Bermuda, British Guiana, Boliva, Brazil, British West Indies, Burundi, Canada, Chile, China, Columbia, Costa Rica, Cuba, Curacao, Cyprus, Czechoslovakia, Dominica, Equator, Estonia, Egypt, Finland, France, Gambia, Germany, Great Britain, Georgia, Greece, Guatemala, Hong Kong, India, Indonesia, Iraq, Ireland, Israel, Italy, Japan, Jordan, Kuwait, Lativa, Lithunia, Madagascar, Malaysia, Malta, Mexico, Mongolia, Myanmar, Netherlands, New Zealand, Nigeria, Pakistan, Peru, Philippines, Poland, Portugal, Romania, Russia, San Salvador, Saudia Arabia, Serbia, Singapore, Sir Lanka, South Korea, Spain, Sudan Slovenia, Suriname, Tajikistan, Thailand, Turkey, Turkmenistan, Taiwan, Ukraine, Uruguay, Venezuela, Viet Nam, Zambia. | Terminated | July 21, 2015 | 519 | 1776 | Computer
Z-1741-2014 | 68269 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | enGen (TM) Laboratory Automation System Configured with TCAutomation" Software v3.5 -- For in vitro quantitative measurement of a variety of analytes of clinical interest --- |  |  | April 17, 2014 | 2014 | Ortho-Clinical Diagnostics | Software anomaly: potential buffer timeout issues when using enGen (TM) Laboratory Automation Systems Configured with TCAutomation (TM) Software v3.5. | Software Design Change | Ortho Clinical Diagnostics sent an Important Product Correction Notification dated April 17, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.On April 17, 2014, foreign affiliates were informed by e-mail of the issue and instructed to notify their customers of the issue and actions.The letters instruct customers: to ensure that samples prematurely routed to Check are detected and processed appropriately, please monitor completion of all pending tests following standard laboratory practice. Customers were informed of the issue and advised to: -- Routinely monitor pending result reports from your Laboratory Information System (LIS) to determine if results are pending for samples. -- Investigate to determine the next action identified for a sample if the sample is routed to the Check location prematurely. If you have any additional questions, please call Customer Technical Services at 1-800-421-3311. |  | 9 | Domestic: 6 units, Foreign: 3 units | Worldwide Distribution - US including the states of Arizona, Delaware, Ohio, Tennessee, and North Carolina; and the countries of Mexico, England and Italy. | Terminated | April 11, 2017 | 1090 | 9 | Computer
Z-1739-2014 | 68135 | Yes | KPO | Hemodialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | NaturaLyteÂż Liquid Bicarbonate Concentrate, 6.4 liter bottle to be used with a hemodialysis unit; Part Number: 08-4000-LB. |  |  | May 01, 2014 | 2014 | Fresenius Medical Care Holdings, Inc. | Container cap may contain foreign material- 0.60% - 0.80% Manganese | Equipment maintenance | Fresenius Medical notified customer by letter dated 5/1/14 to examine stock immediately and instructed to discontinue use and place all Naturalyte Liquid Bicarbonate Concentrate in a secure area for return to FMC-RTG. The customer instructed to contact the Fresenius Medical Care Customer Service Team for instructions on how to return the recalled product. |  | 12 | 12 bottles | Distributed in the state of MO. | Open, Classified |  |  | 12 | Computer
Z-1064-2014 | 65450 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | SL &amp; SLi Series (Mod 15/18/20/25) Linear Accelerator |  |  | May 28, 2013 | 2013 | Elekta, Inc. | Customers are resetting the default values outside recommended factory settings. | Device Design | The firm, Elekta, sent an "Important Field Safety Notice (IFSN 200 01 103 073 - Electron Applicator Auto-Tracking Defaults)" released on May 28, 2013 and is in the process of being distributed to all customers. The Notice requests users to please adhere to the instructions and advice given in manuals and system dialogues, which explicitly state that a change to auto-tracking values could compromise compliance with safety standards and can cause increased unwanted radiation dose to the patient. All machines should be checked against the default settings attached. If the actual settings are greater than the default settings, then Elekta's recommendation is that the default settings should b re-applied and all relevant physics checks (e.g. beam uniformity, output factor) performed. If, however, auto-tracking values are increased beyond their factory-set defaults, then electron applicator radiation leakage measurements should also be made to ensue that the system remains compliant to IEC safety. All users of Desktop Pro and Integrity 1.1 software will receive a mandatory upgrade to Integrity 1.2 (released 28-Jan-14. The upgrade path to Integrity 1.2 differs depending on system configurations at each customers site, therefore Elekta has released three separate Important Field Safety Modifications in order to upgrade all customers successfully. Important Field Safety Notice 200 01 501 036  Important Field Safety Modifications 200 02 501 037  Applicable System Configuration Desktop Pro with Fast Tuning Magnetron, New Patient Support System; Important Field Safety Notice 200 01 501 038  Important Field Safety Modification 200 02 501 039  Applicable System Configuration Desktop Pro with Linac Control System MK3i, Fast Tuning Magnetron, New Patient Support System; Important Field Safety Notice 200 01 502 035  Important Field Safety Modification 200 02 502 036  Applicable System Configuration All Integrity 1.1 sites. All users of Integrity 3.1 will receive a man |  | 449 | 449 | Worldwide Distribution: US (Nationwide) including states of: AK, CA, CT, FL, IL, IN, IA, LA, MD, MA, MI, MN, NJ, NY, OH, OK, OR, PA, TN, TX, VA, WA, WV, and Puerto Rico; and countries of: Argentina, Australia, Austria, Belarus, Belgium, Bosnia &amp; Herzegovinia, Brazil, Canada, China, Czech Republic, Denmark, France, Germany, Greece, Hong Kong, Hungary, India, Indonesia, Iran, Ireland, Italy, Japan, Jordan, Libya, Mexico, Netherlands, New Zealand, Norway, Panama, Philippines, Poland, Portugal, Russia, Senegal, Serbia, Singapore, Slovenia, South Africa, Spain, Sudan, Sweden, Switzerland, Taiwan, Turkey, and United Kingdom. | Open, Classified |  |  | 449 | Computer
Z-1736-2014 | 68271 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | VITROS 5,1 FS Chemistry System and VITROS 5,1 FS Refurbished, Product Codes/Catalog Numbers 6801375, 6801890 for in vitro diagnostic use. |  |  | April 10, 2014 | 2014 | Ortho-Clinical Diagnostics | Potential for biased results when a Multiple Window (MW) code is associated with results generated by the VITROS 4600 Chemistry System, the VITROS 5600 Integrated System and the VITROS 5,1 Chemistry System. MW codes may be associated with results from any of the following VITROS Chemistry Products using multiple point (multipoint) rate measurements. The MicroSlide assays that could be affected | Software design | Important Product Correction Letters (CL2014-116, dated 4/10/2014) were sent on 4/10/2014 via FedEx overnight courier or ORTHO PLUS e-Communications to all customers with VITROS 5,1 FS, 5,1 FS Refurbished, 4600,and/or 5600 Systems. They were informed of the issue and advised to: -- Review VITROS ALT, ALTJ, ALKP, ALKPJ, AST, ASTJ, CHE, CHEJ, CK, CKJ, GGT, GGTJ, LDH and LDHJ results generated on VITROS 4600/5600 and 5,1 FS Systems. Do not report any results associated with an MW code. -- Perform repeat testing in singlicate on all samples with results having an MW code. Only results without an MW code are acceptable. -- Review their current Laboratory Information System (LIS) configuration to determine if changes are required in order to restrict reporting of results with an MW code. Customers with questions should contact Customer Technical Services at 1-800-421-3311. --- Foreign affiliates were informed of the issue by e-mail on 4/10/2014 and instructed to notify their consignees of the issue and actions. |  | 1978 | Domestic - 597 units, Foreign - 1381 units | Worldwide distribution including US Nationwide, Puerto Rico, Australia, Brazil, Canada, Chile, China, Colombia, England, France, Germany, India, Italy, Japan, Mexico, Panama, Singapore, Spain, and Venezuela. | Open, Classified |  |  | 4029 | Computer
Z-1738-2014 | 68270 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | enGen Laboratory Automation Systems (enGen) |  |  | April 10, 2014 | 2014 | Ortho-Clinical Diagnostics | Software anomaly: potential for delay in the reporting of patient sample test results when using enGen with TCAutomation Software v2.6 , v3.2 and v3.5 Configured with Bypass Modules for ADVIA Centaur, Abbott Architect, VITROS and/or Tosoh AIA Systems. | Software design | Urgent Product Correction Letters (CL2014-114, dated 4/10/2014) were sent to all US customers and Foreign affiliates starting on 4/10/2014. The letter informed customers of the issue and provided required actions to be taken: Routinely monitor your Bypass module(s) -- If you detect a stoppage due to the software timing issue, follow the steps in the How to Detect the Bypass Module Time Issue and Release the Samples section on Page 2 of this notification. -- Complete the Confirmation of Receipt Form and return by 4/23/2014. -- Place this notification near each system or with your enGen (TM) Laboratory Automation System user documentation. |  | 108 | Domestic - 31 units; Foreign - 77 units | Worldwide Distribution: US Nationwide including Puerto Rico, and countries of: Australia, Bermuda, Brazil, Canada, Chile, China, Colombia, England, France, Germany, India, Italy, Japan, Mexico, Panama, Singapore, Spain, and Venezuela | Open, Classified |  |  | 108 | Computer
Z-1215-2014 | 66787 | No | NHA | Endosseous dental implant abutment. | Dental | Dental | 510(k) | Dental Abutment labeled in part: "4555 Riverside Drive Palm Beach Gardens, FL 33410 USA 1-800-342-5454***EDAZ , EDAZx and ILDExxx***BellaTek Zirconia Abutment Position Number: XX Abutment; Pilier; Abutment; Abutment; Pilastro; Pilar***Rx only***Non-sterile, Sterilize prior to use.***Biomet 3i Dental Iberica S.L.*** |  |  | August 26, 2013 | 2013 | Biomet 3i, LLC | Reports relating to Biomet 3i's BellaTek Zirconia Dental Abutment fractures were received. | Device Design | The firm, BIOMET 3i, sent an "URGENT: Medical Device Removal - BellaTek Zirconia Abutments" letter dated August 2013 via email, facsimile or postal mail to its customers. The letter described the product, problem and actions to be taken. All customers were instructed to check their inventory for the affected product; return the abutments to BIOMET 3i for replacement and full credit; complete and return the attached Recall Return Response Form, and to advise their customers and forward a copy of the Medical Device Removal Notice and the contained recommendations.BIOMET 3i has decided to initiate a design change to mitigate the fracture risk, and consequently, will cease distribution of BellaTek Zirconia Abutments unit l the corrective actions are completed.If you have any further questions, please contact Customer Service at 1-800-342-5454 for assistance. |  | 10454 | 10,454 | Worldwide Distribution: US (nationwide) including states of: AK, AL, AR, AZ, CA, CO, CT, DE, FL, GA, IA, ID, IL, KS, KY, LA, MA, MD, MI, MN, MO, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OR, PA, RI, SC, SD, TN, TX, UT, VA, WA, WI, WV, Hawaii and Puerto Rico; and countries of: Spain, Australia, Canada, Colombia, Japan, Uruguay, Mexico, Estonia, United Kingdom, Denmark, Germany, Spain, Italy, Poland, Ireland Netherlands, France, Sweden, Portugal, Australia, France, Belgium, Switzerland, Norway, Czech Republic, France, Holland, and Luxembourg. | Terminated | September 26, 2014 | 396 | 10454 | Computer
Z-1729-2014 | 68237 | No | LDS | Medical washer. | General Hospital | General Hospital | 510(K) Exempt | HLD Systems Model 610 Washer/High Level Disinfector. |  |  | April 24, 2014 | 2014 | Cenorin, LLC | Cenorin LLC is voluntarily recalling HLD Systems Model 610 Washer/High Level Disinfector due to a potential risk of the system not injecting the proper amount of cleaning fluid during the wash cycle and not achieving the desired level of debris reduction. | Equipment maintenance | Cenorin sent a letter on April 24, 2014 informing their customers of the potential failure of the pump. This letter also explains the customer's option of upgrading their entire HLD System to the new cleaning fluid delivery system or to continue purchasing a six month PM kit that has a replacement pump added. Customers are instructed to complete and return an enclosed response form within 15 days and to email Jamie Wright at jwright@cenorin.com or call 800-426-1042, extension 224 for questions. |  | 25 | 25 HLD systems | Distribution USA Nationwide | Terminated | June 30, 2015 | 432 | 25 | Computer
Z-1728-2014 | 68431 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Horizon Medical Imaging |  |  | May 29, 2014 | 2014 | Mckesson Medical Immaging | Firm received a complaint that an unreported study performed that day could not be opened from the archive. | Software design | McKesson sent an Urgent Field Safety Notice dated May 20, 2014, to all affected customers. The notice identified the product, the problem, and the action to be taken by the customer. Customers were informed that the potentially affected software versions will be corrected without removing the systems from their location. Customers were instructed to distribute the notice to all those within their organization who need to be aware and alert all those affilliated with their organization that are affected by this field safety notice. All customers receiving the Field Safety Notice should contact McKesson Support as soon as possible at 1-800-663-2533 to arrange installation of the software update. Customers with questions should contact the Customer Support Department and ask to speak to their Support Manager. |  | 659 | 659 | Worldwide Distribution - USA including AL, AK, AZ, AR, CA, CO, CT, FL, GA, HI, ID, IL, IN, IA, KS, KY, LA, ME, MD, MA, MI, MN, MS, MO, MT, NE, NV, NH, NJ, NM, NY, NC, ND, OH, OR, PA, RI, SC, TN, TX, VT, VA, WA, WV, WI, WY, and Puerto Rico. Internationally to Australia, Canada, France, Great Britain,and Guam. | Terminated | March 31, 2017 | 1037 | 659 | Computer
Z-1216-2014 | 67559 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | syngo Dynamics Picture Archiving and Communication System (PACS). Model numbers 10091804, 10091805, 10091806, 10091807, 10091673 |  |  | January 27, 2014 | 2014 | Siemens Medical Solutions USA, Inc | Potential patient data mixup. Improper error handling could allow for two different patients to be displayed in syngo Dynamics. Images shown for one patient and reporting system shown for a different patient, leading to the potential for misdiagnosis. | Software design | A recall notification letter, dated January 27, 2014, was sent to end Users. Customer Safety Advisory Letter SY016/14/S. Siemens is preparing a Service Patch modification of syngo Dynamics v9.5 which will be available in March 2014. |  | 50 | 50 | Nationwide distribution to MN, AR, MA, TX, NJ, WV, GA, CA, PA, OH, MI, WI, TN, NY, KS, CT, NY, MO, IL, IN, OR, IA, and WA. | Terminated | March 16, 2015 | 413 | 50 | Computer
Z-1712-2014 | 68326 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | PATH(R) THREAD CUP ADAPTOR, REF 2007-0013, Wright Medical Technology, Inc. Orthopaedic surgical instrument. |  |  | April 30, 2014 | 2014 | Microport Orthopedics INC. | Difficulty removing the cup adaptor from the impaction handle during surgery | Device Design | Firm has initiated the recall by email on 04/28/2014 to all distributor's offices, including international, then followed with written communication on 05/02/2014 via Fed Ex P1 delivery to all United States consignees. Follow-up to non-responding consignees will commence 30 days after initiation. |  | 140 | 140 units | Nationwide Distribution. | Terminated | December 01, 2014 | 215 | 140 | Computer
Z-1234-2014 | 67584 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | MEDITECH Microbiology; Client Server Releases 5.6, 5.65, 5.66, 5.67, 6.0, 6.05, 6.07, 6.1, 6.13 and 6.14. |  |  | February 21, 2014 | 2014 | Medical Information Technology, Inc. | Potential for erroneous result reporting. | Device Design | Beginning on February 3, 2014 MEDITECH distributed a notification to Client Server customers who have the Microbiology product in TEST and/or LIVE only for the releases noted in this correction report. This notification was made via e-mailed task updates that can be printed by the customer. Task updates sent via e-mail are immediately transmitted to the customer. The issue will be corrected by CS LAB DTS 12279. |  | 438 | 438 | Distributed USA (nationwide) and Canada. | Terminated | November 12, 2015 | 629 | 438 | Computer
Z-1783-2014 | 64507 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Philips IntellVue Info Center iX 866023; PIIC iX Upgrade 866024; IntellVue Server iX Upgrade 866026; PIIC Classic Upgrade 866117 Patient Physiological Monitor. |  |  | March 05, 2013 | 2013 | Philips Healthcare Inc. | If a customer creates customized trend scales in the trend review tile and the iX or primary server reboots for any reason, the iX(s) will enter a reboot loop. If one iX reboots, the issue will impact only those patients monitored on that device. If the primary server reboots, all patients across all iX devices will be impacted. | Software design | On 3/5/13, Philips sent each known affected customer an Urgent Medical Device Correction notification/Field Safety Notice to inform them that Philips has received a customer report that the PIC iX (release A.01.XX) central stations that are connected to a database server (DBS) may not be able to complete a restart/reboot resulting in the loss of central monitoring. In order to prevent this from occurring, customers were advised to not modify/customize the default trend scales for reviewing stored patient data in the retrospective review applications. Recall letter also states that a Philips Healthcare representative will contact customers with affected devices to arrange for the installation of updated PIIX iX software resolving this issue on affected units. Philips will conduct these updates for all affected devices at no charge. |  | 899 | 899 | Worldwide Distribution: US (nationwide) and country of: Canada. | Terminated | May 17, 2017 | 1534 | 899 | Computer
Z-1711-2014 | 68110 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | AXIOM Luminos dRFsystem |  |  | April 11, 2014 | 2014 | Siemens Medical Solutions USA, Inc | It was determined that under rare environmental conditions (i.e., extremely dry air in combination with a highly isolated floor) electrostatic discharges (ESD) &gt;8 kV may trigger an unintended movement of the AXIOM Luminos dRFsystem which may lead to an emergency situation and pose danger to the patients, operating personnel or to the unit. | Device Design | Siemens sent an Safety Advisory Notice dated April 11, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.What steps can the user take to avoid the potential risk of this issue?Press one of the red emergency STOP buttons immediately in case of an unintended movement. The system needs to be restarted afterwards to be fully functional again.How will the issue finally be resolved?Siemens is preparing a modification of the Luminos dRF system control console that will resolve this potential malfunction. The field modification will be available by end of May 2014.We appreciate your understanding and cooperation with this safety advisory notice and ask you to immediately instruct your personnel accordingly. Please ensure that this safety advisory notice is placed in the system's instructions for use until the update has been installed.If you have sold this device/equipment and it is no longer in your possession, we kindly ask that you forward this safety notice to the new owner of this device/equipment. Please inform us about the new owner of the device/equipment.For further questions please call (610) 219-6300. |  | 37 | 37 | Class II Recall - US Distribution including the states of IL,VA, TN, PA, TX, MN, GA, MI, IN, OH, FL, CA, OR, NY, ME, MO, OK, CO and NJ. | Terminated | December 15, 2014 | 248 | 37 | Computer
Z-1231-2014 | 67630 | Yes | BTS | Bronchial tube. | Anesthesiology | Anesthesiology | 510(k) | Rusch Brochopart and Rusch Brochopart White |  |  | February 28, 2014 | 2014 | Teleflex Medical | The product may fail to achieve seal of right lung due to the cuff inflating to one side | Device Design | Teleflex sent a Urgent Medical Device Recall Notification dated February 28, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Customers were instructed:1. If you have affected stock, immediately discontinue use and quarantine any products with the catalog numbers listed above.2. To return product, complete the enclosed Recall Acknowledgement Form and fax it to 1-866-804-9881, Attn: Customer Service. This will allow us to document the amount of product you have on hand for return. A customer service representative will contact you with a Return Goods Authorization (RGA) Number and will provide instructions for the return of product to Teleflex Medical.3. If you have no affected stock, please complete the enclosed Recall Acknowledgment Form and fax it to 1-866-804-9881, Attn: Customer Service. This will allow us to document your receipt of this letter.Teleflex Medical is committed to providing high quality, safe and effective products. We sincerely apologize for any inconvenience this action may cause your operations. If you have any other questions, feel free to contact your local sales representative or Customer Service at 1-866-246-6990. |  | 63568 | 63,568 ea. | Worldwide Distribution - USA (nationwide) Distribution including Puerto Rico and the states of AL, AZ, CA, CO, CT, DC, FL, GA, ID, IL, IN, KS, KY, LA, MD, MA, MI, MN, MS, MO, NH, NJ, NM, NY, NC, ND, OH, OK, OR, PA, RI, SC, SD, TN, TX, VA, WA, WV and WI,and the countries of United Arab Emirates, Austria, Australia, Azerbaijan, Belgium, Belgrade, Bahrain, Canada, Switzerland, Cameroon, China, Costa Rica, Cyprus, Czech Republic, Germany, Denmark, Dominican Republic, Algeria, Ecuador, Estonia, Egypt, Spain, Finland, France, Great Britain, Greece, Guyana, Croatia, Hong Kong, Hungary, Ireland, Israel, India, Iraq, Italy, Japan, Jordan, Kenya, Kuwait, Lebanon, Latvia, Libya, Morocco, Serbia, Malaysia, Martinique, Nigeria, Netherlands, Norway, Oman, Panama, French Polynesia, Philippines, Poland, Portugal, Qatar, Reunion, Russia, San Salvador, Saudi Arabia, Sweden, Singapore, Slovenia, Slovakia, San Marino, Thailand, Trinidad, Tunisia, Turkey, Venezuela, Yemen and Zambia | Open, Classified |  |  | 63568 | Computer
Z-1706-2014 | 68219 | No | KPY | Personnel protective shield. | Radiology | Radiology | 510(K) Exempt | Kenex Radiation Shield and surgical lamps installed with GE Healthcare Interventional, Surgery, CT and X-Ray Imaging systems. |  |  | May 02, 2014 | 2014 | GE Healthcare, LLC | GE Healthcare has recently become aware of a potential safety issue due to the improper installation of a Kenex ceiling suspended radiation shield and surgical lamp. This issue is not related to the design of the Kenex product. The affected products are limited to installations completed or contracted by GE Healthcare.There has been a reported incident of a radiation shield falling from the ceil | Equipment maintenance | GE Healthcare sent an "Urgent Medical Device Correction" letter # 12209 dated May 2, 2014. The letter was addressed to Director of Biomedical Engineering, Director of Radiology and Chief of Cardiology. The letter described the Safety Issue, Safety Instructions, Affected product Details, Product Correction and Contact Information. For questions they should call 1-800-345-2700. For other countries, they should contact their local GE Healthcare Service Representative. . |  | 3221 | 3221 (108 US, 3113 OUS). | Worldwide Distribution - US Nationwide, ALBANIA, ALGERIA, ANGOLA, ARGENTINA, AUSTRALIA, AUSTRIA, AZERBAIJAN, BAHRAIN, BELARUS, BELGIUM, BOLIVIA, BOSNIA &amp; HERZEGOVINA, BRAZIL, BULGARIA, CANADA, CHILE, CHINA, COLOMBIA, COSTA RICA, CROATIA, CZECH REPUBLIC, DENMARK, DOMINICAN REPUBLIC, ECUADOR, EGYPT, FINLAND, FRANCE, GEORGIA, GERMANY, GREECE, GUADELOUPE, GUATEMALA, GUYANA, HONDURAS, HONG KONG, HUNGARY, ICELAND, INDIA, INDONESIA, IRAN, IRAQ, IRELAND, ISRAEL, ITALY, JAMAICA, JAPAN, JORDAN, KAZAKHSTAN, REPUBLIC OF KOREA, KUWAIT, LATVIA, LEBANON, LIBYAN ARAB JAMAHIRIYA, LITHUANIA, MACEDONIA, MALAYSIA, MALTA, MARTINIQUE, MAURITIUS, MEXICO, MOROCCO, NETHERLANDS, NEW ZEALAND, NICARAGUA, NIGERIA, NORWAY, OMAN, PAKISTAN, PANAMA, PARAGUAY, PHILIPPINES, POLAND, PORTUGAL, REUNION, ROMANIA, RUSSIA, SAUDI ARABIA, SENEGAL, SERBIA, SINGAPORE, SLOVAKIA, SLOVENIA, SOUTH AFRICA, SPAIN, SUDAN, SWEDEN, SWITZERLAND, SYRIA, TAIWAN, THAILAND, TUNISIA, TURKEY, UKRAINE, UNITED KINGDOM, VENEZUELA, VIETNAM, and YEMEN. | Terminated | February 25, 2015 | 299 | 3221 | Computer
Z-1707-2014 | 67926 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | ExacTrac 5.5, an Image Processing System used for patient positioning for radiation therapy or radiosurgery. |  |  | April 15, 2014 | 2014 | Brainlab AG | When using a specific plan update workflow in combination with a non-Brainlab treatment planning system (TPS), there is a possibility in ExacTrac 5.5.5 to generate an unintended blend of two initially separate plans of the same patient. | Software design | Brainlab sent a "FIELD SAFETY NOTICE / PRODUCT NOTIFICATION" letter dated March 13, 2014 was sent to all affected customers. The letter identifies the affected product, problem, and actions to be taken. The letter includes specific instructions to users that are intended to reduce the likelihood of similar events occurring when using ExacTrac software version 5.5.5. The letter further states that Brainlab will be providing a software update to all customers. Brainlab will contact all customers to arrange for the software update within the following 6 months. Customers with any questions about this issue can contact Brainlab at 800-597-5911 or via e-mail at us.support@brainlab.com. |  | 40 | 40 systems (worldwide) | Worldwide Distribution - US Nationwide in the states of AZ, AR, CA, FL, KY, MA, NE, NY, TN, TX and the countries of Argentina, Australia, Brazil, Finland, India, Japan, Netherlands, Oman, Russia, and Taiwan | Open, Classified |  |  | 40 | Computer
Z-1709-2014 | 68266 | [6 empty] | CARINAsim MAN-1236 Rev 2 and CARINAiso MAN-1237 Rev 2 laser software. LAP isocenter marking control software. Laser systems for patient alignment. |  |  | October 07, 2013 | 2013 | LAP of America Laser Applications, LLC | Potential for patient to be marked incorrectly. Customers currently using CARINAsim or CARINAiso systems can encounter two safety relevant errors in the operating instructions (CARINAsim MAN-1236 Rev 2 / CARINAiso MAN-1237 Rev 2) | Software design | On 10/16/2013 Customers were sent an Urgent field safety notice for Corrective action concerning the operating instructions of CARINAsim and CARINAiso. The letter identified the affected product, as well as the issues involved. Users were to review their current operating steps and validate their operating sequence. Users should have received a modified software version with software update in February 2014. All users should have been notified of this letter and a copy should be maintained in the users' records. Questions should be directed to Thomas Gaudet on 561-416-9250 at the Service Center. |  | 280 | 280 Manuals | Worldwide Distribution -- USA, including the states of AK, AZ, CA, CT, DC, FL, GA, IA, ID, IL, KS, LA, MA, MD, MI, MN, MO, MS, NC, NH, NJ, NM, NY, OH, OK, OR, PA, SC, TN, TX, UT, VA, VT, WA, WY and the territory of Puerto Rico; and, the countries of Brazil, Canada, Chile, Colombia, Cuba, Honduras, and Mexico. | Open, Classified |  |  | 280 | Computer
Z-1705-2014 | 68357 | No | DSI | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Spacelabs Medical Ultraview SL Command Module, Model 91496. |  |  | May 08, 2014 | 2014 | Spacelabs Healthcare Inc | A software anomaly causes temperature alarm limits to change to limits which are not physiologically possible and result in nuisance false positive alarms. | Software Manufacturing/Software Deployment | Spacelabs Healthcare identified a software anomaly in the Ultraview SL" Command Module, Model 91496. During an internal audit, Spacelabs identified a reportable recall that was not reported. We have a product called the Spacelabs Medical Ultraview SL" Command Module, Model 91496. A software error was discovered at a customer site which was not a safety issue, but in the interest ofcustomer service it was decided to downgrade all their modules with an older software version. |  | 64 | 64 devices | US Distribution to GA. | Terminated | June 03, 2014 | 26 | 64 | Computer
Z-1676-2014 | 68091 | No | JXG | Central nervous system fluid shunt and components. | Neurology | Neurology | 510(k) | Aesculap Inc.(AIC) Miethke Ventricular Catheter with Deflector/Miethke Shunt System |  |  | April 21, 2014 | 2014 | Aesculap, Inc. | AIC (USA) received information regarding 3 complaints in which the deflector did not move freely on the ventricular catheter. | Device Design | Aesculap sent an Important Correction &amp; Removal recall letter dated April 21, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. An Aesculap Sales Representative will contact you to schedule return of any affected inventory and replace your product. Please complete the attached Inventory Sheet which is necessary to comply with FDA regulations. When completing the inventory sheet, please be sure to fill in the quantitybeing returned. If you cannot locate the product, please provide an explanation as to why the inventory will not be returned (discarded, etc.).AIC (USA) appreciates your cooperation on this matter and apologizes for the inconvenience this may cause. Thank you for your patience and continued support of this product. Please contact (610) 984-9414 with any questions. |  | 48 | 48 | US Distribution including the states of PA, KY, NY, RI, SC, KS, TX, CT, TN and WA. | Terminated | October 20, 2014 | 182 | 48 | Computer
Z-1235-2014 | 67416 | No | INI | Motorized three-wheeled vehicle. | Physical Medicine | Physical Medicine | 510(k) | Buzzaround XL scooters' |  |  | January 21, 2014 | 2014 | Golden Technologies, Inc. | Buzzaround XL scooters' front to rear lockup can become unintentionally disengaged due to possible improper fit of the front alignment cup to the rear alignment screw. | Device Design | Golden Technologies sent a recall notification letter, dated January 14, 2014 to providers/dealers. A second letter, dated January 21, 2014 was sent to these same customers. Based on a slow initial response, the firm intends to send out the correction kit with the inspection acknowledgement form as soon as possible. Based on the overall response from dealers, a third consumer letter was drafted (notification3) for mailing directly to known consumers, if necessary. For help contact a special technical service hotline at 1-800-624-6374 ext. 392. |  | 5044 | 5044 | Worldwide Distribution - USA Nationwide | Terminated | July 23, 2015 | 548 | 5044 | Computer
Z-1245-2014 | 67592 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | GE Healthcare Optima CT580/Discovery CT590RT scanners running software version 13BW16.3. |  |  | February 14, 2014 | 2014 | GE Healthcare, LLC | GE Healthcare has recently become aware of a potential safety issue due to a software issue associated with the Head Scan Protocols used on your Optima CT580 or Discovery CT590RT scanner.A potential hazardous situation may occur during a CT head scan causing a novel and rare artifact that may hide pathology. When using the head and large head SFOV, which applies the Iterative Bone Option (IBO) | Software design | The firm, GE Healthcare, sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated February 14, 2014 to its Consignees. The letter was addressed to Hospital Administrators / Risk Managers, Radiology Directors and Radiologists. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction and Contact Information. The consignees were instructed to ensure that all potential users in your facility are made aware of this safety notification and recommended actions. A GE Healthcare Service representative will perform the required software update on each affected system. This activity will be performed at no cost to consignees.Please contact your local GE Healthcare Service Representative if you have any questions or concerns regarding this notification. You may also contact GE Healthcare Call Center: United States or Canada: 800-437-1171 and in other countries, please contact your local GE Healthcare Service Representative. |  | 33 | 33 (21 US, 12 OUS) | Worldwide Distribution: US (nationwide) to states of: CO, IN, MA, NJ, NC, OH, PA, TX, and VA; and countries of: Algeria,France, Germany, Poland, Russia, and Taiwan. | Terminated | April 29, 2015 | 439 | 33 | Computer
Z-1677-2014 | 68263 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | enGen (TM) Laboratory Automation System configured with Data Innovations, LLC Instrument Manager (TM) (IM) v8.06 --- For in vitro quantitative measurement of a variety of analytes of clinical interest --- |  |  | March 20, 2014 | 2014 | Ortho-Clinical Diagnostics | Software Anomaly: Ortho Clinical Diagnostics (OCD) has become aware that if a sample requires follow-up action, such as dilution or retesting due to a mechanical issue, it is possible for that sample to be routed to the sample storage drawer (Storage) prior to the completion of the necessary pending action. A delay in patient test result(s) reporting could occur as operator is not alerted that results have not been produced and that further actions are required. | Software design | US consignees were sent Urgent Product Correction Letters (Cl2014-102, dated 3/20/2014) on 3/20/2014 by FedEx or USPS (for PO Box addresses only) overnight mail, informing them of the issue and instructing them to ensure that all required data elements are entered and in the correct format when using the auto verification feature or custom rules. J&J; Foreign affiliate consignees were notified by e-mail on 3/20/2014 informing them of the issue. |  | 81 | Domestic: 20 units, Foreign: 61 units | Worldwide Distribution-USA (nationwide) including Puerto Rico, and the countries of Australia, Bermuda, Brazil, Canada, Chile, China, Colombia, England, France, Germany, India, Italy, Japan, Mexico, Panama, Singapore, Spain, and Venezuela. | Terminated | April 11, 2017 | 1118 | 81 | Computer
Z-1678-2014 | 68179 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | IntelliSpace Portal (ISP), (ISP DX/HX/EX, Model number 881001; ISP IX, Model number 881030; ISP LX SPECT, Model number 881035), Philips Medical Systems. A medical software system for diagnostic imaging data viewing. |  |  | March 11, 2014 | 2014 | Philips Medical Systems (Cleveland) Inc | During SPECT reconstructions using Attenuation Correction and Scatter Correction, no Scatter Correction is being applied in the AutoSPECT Pro application on IntelliSpace Portal. In addition, resolution recovery is not applied correctly in SPECT reconstructions using Astonish in AutoSPECT Pro on Intellispace Portal. | Software design | On 03/10/2014 the firm sent Urgent Medical Device Correction Letters to their customers. |  | 157 | 157 units | Worldwide distribution, including US nationwide, Argentina, Australia, Austria, Canada, Cayman Islands, Chile, Denmark, Finland, France, Germany, India, Iraq, Israel, Italy, Latvia, Malaysia, Mexico, Netherlands, Panama, Peru, Poland, Saudi Arabia, Singapore, Slovakia, Spain, Switzerland, Turkey and United Kingdom. | Terminated | August 04, 2015 | 511 | 157 | Computer
Z-1263-2014 | 67685 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | Philips Model EPIQ 5 Ultrasound System. Intended for diagnostic ultrasound imaging and fluid flow analysis. |  |  | February 28, 2014 | 2014 | Philips Medical Systems North America Inc. | A problem has been detected with Philips Model EPIQ 5 and 7 Ultrasound System. The system will report incorrect Doppler velocities if you use High Pulse Repetition Frequency (HPRF) with multiple sample volume gates. While in High Pulse Repetition Frequency (HPRF) in Display Zoom Velocity Scale using manual spectral Doppler trace, the velocities in the velocity results table may be overstated by | Software design | The firm, Philips Healthcare Ultrasound, sent an "Medical Device Correction" letters dated 2014 Feb 28 via Certified Mail via the USPS by March 5, 2014. The letters described the product, problem and actions to be taken. The customers were instructed to not use HPRF in PW Doppler mode, to avoid the noted issues. Philips Field Service Engineer will be contacting them to schedule software upgrade to the systems free of charge.If you need any further information or support concerning this issue, please contact the Quality Specialist II at 978-659-4519 or your local Philips representative. |  | 262 | 489 total =227 units in US and 262 units Internationally (no Canadian consignees) | Worldwide Distribution: US (nationwide) including states of: AL, AK, AZ, CA, CO, CT, DC, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, NE, NH, NM, NJ, NM, NY, OH, OK, PA, SC, TN, TX, VA, WA, and WV; and countries of: Australia, Austria, Belgium, Canada, Chili, Ecuador, Egypt, Finland, France, Germany, Greece, India, Ireland, Italy, Japan, Korea Rep, Malaysia, Netherlands, New Zealand, Norway, Peru, Philippines, Portugal, Saudi Arabia, Singapore, South Africa, Spain, Sweden, Switzerland, Thailand, Turkey and United Kingdom. | Open, Classified |  |  | 262 | Computer
Z-1238-2014 | 67588 | No | KOD | Urological catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | SLC 2-way standard Lubricious Coated |  |  | February 07, 2014 | 2014 | Convatec Inc. | The lumen of the Medline Sterile Foley Catheter was found to be blocked by excess lubricious coating material. | Device Design | Convatec sent Urgent: Voluntary Medical Device Recall letters dated 2/7/2014 via UPS and email. Foreign customers were provided the notice via DHL. The letter identified the affected product, problem and actions to be taken. Customers are asked to perform a count of recalled product in inventory. All consignees should be contacted and be instructed to return product based on their procedures. Customers were instructed to complete and return the enclosed respoonse form. |  | 5540 | 5,540 catheters | Worldwide Distribution and in the state of TX | Terminated | December 18, 2014 | 314 | 100580 | Computer
Z-1198-2014 | 67464 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | GE Healthcare Infinia and Infinia Hawkeye Nuclear Medicine Systems. For Dual-Head Variable Angle Gamma Camera, Hawkeye Option for Dual-Head Variable Angle Gamma Camera (this includes GE Quasar Nuclear Medicine Systems). Nuclear Medicine Systems Infinia devices Product Usage: The intended use of the Quasar system is to perform general Nuclear Medicine imaging procedures for detection of radioisotope tracer uptake in the patient body, using a variety of scanning modes supported by various acquisition types and optional imaging features designed to enhance image quality in Oncology, Cardiology, Neurology and other clinical diagnostic imaging applications. |  |  | February 07, 2014 | 2014 | GE Healthcare, LLC | Potential safety issue related to unintended radial detector motion, which may occur during patient setup or at the end of a patient scan. This failure mode has been observed on a single system, in which the detector moved to the home position and then slowly starting drifting inward toward the patient. This generated both a visual and audio alert condition on both the console and the gantry. N | Device Design | GE sent an Urgent Medical Device Correction letter, dated 2/7/2014 to medical facilities beginning 2/7/2014. The letter was addressed to Manager of Nuclear Medicine, Hospital Administrators/Risk Managers, and Managers of Radiology/Cardiology. The letter described the issue and provided safety instructions. For Infinia and Infinia Hawkeye, Millennium VG, VG Hawkeye and VariCam nuclear Medicine Systems, the letter stated that a software upgrade would correct the issue. For the Helix Nuclear Medicine Systems, GE stated that a software upgrade can not be applied to the older systems and that additional information will be provided at a later date. UPDATE ON THE HELIX DEVICES: An Urgent Medical Device Correction letter, dated 4/7/2014 was sent to customers who still have a GE Helix Nuclear Medicine System. The letter requested to immediately stop using the device, and reiterated the March 2008 letter in which GE stated that they would no longer be able to offer full service support. The main reason for this was due to unavailability of parts required for maintenance and repair of the device because outside suppliers cease to produce the older parts. GE offered assistance in removal and disposal of the Helix system. |  | 3062 | 3062 units (1182 US, 1880 OUS) | Worldwide Distribution-USA (nationwide) including Puerto Rico and Guam and the countries of Algeria, Argentina, Australia, Austria, Belgium, Brazil, Canada, Chile, China, Colombia, Costa Rica, Croatia, Cyprus, Czech Republic, Denmark, Dominican Republic, Ecuador, Egypt, Finland, France, Georgia, Germany, Greece, Hong Kong, Hungary, India, Indonesia, Iran, Ireland, Israel, Italy, Japan, Jordan, Korea (Republic of), Kuwait, Latvia, Lebanon, Lithuania, Malaysia, Malta, Martinique, Mexico, Morocco, Namibia, New Zealand, Norway, Pakistan, Panama, Philippines, Poland, Portugal, Russia, Saudi Arabia, Singapore, Slovakia, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, Tunisia, Turkey, Ukraine, and United Kingdom. | Terminated | November 03, 2014 | 269 | 3062 | Computer
Z-1781-2014 | 68111 | No | HTZ | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Synthes Matrix Mandible Short Cut Plate Cutter |  |  | April 22, 2014 | 2014 | Synthes, Inc. | It was discovered internally that the face of the Synthes Matrix Mandible Short Cut Plate Cutter has the potential for discoloration/corroded material in the affected lot. | Device Design | Synthes Inc. sent an Urgent Medical Device Recall Notice dated April 22, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.If you DO have any of the identified devices, please take the following steps:Contact DePuy Synthes at 1-800-479-6329 to obtain a Return Authorization (RA) Number . Complete the Verification Section at the end of this letter by checking the appropriate box indicating affected product has been located. Also, please indicate the number of devices found. Please include your name, title, telephone number and signature in the spaces provided.Return the Verification Section (page 3 of this letter) with the product to: Credit/Returns, DePuy Synthes, 1101 Synthes Avenue, Monument, CO 80132. Return a copy of the Verification Section (page 3 of this letter) to DePuy Synthes by: Fax: 866-324-3731 or Scan/email: Synthes7726@stericycle.comIf you DO NOT have the identified product, please take the following steps:Complete the attached Verification Section at the end of this letter by checking the appropriate box indication that no affecteded product has been located. Please include your name, title, telephone number and signature in the spaces provided this return documentation acknowledges your receipt of medical device removal information.Return the Verification Section (page 3 of this letter) to DePuy Synthes by:Fax: 866-324-3731 or Scan/email: Synthes7726@stericycle.comNote: If the Verification Section is answered on behalf of more than one facility and/or individual, please clearly indicate the name and address of the facility and/or individual on page 3 of the notification.If you have any questions, please call 610-719-5450 or contact your DePuy Synthes sales consultant. |  | 14 | 14 | US Distribution including the states of CA, CT, FL, GA, NY, SD and TX. | Terminated | September 10, 2015 | 506 | 14 | Computer
Z-1147-2014 | 67430 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | GE Healthcare Solar 8000M and Solar 8000i Patient Monitor software version 5.4, 5.5, or 5.6 with the Automatic-View-On- Alarm (AVOA) feature enabled. |  |  | February 13, 2014 | 2014 | GE Healthcare, LLC | GE Healthcare has recently become aware of a potential safety issue with the Solar 8000M/i Patient Monitor Automatic-View-On-Alarm (AVOA) feature.The Solar 8000M/i AVOA feature will stop functioning if the patient monitor is discharged from the CIC Central Station while a remote automatic bedside alarm view is currently being displayed on the Solar. Once this issue occurs, subsequent remote AV | Software design | GE Healthcare sent a "Urgent Medical Device Correction" letter dated February 12, 2014. The letter was sent to all affected customers. It identified the product the problem and the action needed to be taken by the customers. This issue affects all Solar 8000M/i Patient Monitors running software version 5.4, 5.5, or 5.6 with the Automatic-View-On-Alarm (AVOA) feature enabled.GE Healthcare will provide a correction at no charge once it is available. We will contact you to arrange for this correction.Please contact Technical Support at 1-800-558-7044 or your local Service Representative. |  | 10211 | 10,211 (7,549 units US, 2,662 units OUS) | Worldwide Distribution - USA (nationwide) including the states of AL, AZ, AR, CA, CO, CT, DE, DC, FL, GA, ID, IL, IN, KS, KY, LA, ME, MD, MA, MI, MN, NS, NO, NE, NH, NH, NM, NY, NC, OH, OK, OR, PA, SC, TN, TX, UT, VA, WA, WV, WI and WY., and the countries of : AUSTRALIA, CANADA, AUSTRIA, BANGLADESH, BELGIUM, BOSNIA AND HERZEGOVINA, BRAZIL, CHILE, CHINA, COSTA RICA, CYPRUS, EGYPT, FINLAND, FRANCE, GERMANY, GREECE, INDIA, INDONESIA, JAPAN, MALAYSIA, MEXICO, NAMBIA, NETHERLANDS, NORWAY, POLAND, RUSSIA, SAUDI ARABIA, SINGAPORE, SOUTH KOREA, SPAIN, SWEDEN, TAIWAN, THAILAND, TURKEY, UNITED KINGDOM, SERBIA AND YUGOSLAVIA. | Terminated | July 10, 2015 | 512 | 10211 | Computer
Z-1839-2014 | 67610 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Altrus Thermal Tissue Fusion Handpieces under the following specifications: 1) 10mm x 16 cm Catalog Number: 60-9510-001, 2) 10 mm x 23 cm Catalog Number: 60-9511-001, 3) 10 mm x 36 cm Catalog Number: 60-9512-001, 4) 5 mm x 16 cm Catalog Number 60-9520-001, 5) 5 mm x 23 cm Catalog Number 60-9521-001 and 6) 5 mm x 36 cm Catalog Number 60-9522-001. Products are packaged within cartons, 6 units per carton. |  |  | February 20, 2014 | 2014 | ConMed Corporation | ConMed Corporation initiated an Urgent Medical Device Correction for Altrus Thermal Tissue Fusion Handpieces due to risk of burns. | Device Design | On February 20, 2014 and February 28, 2014 ConMed Corporation distributed Urgent Medical Device Correction notices to their domestic and foreign customers via FedEx Priority. The notice was generated by ConMed for information purposes only. Customers are to ensure that the jaws of the energy vessel sealing device does not come in contact with unintended tissue, surfaces, or objects until the jaws have cooled sufficiently. Customers who have further distributed this product, please forward a copy of the correction notice to the recipient as appropriate. Customers will questions should contact ConMed Regulatory Affairs at (315) 624-3237 or by email at altrus@conmed.com. |  | 12653 | 12,653 units (10,769 domestically &amp; 1,884 internationally) | Worldwide Distribution. | Terminated | April 07, 2017 | 1142 | 12653 | Computer
Z-1812-2014 | 68419 | No | DQX | Catheter guide wire. | Cardiovascular | Cardiovascular | 510(k) | HydroFinity Hydrophilic guidewire; Nitinol Guidewire with Hydrophilic Coating, Sterile. |  |  | May 28, 2014 | 2014 | Nitinol Devices and Components, Inc. | Reports of outer polymer jacket to the core wire being damaged during use on the HydroFinity Guidewire. | Device Design | Covidien sent an Urgent Product Recall letter dated May 30, 2014, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to return any unused devices to Covidien, their account would be credited. Customers were also instructed to::1) Stop using product listed in this letter immediately2) Segregate this product from other inventory3) Fill out the reply (verification) form at the end of this letterÂż If you do not have any product identified in this letter, please fax oremail the completed form to Covidien at (877) 523-9109 orcovidien8770@stericycle.com.Âż If you do have product, your sales representative will assist you incompleting the verification form and arranging for return of theproduct.Âż Please fax the completed form to Covidien (877) 523-9109 or email tocovidien8770@stericycle.com.4) Your sales representative will be available to answer any questions regarding this recall and assist you in completing the verification form, returning product and addressing any account credits.This action is being conducted with the knowledge of the United States FDA and other regulatory authorities. Âż Customers with questions were instructed to call (202) 310-5120.For questions regarding this recall call 202-310-5120. _ |  | 8485 | 8485 in US, 9595 - ROW - total, all models | Worldwide Distribution - USA (nationwide) and Internationally to Italy, France, Norway, Sweden, Germany, Belgium, United Kingdom and Canada. | Terminated | January 13, 2015 | 230 | 8485 | Computer
Z-1082-2014 | 67223 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Desktop Pro TM linear accelerator control software intended to assist a licensed practitioner in the delivery of radiation to defined target volumes. |  |  | November 25, 2013 | 2013 | Elekta, Inc. | During the treatment table, gantry, and collimator calibration procedures, it is possible to type incorrect values or invert a value. | Software design | Corrective Action #1 - Release of Important Field Safety Notice 200 01 502 040 - "Recommended quality assurance after a calibration procedure" Released 25 November, 2013. All software versions (r7.01) detailed in this notice are affected by this recall. 200 01 501 041 - Recommended quality assurance after a calibration procedure Released November 25, 2013. Only software versions R6.0 and R6.1 are affected by this recall. Important Field Safety Notices are in the process of being distributed to customers advising them of the recommended Quality Assurance checks to avoid/detect the problem. The permanent solution to this problem is currently been developed. Corrective Action #2 - Permanent Solution - All users of Desktop Pro software will be informed of a mandatory upgrade to Integrity 1.2 upon its release which will prevent calibration from being inverted. |  | 1309 | 1309 | Worldwide Distribution: US (nationwide): AL, AK, AZ, AR, CA, CO, CT, FL, GA, ID, IL, IA, KY, LA, MD, MA, MI, MN, MS, MO, MT, NE, NV, NH, NJ, NM, NY, NC, ND, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VT, VA, WA, WV, WI, including Puerto Rico, DC; and countries of: Albania, Argentina, Australia, Austria, Bangladesh, Belarus, Belgium, Bosnia &amp; Herzego, Brazil, Bulgaria, Canada, Chile, China, Colombia, Croatia, Cyprus, Czech Republic, Denmark, Ecuador, Egypt, Estonia, Finland, France, Gabon, Germany, Ghana, Greece, Hong Kong, Hungary, India, Ireland, Israel, Italy, Japan, Jordan, Libya, Lithuania, Malaysia, Malta, Mexico, Morocco, Netherlands, New Zealand, Norway, Pakistan, Panama, Peru, Philippines, Poland, Portugal, Romania, Russia, Saudi Arabia, Serbia, Singapore Slovakia, South Africa, South Korea, Spain, Sudan, Suriname, Sweden, Switzerland, Syria, Taiwan, Thailand, Turkey, United Kingdom, Ukraine, Venezuela, and Vietnam. | Open, Classified |  |  | 1309 | Computer
Z-1869-2014 | 68386 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | MOSAIQ |  |  | May 29, 2014 | 2014 | Elekta, Inc. | MOSAIQ does not re-calculate the agent volume when the user changes the ordering dose (to other than what was calculated) prior to saving the order. | Software design | Elekta issued an Important Safety Notice dated May 29, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Corrective action that shall be taken:Notify the affected customers of the issue - To be completed 06/22/2014Follow up with the affected customers for:Confirmation receipt of the LIN, ANDConfirmation of installation of patch for the fix- To be completed 11/30/2014The notice includes a confirmation which is to be returned.For further questions call (770) 670-2422. |  | 101 | 101 | Worldwide Distribution - USA (nationwide) including Puerto Rico and the states of AL, AK, AZ, CA, CO, CT, DE, FL, GA, HI, IL, IN, IA, KY, LA, ME, MD, MA, MI, MO, MT, NH, NJ, NM, NY, NC, ND, OH, OK, OR, PA, SC, SD, TX, VT, VA and WA,. and the countries of Australia, Canada, New Zealand and United Kingdom. | Open, Classified |  |  | 101 | Computer
Z-1867-2014 | 63486 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Siemens StreamLAB(R) Analytical Workcell; Catalog number 746002.901. |  |  | October 11, 2012 | 2012 | Siemens Healthcare Diagnostics, Inc. | Firm has received reports from customers that the small pins at the bottom of the center door panel of the Input/Output Module, which protrude approximately 5/8 inch (16 mm), have contributed to trip and fall incidents. Customers are advised that when working in the area of the Input/Output Module to be aware of the two pins that protrude and use caution. | Device Design | The firm initiated their recall/correction of this product on October 11, 2012 by sending an Urgent Customer Notification letter to all StreamLAB(R) Automation Solutions customers. The letter informed customers that the firm received reports from customers that the small pins at the bottom of the center door panel of the Input/Output Module have contributed to two trip and fall incidents. Customers were advised that when working in the area of the Input/Output Module to be aware of the two pins that protrude and use caution. Also, StreamLAB operators were to be notified of the potential safety issue and the notice was to be posted on or near the system. Additionally, customers were to complete and return the attached Field Correction Effectiveness Check form. |  | 317 | 317 | Nationwide Distribution -- AL, AR, AZ, CA, CO, CT, FL, GA, ID, IL, IN, KS, LA, MA, MD, MI, MN, MO, MS, MT, NC, NE, NH, NJ, NM, NY, OH, OK, PA, SC, SD, TN, TX, UT, VA, WA, WI, and WV. | Terminated | August 26, 2014 | 684 | 317 | Computer
Z-1866-2014 | 68259 | [6 empty] | LAP LASER PICTOR 3D Multi Color Laser System |  |  | March 21, 2014 | 2014 | LAP of America Laser Applications, LLC | While projecting yellow lines it was notice that the yellow projection has the potential to splits into two separate red and green projections. In this case the difference between the red and green projection was up to 2.5mm. | Software design | LAP of America Laser sent an Urgent safety Information FFST 2014-2-3 via UPS Next Day Air to the customer. |  | 1 | 1 system | North Carolina | Open, Classified |  |  | 1 | Computer
Z-1065-2014 | 65451 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | All Elekta digital accelerators with electrons and beam modulator, MLCi/MLCi2, Agility or Asymmetric Heads |  |  | May 28, 2013 | 2013 | Elekta, Inc. | Customers are resetting the default values outside recommended factory settings. | Device Design | Elekta sent an Important Field Safety Notice dated May 2013 to all affected customers. The Notice identified the affected product, problem, clinical impact and actions to be taken. Customers were instructed to adhere to the instructions and advice given in manuals and system dialogues. For question contact your local Elekta office. |  | 472 | 472 units | USA Nationwide Distribution in the states of: AL, AK, AZ, AR, CA, CO, CT, FL, GA, ID, IL, IN, IA, KS, KY, LA, MD, MA, MI, MN, MS, MO, MT, NE, NV, NH, NJ, NY, NC, ND, OH, OK, OR, PA, RI, SC, TN, TX, UT, VT, VA, WA, WV, WI, WY, DC and including Puerto Rico. | Open, Classified |  |  | 472 | Computer
Z-1850-2014 | 68128 | No | FSY | Surgical lamp. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Vertical Spring Arm for Single or Dual Monitor Mount (Harmony LL 500 and Harmony LL 700 surgical lighting systems), Steris Corporation. Model numbers YG19051 and YG19262. |  |  | April 04, 2014 | 2014 | Steris Corporation | Potential for the welded seam in affected monitor spring arms to crack and/or break. | Device Design | On April 11, 2014 the firm sent Urgent Voluntary Field Correction Notices to their customers. |  | 68 | 68 units | Distribution to US (states: MA, NY, OH, PA &amp; TX) and Canada. | Terminated | October 15, 2014 | 194 | 68 | Computer
Z-1849-2014 | 68466 | No | CAN | Pressure regulator. | Anesthesiology | Anesthesiology | 510(K) Exempt | Model RG3000 and RG4000 Series Valve Integrated Pressure Regulator (VIPR) |  |  | April 07, 2014 | 2014 | Western / Scott Fetzer Company | Separation events have occurred, including the VIPR separating from the Aluminum cylinder. | Device Design | All initial consignees of subject product were notified by letter dated June 5, 2014 via U.S. mail. Initial consignees were directed to forward copies of the recall notification letters to their contacts in the event the subject product has been further distributed. |  | 118325 | 118,325 units | Worldwide Distribution-USA (nationwide) including Puerto Rico and the states of Alaska, Arkansas, Arizona, California, Colorado, Connecticut, Delaware, Florida, Georgia, Hawaii, Idaho, Illinois, Indiana, Kansas, Kentucky, Louisiana, Massachusetts, Maryland, Michigan, Minnesota, Missouri, Mississippi, Montana, North Carolina, North Dakota, Nebraska, New Hampshire, New Jersey, New Mexico, Nevada, New York, Ohio, Oklahoma, Oregon, Pennsylvania, Rhode Island, South Carolina, South Dakota, Tennessee, Texas, Utah, Virginia, Washington, Wisconsin, and West Virginia and the countries of Canada,China, Costa Rica, France, Honduras, India, Ireland, Italy, Japan, Korea (South), Malaysia, Panama, Spain and Taiwan. | Open, Classified |  |  | 118325 | Computer
Z-1848-2014 | 68456 | No | FCN | Urine collector and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | SteriGear Urinary Drain Bag with Fig Leaf Cover, 2000 ml, REF 10270, The Fig Leaf, SteriGear 362 S. University Avenue, provo, UT 84601, 1-800-398-3259, SteriGear.com, Sterile, EO |  |  | January 11, 2011 | 2011 | SteriGear, LLC | SteriGear notifed customers of circumstances in which The Fig Leaf urinary drain bag valve can close creating a negative pressure that can suck the non-return valve closed. | Device Design | SteriGear, LLC sent a special note to all affected customers between January 12, 2011 and September 28, 2012. The note identified the product, the problem, and the action to be taken by the customer. Customers were instructed on how to reposition the inlet tube. Customers with questions were instructed to contact Director of Product Development at 801-607-3211. No correction or removal was performed.For questions regarding this recall call 801-607-3209. |  | 110560 | 110,560 | Nationwide Distribution | Terminated | January 02, 2015 | 1452 | 110560 | Computer
Z-1847-2014 | 68256 | No | NEP | Intravascular administration set. | General Hospital | General Hospital | 510(K) Exempt | Abacus 3.1, Pharmacy Calculator |  |  | May 28, 2014 | 2014 | Baxter Corporation Englewood | Baxter Corporation is initiating a field correction following an upgrade to ABACUS Software v3.1, when Clinimix 4.25110 is used in an order, the Nutritional Summary and any labels containing energy information may display 1,000 times less protein than is actually present in the solution. | Software design | Baxter sent an Urgent Device Correction letter dated May 27, 2014 via certified mail to affected customers. The letter identified the affected product, problem and actios to be taken. Customers were instructed to: 1. Review and distribute this communication to pharmacy staff and compounding staff that use ABACUS software. This notice is intended only for customers that have upgraded to version 3.1 of the software.2. Please acknowledge receipt of this notification by completing the enclosed customer reply form and return it to Baxter by either faxing it to 1-224-270-5457 or scanning and e-mailing it to fca@baxter.com.3. If you have distributed ABACUS to other facilities, please notify the end users in accordance with your customary procedures. For questions call Baxter Techenical Support line at 1-800-678-2292. |  | 42 | 42 | US Nationwide Distribution and the state of VA. | Open, Classified |  |  | 42 | Computer
Z-1845-2014 | 66315 | No | KPR | Stationary x-ray system. | Radiology | Radiology | 510(k) | Siemens RAD Fluoro Uro System (Ysio, Luminos dRF, and Uroskop Omnia) with software versions VB10G. |  |  | August 28, 2013 | 2013 | Siemens Medical Solutions USA, Inc | When using systems operating with software versions VD10A/G during a RAD examination, a malfunction can result in the image appearing very dark, which may prompt unnecessary repetition of the examination. | Software design | Siemens sent a notification letter, dated August 28, 2013, to Affected customers. The letter identified the affected product, problem and actions to be taken. The letter provided work-around instructions and announced that a field modification was prepared to resolve the issue. |  | 76 | 76 | USA Nationwide Distribution in the states of OH, NJ, MA, CO, NY, CA, MD, OR, KY, PA, IL, VA, TX, MN, KS, GA, and WY. | Terminated | February 24, 2015 | 545 | 76 | Computer
Z-1156-2014 | 67495 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Panaroma Central Station |  |  | January 08, 2014 | 2014 | Mindray DS USA, Inc. dba Mindray North America | Mindray has identified two issues with the Panorama Central Station. Full Disclosure data may lag behind the Panorama system time clock when the Panorama CentralStation is in use with MindrayÂżs V-Series Monitor. Additionally, event indicators lag behind the Panorama system time clock when the Panorama Central Station is in use with a Panorama Telepack. | Software design | Mindray DS USA Inc. sent an Urgent Corrective Action letter dated January 8, 2014 via certified mail to their customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to contact Mindray Service Representative, at 1-800-288-2121 (Monday  Friday 8:30 am-5:30 pm), to arrange for the software update of the Panorama System(s) in their facility. This work will be performed at no cost. |  | 732 | 732 units in the US | Worldwide Distribution - USA Nationwide Distribution | Terminated | January 21, 2015 | 378 | 732 | Computer
Z-1154-2014 | 67470 | No | HKI | Ophthalmic camera. | Ophthalmic | Ophthalmic | 510(k) | RetCam 3, RetCam Portable and RetCam Shuttle; |  |  | January 24, 2014 | 2014 | Clarity Medical Systems Inc | Software anomaly for RetCam 3, RetCam Shuttle and RetCam Portable with software versions 6.0.x, 6.1. x and 6.2x where the patients age is calculated incorrectly. | Software Manufacturing/Software Deployment | Urgent Safety Alert letters were sent to Risk Manager and to NICU US consignees via FedEx on 1/24/2014.Template of letters to OUS distributors to be sent on distributors letterhead and sent to device recipients via electronic communication and./or hard copy on 1/24/2014. |  | 480 | 166 Ret Cam Shuttles, 288 RetCam 3 and 23 RetCam Portables. | Worldwide Distribution-USA (nationwide) including DC and the states of NM, LA, CA, MI, PR, GA, MD, MT, AL, MA, NY, IL, OH, TX, VA, MI, FL, MO, KY, TN, CT, NJ, PA, UT, SD, OK, MS, and NC and the countries of Algeria, Australia, Austria, Bangladesh, Belarus, The Netherlands, Brazil, Bulgaria, Chile, China, Columbia, Czech Republic, Denmark, Ecuador, Egypt, Finland, France,India, Indonesia, Iran, United Kingdom, Israel, Italy, Kuwait, Lebanon, Malaysia, Mexico, Sweden, Pakistan, Peru, Philippines, Poland, Portugal, Saudi Arabia, Taiwan, Thailand, Turkey, UAE, Argentina, Uzbekistan, Oman, Canada, Germany, and Switzerland. | Terminated | August 11, 2014 | 199 | 480 | Computer
Z-1163-2014 | 67257 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | GE Healthcare, Innova 2100IQ, Innova 3100, Innova 3100IQ, Innova 4100, Innova 4100IQ Innova 2121IQ, Innova 3131IQ; Cardiovascular X-ray imaging systems (X-ray generator installed with Rotor V3 version and Heater V4 version boards). |  |  | December 11, 2013 | 2013 | GE Healthcare, LLC | GE Healthcare has become aware of a potential safety issue involving the innova systems. The Innova system can unexpectedly stop delivering X-rays after a power-on or a reset cycle which may result in the loss of real-time interventional imaging. No injuries have been reported due to this issue. | Software design | GE Healthcare sent an "Urgent Medical Device Correction " letter dated December 10, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. The letter described the Safety Issue, The long term solution is to provide a software correction for all of the affected systems that were distributed. The plan for public warning is to contact all customers with an affected system by means of an Urgent Medical Device Correction letter within the US. In the US this will be sent by a traceable means (either US Postal Service with receipt acknowledgement (registered mail) or equivalent courier service such as FEDEX). This letter will be sent to the following titles within the affected accounts:GE Healthcare will correct all affected systems at no cost to you. A GE Healthcare service representative will contact you to arrange for this correction. If you have any questions or concerns regarding this notification, please call one of the following phone numbers: United States: 800 437 1171For other countries, please contact your local GE Healthcare Service Representative. |  | 98 | 257(159 US, 98 OUS) | Worldwide Distribution - USA (nationwide) including the states of AL, AK, AZ, AR, CA, CO, CT, FL, GA, IL, IN, KS, KY, LA, ME, MD, MA, MI, MS, MO, NV, NJ, NY, NC, ND, OH, OK OR, PA, PR, RI, SC, TN, TX, UT, VA,WA,WY and WI., and the countries of ARGENTINA, AUSTRALIA, CANADA, CHINA, CZECH REPUBLIC, DENMARK, EGYPT, FINLAND, FRANCE, GERMANY, HUNGARY, INDIA, IRELAND, ITALY, JAPAN, JORDAN, REPUBLIC OF KOREA, MALAYSIA, MEXICO, NEW ZEALAND, NORWAY, PHILIPPINES, POLAND, RUSSIA, SAUDI ARABIA, SPAIN, SRI LANKA, TURKEY and UNITED KINGDOM. | Terminated | July 29, 2014 | 230 | 98 | Computer
Z-1177-2014 | 67463 | No | LIB | Microorganism differentiation and identification device. | Microbiology | Microbiology | 510(K) Exempt | bioMerieux PREVI Isola System (General purpose,microbiology, diagnostic Medical Device), (REF 29500/29500R), bioMerieux, Hazelwood, MO. |  |  | February 10, 2014 | 2014 | Biomerieux Inc | The firm has determined the product may fail to dispense the sample to the agar plate resulting in a "failure to dispense" in conjunction with urine specimens. This could lead to a "false" negative growth result. | Software design | The firm sent a recall notification letter to all of their consignees on 2/10/2014. |  | 471 | 471 instruments | Distribution was made to AK, CA, CO, CT, DC, FL, GA, IL, IN, KY, LA, MA, MO, NC, NJ, NY, OH, OR, PA, RI, SC, TN, TX, VA and WI. | Terminated | March 16, 2017 | 1130 | 471 | Computer
Z-1832-2014 | 68320 | equipment, laboratory, general purpose, labeled or promoted for a specific medical use | LXG | General purpose laboratory equipment labeled or promoted for a specific medical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | VersaCell X3 Solution and VersaCell X3 Expansion Pack, a robotic laboratory sample-management system. It connects up to 3 instruments and provides a user interface that routes samples, reduces manual tasks and consolidates results. |  |  | April 25, 2014 | 2014 | Siemens Medical Solutions Diagnostics | The keyboard and monitor may not be properly secured to the monitor arm. | Device Design | Siemen's Healthcare Diagnostics sent an Urgent Medical Device Correction letter dated April 25, 2014, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to complete and return teh Field Correction Effectiveness Check form attached to the letter within 30 days by fax to 302-631-7597.Inspect the monitor, keyboard and bracket assembly. If the screw and washer assembly is secure as shown in Figure 1, the monitor and keyboard assembly is safe to use. After making adjustments or repositioning the monitor, keyboard and bracket assembly, inspect the screw adn washer to ensure it is ssecure as shown in Figure 1.If there is a gap between the underside of the monitor support bracket and cylindricalmounting feature as shown in Figure 2, contact your Siemens technical supportrepresentative to schedule a service visit as soon as possible. Front load and manage samples directly on the connected instruments until your Siemens technical support representative addresses the issue.Customers were instructed to retain the letter with their laboratory records, and forward the letter to those who may have received the affected product. Customers with questions were instructed to contact their Siemens Customer Care Center or their local Siemens technical support representative.For questions regarding this recall call 877-229-3711. |  | 19 | 19 units | Worldwide Distribution - USA including Iowa and New York and Internationally to Canada, Chile, France, India, Italy, Malaysia, Peru, Korea, and the United Kingdom. | Terminated | April 16, 2015 | 356 | 19 | Computer
Z-1165-2014 | 67515 | Yes | DTR | Cardiopulmonary bypass heat exchanger. | Cardiovascular | Cardiovascular | 510(k) | MAQUET PLEGIOX Cardioplegia Heat Exchanger |  |  | January 31, 2014 | 2014 | Maquet Cardiovascular Us Sales, Llc | The MAQUET post-marketing vigilance program has identified a discrepancy that thedesired cardioplegia temperature may not be obtained when the PLEGIOX CardioplegiaHeat Exchanger (PLEGIOX) is used in conjunction with a heater/cooler unit. | Device Design | Maquet Inc.sent an" Field Safety Notice" dated January 31, 2014, to all affected customers via Fed Ex. The letter identified the product the problem and the action needed to be taken by the customer.Any questions or additional information, can be obtained from the MAQUET representative, or MAQUET Customer Service at 1-888-627-8383 (press option 2, followed by option 3), Monday through Friday, between the hours of 8:00 a.m. and 6:00 p.m. EST.Please complete and return the enclosed confirmation form, acknowledging your receipt and understanding of this communication, by emailing a scanned copy to Fieldactions@maquet.com or via fax at (973) 807-9290. |  | 1424 | 1,424 units | Worldwide Distribution - USA (nationwide) and the countries of Brazil, Germany and Turkey | Terminated | April 24, 2015 | 448 | 1424 | Computer
Z-1827-2014 | 68311 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | Philips MR systems: Trolley Variable Height IRF (for Ingenia 1.5T &amp; 3.0T products) |  |  | May 15, 2014 | 2014 | Philips Medical Systems, Inc. | Ingenia customers have experienced clamping of the foot under the central column of the Height Adjustable Flex Trak trolley and the Trolley Variable Height IRF. | Device Design | Philips sent an Urgent Field Safety Notice (FSN78100414) dated May 2014 via certified mail to US and Canadian consignees. Philips Key Market representatives will distribute the letter internationally. The letter identified the affected product, explains the problem, actions that should be taken in order to prevent risks for users, and the actions planned by Philips to correct the problem. On May 31, 2014, Philips released a Field Change Order that requires service engineers to replace the front and rear covers of the Ingenia MR FlexTrak Trolley with new covers with protection bumper that prevents foot clamping. For further information and support concerning this issue, consignees should contact their local Philips representative at 1-800-722-9377. |  | 5 | 5 | Worldwide Distribution - US (nationwide), Canada, Mexico Australia, Austria, Germany, Belarus, Bulgaria, Brazil, China, Denmark, Estonia, Finland, France, Germany, Greece, Italy, Jordan, Korea, Republic of Korea, Netherlands, New Zealand, Norway, Oman, Pakistan, Russian Federation, Singapore, Slovakia, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, United Arab Emirates, United Kingdom, and Vietnam. | Terminated | August 07, 2015 | 449 | 425 | Computer
Z-1825-2014 | 68535 | No | ISW | External assembled lower limb prosthesis. | Physical Medicine | Physical Medicine | 510(K) Exempt | RUSH81 Prosthetic Foot |  |  | June 02, 2014 | 2014 | Ability Dynamics LLC | Ability Dynamics is recalling RUSH81 Prosthetic Foot because the existing bolts may be bottomed out in the threaded hole and this could lead to bolt failure. | Device Design | Ability Dynamics sent an Urgent Medical Device Correction Notice letter dated June 2, 2014 to all affected customers (Prosthetist) who purchased the RUSH Prosthetic Foot. The letter informs the customers of the problems identified and the actions to be taken. Customers are instructed to fill out the Bracket Bolt Replacement Form and return to Ability Dynamics. Customers with questions are instructed to contact Rich Sainz, National Clinical Manager, 2082 E. University Dr. Tempe AZ 85281 (480) 717-1412, rsainz@abilitydynamics.com. |  | 175 | 175 units | Worldwide Distribution - US (nationwide) and in the countries of Australia and Canada. | Terminated | September 22, 2014 | 112 | 175 | Computer
Z-1173-2014 | 67528 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Device Name: ABX PENTRA ML |  |  | February 14, 2014 | 2014 | Horiba Instruments, Inc dba Horiba Medical | HORIBA Medical initiated this recall of all software versions of ABX PENTRA ML (Model A11A06152) due to a software bug with the potential of transmitting incorrect results to the Laboratory Information System. | Software design | Horiba sent an "URGENT SAFETY NOTICE letter datedFebruary 13, 2014 via by mail to customers. The letter identified the affected product, issue, impact, actions to be taken. Customers were instructed to share this information with laboratory staff and retain this notification as part of their laboratorydocumentation. Customers were requested to complete and return the enclosed response form via one of the following methods:Fax: 949-453-0600Email: jim.knowles@horiba.comMail:HORIBA MedicalAttention: Jim Knowles, Recall Coordinator34 Bunsen DriveIrvine, CA 92618.For questions contact your local Horiba Medical representative. |  | 15 | 15 | USA Nationwide Distribution | Terminated | August 22, 2014 | 189 | 15 | Computer
Z-1807-2014 | 68430 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | iGUIDE System |  |  | May 29, 2014 | 2014 | Elekta, Inc. | A bug in the software prevents the iGUIDE software from logging off. | Software design | Elekta sent an Important Field Safety Notice 618 01 303 024 dated May 29, 2014 to affected customers. The letter identified the affected product, problem, potential clinical impact and actions to be taken. It also gives a recommendation how the risk can be mitigated - Users should only use the iGUIDE software when the key is in the on position and the Position Guard is active. Users are advised to restart the workstation if any error occurs to ensure the software is running properly. The notice also advises the customer that a patch will be released that corrects the behavior of the iGUIDE software. This solution will be implemented on all affected products when available. For questions contact your local Elekta office. |  | 224 | 224 | Worldwide Distribution - USA (nationwide) in the states of AZ, CA, CT, HI, IL, IA, KY, MD, MA, MI, MO, NH, NJ, NC, ND, OH, OK, OR, PA, TX, VA, WA, WI, and the countries of Algeria, Australia, Austria, Brazil, Canada, China, Colombia, Denmark, Finland, France, Germany, Hong Kong, India, Italy, Japan, Kazakhstan, Mexico, Netherlands, New Zealand, Peru, Poland, Russia, Saudi Arabia, South Korea, Spain, Sweden, Switzerland, Taiwan, Thailand, United Kingdom and Vietnam | Open, Classified |  |  | 224 | Computer
Z-1175-2014 | 67157 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | GE Healthcare CARESCAPE Monitor B650 and CARESCAPE Monitor B850. |  |  | February 28, 2014 | 2014 | GE Healthcare, LLC | GE Healthcare has recently become aware of potential safety issues associated with the CARESCAPE Monitor B850 and CARESCAPE Monitor B650. These issues include both NIBP (Non-Invasive Blood Pressure) and Bed to Bed Issues. | Software design | GE Healthcare sent an Urgent Field Safety Notice dated February 26, 2014, to all International consignees. The letter was addressed to Healthcare Administrator / Risk Manager, Chief of Nursing and Director of Biomedical Engineering. The letter described the Safety Issues and Instructions, Affected Product Details Product Correction and Contact Information ( Contact Technical Support at 1-800-558-7044 or their local Service Representative).GE Healthcare determined that there were no US sites affected since all units were under GE control therefore, no customer letter was deployed in US.For questions regarding this recall call 847-277-4472. |  | 676 | 676 units OUS | Worldwide Distribution - USA (nationwide) and Internationally to AUSTRIA, BELGIUM, CZECH REPUBLIC, FINLAND, FRANCE GERMANY, GREECE, ITALY, LITHUANIA, POLAND, SPAIN, SWEDEN, SWITZERLAND, UNITED KINGDOM. | Terminated | April 29, 2015 | 425 | 676 | Computer
Z-1806-2014 | 68162 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | RayStation software version 4.0.3.4 |  |  | April 25, 2014 | 2014 | RAYSEARCH LABORATORIES AB | This notice concerns a problem with the computation of ROI voxel volumes for cases where the CT has variable slice spacing. The dose grid volumetric representation of an ROI can be wrong. This affects all dose-volume properties for such ROIs, including DVH, dose statistics, clinical goals and constraints or objective functions. Furthermore, if material override ROIs are used, or if the external is not represented by contours, the bug can trigger an error in the 3D dose distribution. | Software design | On April 25, 2014, RaySearch emailed all consignees of the affected product a copy of the recall notice as well as the response forms. Back up notification methods include phone calls and site visits. |  | 6125 | 25 sites/245 units (US), 25 sites/124 units (foreign) | Worldwide Distribution-USA (nationwide) including the states of AZ, CA, FL, IL, MA, MD, MO, MS, NJ, NV, NY, OH, PA, TN, VA, and WA, and the countries of Australia, Belgium, Germany, France, Italy, New Zealand, Republic of Korea, Switzerland, Spain, and United Kingdom. | Terminated | June 26, 2017 | 1158 | 6125 | Computer
Z-1194-2014 | 67459 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | Philips HeartStart MRx Monitor/Defibrillator |  |  | February 18, 2014 | 2014 | Philips Medical Systems, Inc. | When the HeartStart MRx.is used with the Q-CPR Meter in defibrillation mode, the Q-CPR Meter may incorrectly display the Do Not Touch the Patient icon. | Software Design Change | The firm, Philips Healthcare, sent an "URGENT -Medical Device Correction/Field Safety Notice" letter dated February 2014 on February 18, 2014 to its domestic and foreign customers. The letter describes the product, problem and actions to be taken. The customers were instructed to follow the Action to be Taken by Customer/User section of the attached Field Safety Notice; to continue to use your HeartStart MRx without the Q-CPR Meter attached. If you choose to use the HeartStart MRx with the Q-CPR Meter attached and you observe that the Q CPR Meter is displaying the Do Not Touch the Patient icon at an unexpected time, follow the MRx Instructions for Use and remove the meter from the patient and continue CPR according to your organization's protocol.Philips plans to take to remedy the issue. Philips will provide a software upgrade free of charge.If you have questions regarding this notification or need any further information or support, please contact your local Philips representative or call at 1-800-722-9377. |  | 6583 | 6583 units | Worldwide distribution: US (nationwide) and to countries of: ARUBA AUSTRALIA AUSTRIA BELGIUM BRAZIL CANADA ECUADOR EGYPT FINLAND FRANCE GERMANY HONG KONG INDIA INDONESIA IRELAND ISRAEL ITALY KENYA KOREA, REPUBLIC OF KUWAIT LATVIA LIBYAN ARAB JAMAHIRIYA MALAYSIA MEXICO MOROCCO NETHERLANDS NORWAY OMAN PAKISTAN PHILIPPINES POLAND QATAR RUSSIAN FEDERATION SAUDI ARABIA SINGAPORE SOUTH AFRICA SPAIN SWEDEN SWITZERLAND TAIWAN, PROVINCE OF CHINA THAILAND TURKEY UKRAINE UNITED ARAB EMIRATES and UNITED KINGDOM. | Open, Classified |  |  | 6583 | Computer
Z-1800-2014 | 68355 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | MOSAIQ, an oncology information system used to manage workflows for treatment planning and delivery. It supports information flow among healthcare facility personnel and can be used wherever radiotherapy and/or chemotherapy are prescribed. |  |  | May 23, 2014 | 2014 | Elekta, Inc. | When using CMA, if the user hovers the mouse pointer exactly above the direction dropdowns and then uses the mouse thumbwheel or touchpad scroll bar, the direction selection might change when the intention was only to move the mouse to a different area of the screen. | Software design | Elekta sent an Important Safety Notice on May 23, 2014, to all affected customers. The notice identified the product the problem and the action needed to be taken by the customer."Notify the affected customers of the issue - To be completed 06/22/2014"Provide a fix to the defect in an expeditious timeframe. It is mandatory for all customers to apply the fix - Version 2.50.05 SP8 &amp; Version 2.60 - To be released 05/16/2014"Follow up with the affected customers for: -Confirmation receipt of the LIN, AND -Confirmation of installation of patch for the fix- To be completed 11/30/2014. For further questions please call (404) 993-5886. |  | 723 | 723 | Worldwide Distribution - USA (nationwide) including Puerto Rico, and the countries of Albania, Algeria, Argentina, Australia, Austria, Bahrain, Bangladesh, Belarus, Belgium, Bosnia and Herzegovina, Brazil, Bulgaria, Canada, Chile, China, Colombia, Croatia, Cyprus, Czech Republic, Denmark, Ecuador, Egypt, Estonia, Finland, France, Gabon, Germany, Ghana, Greece, Honduras, Hungary, India, Iran, Iraq, Ireland, Israel, Italy, Japan, Jordan, Libya, Lithuania, Malaysia, Malta, Mexico, Morocco, Myanmar, Netherlands, New Zealand, Norway, Pakistan, Panama, Peru, Poland, Portugal, Romania, Russia, Saudi Arabia, Serbia, Slovakia, Slovenia, South Africa, Spain, Sudan, Suriname, Sweden, Switzerland, Taiwan, Thailand, Turkey, United Kingdom, Ukraine and Venezuela. | Open, Classified |  |  | 723 | Computer
Z-1724-2014 | 68255 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Synergy XVI |  |  | April 23, 2014 | 2014 | Elekta, Inc. | XVI can incorrectly calculate the target position of the treatment table. | Software design | Elekta is responsible for the correction of all products with this defect.1. Notifications regarding the possibility of this software error and subsequent table location error and potential hazards have been sent out. The notifications include steps to take to reduce the possibility of the occurrence of an injury.2. A software correction has been released for some versions of the XVI software. All versions of software with this error will be corrected by the end of July 2014.3. The software correction will enable verification that values have been assigned to table position parameters; failure to pass verification will inhibit the system. The software correction will also detect internal communication errors (incomplete registrations) and inhibit the system when those errors are detected.4. The update will be mandatory and will be made at no cost to the users.CDRH approves this CAP. You may proceed with implementation of the CAP.For further questions please call (770) 670-2422. |  | 380 | 380 US | Worldwide Distribution - USA including the states of AZ, CT, IA, OH, MD, MI, MN, NV, TX and VA., and the countries of India, Ireland, Malta, Netherlands, South Africa and UK. | Terminated | July 17, 2017 | 1181 | 380 | Computer
Z-1801-2014 | 68216 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | EZPass Suture Retriever, Nitinol, Single Pack, Part number 10007380. Orthopedic Manual Surgical Instrument, Sterile, single use arthroscopic instrument utilized in conjunction with a nitinol wire to aid in passing suture through soft tissue. |  |  | April 15, 2014 | 2014 | Biomet, Inc. | The nitinol wire or monofilament may buckle within the EZPass Suture Passer and become kinked making the wire/monofilament difficult to use. | Device Design | On 2/28/14, customers received a letter stating these devices have been recalled following an investigation which identified that the nitinol wire or monofilament may buckle within the EzPass Suture Passer and become kinked making the wire/monofilament difficult to use. The letter also stated, there is no adverse health outcome expected for the patient and this action requires the immediate location and discontinued use of the product and its return to Biomet. |  | 100 | 100 units | Nationwide Distribution. | Terminated | July 07, 2015 | 448 | 202 | Computer
Z-1785-2014 | 68292 | No | KNT | Gastrointestinal tube and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Convatec Flexi Seal Control Fecal Management System Rx Only |  |  | April 29, 2014 | 2014 | Convatec Inc. | Convatec Inc. is recalling Flexi-seal FMS Control Kit since it does not have a 510K. In addition, an internal assessment of product performance, including a review of customerfeedback, has confirmed that Flexi-Seal Control is not meeting expectations or those of customers. Specifically, the Auto-Valve feature that is unique to the Flexi-Seal Controlhas not been consistently performing relative | Device Design | Convatec Inc. sent an Urgent Voluntary Medical Device Recall letter dated April 29, 2014 via UPS. The letter identified the affected product, problem and actions to be taken. Customers are asked to immediately stop distributing and quarantine all recalled lots. Affected products must be returned to Distributor. Contact your Distributor for shipping instruction. Customers should completed the attached return response form and return it to the address on the response form. Questions can be directed to Gregory Kowzun at (336) 547-3710. Convatec issued a press release on 7/3/2014. |  | 97887 | 73,021 units US; 24,866 units OUS | Worldwide Distribution - USA Nationwide | Terminated | August 28, 2015 | 486 | 97887 | Computer
Z-1668-2014 | 68197 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | ACUSON SC2000 Ultrasound Systems with software version 3.5 (VA35) and using Stress Echo Feature. Ultrasound imaging system. |  |  | April 23, 2014 | 2014 | Siemens Medical Solutions USA, Inc. | Potential malfunction due to a software issue for ACUSON SC2000 ultrasound systems using software version 3.5 (VA35) when using the stress echo feature which could cause the need for the study to be repeated. | Software design | Siemens sent a Consumer Safety Advisory Notification letter on April 23, 2014, to all affected consignees. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to discontinue use of the stress echo feature with software version .5. For studies acquired with systolic-only captures: review all patient studies on exams performed after software revision 3.5 was installed on your system. If you are unsure when this software was installed on your system,please contact your local Customer Service representative. Foreign distributors will implement distribution of the field correction to their customers.For questions regarding this recall call 650-694-5398. |  | 177 | 177 units | Worldwide Distribution - USA (nationwide) and Internationally to Australia, Belarus, Dominican Rep., France, Germany, India, Indonesia, Israel, Japan, Mexico, New Zealand, Norway, Philippines, Poland, Republic of Korea, Romania, Russian Fed., Saudi Arabia, Singapore, Sweden, Switzerland, Turkey, U.A.E., Ukraine, and United Kingdom. | Terminated | November 12, 2015 | 568 | 177 | Computer
Z-1667-2014 | 68195 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | ACUSON SC2000 Ultrasound Systems with software version 3.5 (VA35); ultrasound imaging system. |  |  | April 23, 2014 | 2014 | Siemens Medical Solutions USA, Inc. | Software issues with ACUSON SC2000 ultrasound with software version 3/5 (VA35) - may incorrectly display the color flow data in the wrong position and when Doppler Frequency is adjusted, the system does not restart the pulsed wave Doppler or continuous wave Doppler waveform. | Software design | Siemens sent an Customer Safety Advisory Notification letter on April 23, 2014, to all affected customers. Distributors in foreign countries will implement delivery of notification to their customers." Use the factory default presets with software version 3.5 instead of user-defined presets restored from software version 2.0or 3.0. Review all patient studies on exams performed after software revision 3.5 was installed on your system. If you are unsure when this software was installed on your system, please contact your local Customer Service representative." After adjusting the Doppler frequency, freeze and unfreeze the system. This action refreshes the display of the Doppler spectrum resulting in accurate measurements. Additional measurements performed after system freeze results in correct measurement values.Your local Customer Service Engineer will contact you to schedule an appointment to replace your existing version of software, free of charge. If you have any questions, please contact your local service support person forinformation regarding timelines and status.Until the new software has been installed on your system, please share this information with all personnel within your organization who need to be aware of this issue. As always, patient safety issues are a high priority, and we alert customers of product concerns which may occur under various conditions. To date, no patient injury has been reported as a result of these issues.We sincerely regret any inconvenience this condition may cause in your daily operations. For further questions call (650) 694-5398. |  | 257 | 257 units | Worldwide Distribution - USA including the states of AL, AZ, AR, CA, CO, CT, FL, GA, ID, IL, IN, KS, LS, MD, MI, MN, MO, NH, NY, NC, OH, OR, PA, SC, SD, TN, TX, UT, VA, WA and WI., and the countries of Australia, Austria, Dominican Rep, France, Germany, Iceland, India, Indonesia, Ireland, Israel, Italy, Japan, Mexico, New Zealand, Norway, Philippines, Poland, Portugal, Republic Korea, Romania, Russian Fed, Saudi Arabia, Sweden, Switzerland, Turkey, U.A.E. and United Kingdom. | Terminated | November 12, 2015 | 568 | 257 | Computer
Z-1665-2014 | 68092 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | The Centricity PACS Workstation is intended for use as a primary diagnostic and analysis tool for diagnostic images by trained healthcare professional, including radiologists, physicians, technologists, clinicians, and nurses. It is also intended for use as a clinical review workstation throughout the healthcare facility. The workstation interface provides the user with a means to display, manipulate, archive, print, and export images when connected wit the Centricity PACS infrastructure. |  |  | April 03, 2014 | 2014 | GE Healthcare | It has been identified that in a reject image workflow where multiple studies are simultaneously opened on the Centricity PACS RA1000 workstation and images are rejected from the active displayed study, images for the previously displayed study may also be rejected. The issue can occur with a previously displayed study that remains open in the background. If one or more, but not all, images in | Software design | An URGENT MEDICAL DEVICE CORRECTION letter dated April 3, 2014 was sent to all direct accounts. The letter included instruction to customers to use until the permanent fix for this issue is deployed: 1) Check to see if the DelRejImg task is active and running. a) Log into CA Tool. b) From the Navigation menu, select System Monitor. c) From the pull down menu, select Dbase Maintenance. d) Scroll to the DelRejImg task (this list is alphabetical). e) Review the Task Status. Possible statuses are IDLE or SUSPEND f) If the task status is set to SUSPEND The task is not running. g) If the task is set to IDLE, move to Step 2. 2) Check the Minimum Rejected Image Deletion Time. a) From the Navigation menu, select Sys. Configuration. b) From the pull down menu, select PACS Configuration. c) Scroll to Minimum Rejected Image Deletion Time (this list is alphabetical). d) Review the Value field. e) If the value field is a negative integer (such as the default setting of -1), the Delete Rejected Image process is NOT running and there is no patient safety hazard. If the value field is a positive integer (such as 1 or higher), the Delete Rejected Image process IS running and there is a potential patient safety hazard. 3) A GE Healthcare Service Representative will remove the DelRejImg task from the task table in the Centricity PACS database. This will prevent the DelRejImg task from running, and prevent any rejected images from being deleted. This feature is being removed entirely from the 3.2 release family of Centricity PACS and will be reinstated in a future 4.X workstation release of Centricity PACS. 4) Perform the reject image workflow with only one patient study open at a time. A future release of Centricity PACS will address the unintentional rejected image issue. Until a correction for the unintentional rejected image issue is available, a GE Healthcare service representative will remove the DelRejImg task table in the Centricity PACS database. This will remove the re |  | 297 | 297 units | Worldwide Distribution-USA (nationwide) including the states of AK, AL, AZ, CA, CO, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NE, NH, NJ, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, VA, VT, WA, WI, and WV, and the countries of Australia, Canada, Chile, France, Germany, India, Ireland, Israel, Italy, Japan, Korea (South), Kuwait, Lebanon, Netherlands, Saudi Arabia, Singapore, Spain, Sweden, Switzerland, Taiwan, Turkey, and United Kingdom. | Terminated | June 05, 2015 | 428 | 297 | Computer
Z-1445-2014 | 67755 | No | OOE | Ophthalmic laser. | Ophthalmic | Ophthalmic | 510(k) | Alcon LenSx Ophthalmic Laser, for use in cataract surgery for the creation of corneal cuts/incisions, anterior capsulotomy and laser phacofragmentation. |  |  | March 14, 2014 | 2014 | Alcon LenSx, Inc. | Alcon LenSx received reports of unexpected downward motion of the gantry. | Device Design | Alcon LenSx Inc advised customers of a Medical Device Correction being initiated for specific serial numbers of the LenSx Laser System for use in cataract surgery for the creation of corneal cuts/incisions, anterior capsulotomy and laser phacofragmentation.Alcon sent customer notification letters via mail during the week of March 17, 2014. The letter titled "Medical Device Correction", advised customers of the correction and provided: the product's description with serial number(s), description of the potential condition, advise on action to be taken by the user, transmission of the notice, and a contact reference person. The letter also contained a section for customers to sign and return the page of the Notice to Alcon. |  | 75 | 75 | Worldwide Distribution. | Terminated | October 16, 2014 | 216 | 75 | Computer
Z-1315-2014 | 67549 | [6 empty] | Omni Micro-electrode/reference electrode for cobas b221 analyzer, Model/Catalog/Part Number: |  |  | February 28, 2014 | 2014 | Roche Diagnostics Operations, Inc. | Reference electrode used beyond the guaranteed in-use 52 week lifetime,may leak and potentially cause erroneous pH and/or sodium results. Other Ion parameters are not affected. | Device Design | Roche Diagnostics Operations sent an Urgent Medical Device Correction letter dated February 27, 2014 via UPS Ground (receipt required) to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to complete the enclosed fax form and fax it to 1-888-345-0253. For questions contact Roche Customer Business Support Center at 1-800-428-5076. |  |  |  | Worldwide Distribution - USA nationwide including PR and the country of Guam. | Open, Classified |  |  |  | Computer
Z-1569-2014 | 67927 | No | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | Shimadzu Digital Radiography X-Ray System, Catalog No. DAR-8000f. |  |  | April 03, 2014 | 2014 | Shimadzu Medical Systems | Shimadzu Corporation is recalling the Shimadzu Digital X-Ray System because they have received a report of a lost image under certain circumstances. | Software design | A Service Bulletin and a customer notification letter were sent on 4/3/2014 for all the customers who purchased the Shimadzu Digital X-Ray System (Catalog No. DAR-8000f). The notification informs the customers of the problems identified and the actions to be taken. Customers with questions or concerns about the device correction, are instructed to contact their local Shimadzu representative or contact Shimadzu National Technical Support at (800) 228-1429. |  | 685 | 685 units total (14 units in US) | Worldwide Distribution-USA (nationwide) and the countries of Japan, China, Germany, Australia, Thailand, Sri Lanka, South Africa, UAE, Saudi Arabia, Korea, and Taiwan. | Terminated | June 25, 2015 | 448 | 685 | Computer
Z-1565-2014 | 63145 | No | MAX | Intervertebral body fusion device. | Orthopedic | Orthopedic | 510(k) | SpineFrontier S-Lift Lock Shaft Inserter, an instrument used with S-LIFT Intervertebral Body Fusion Device in spinal fusion surgery. Product Usage: The S-Lift Lock Shaft SI50001 mates with SI50002, Implant Inserter, to aid in securing the implant to the whole implant inserter assembly, SI50000. |  |  | December 23, 2011 | 2011 | SpineFrontier, Inc. | It has been verified through field use that there are two possible breakages that may occur when using SI50001 Revision D Lock Shaft Inserter, possibly resulting in instrument fragments entering the sterile field:Impact Cap: The impact cap on the Lock Shaft Inserter can shear off under high-impact force during insertion and removal.Pin: The pin used to guide the Lock Shaft Inserter into the ma | Device Design | Spine Frontier sent an Urgent Notification dated December 23, 2011, consignees. The letter identified the affected product, problem and actions to be taken. Customers were instructed to complete the attached acknowledgement form. For questions call 866-914-7717 |  | 30 | 30 devices | US Nationwide - OR, FL, MA, TX, CO, OK, IL, and MO | Terminated | May 20, 2014 | 879 | 30 | Computer
Z-1318-2014 | 67716 | No | FLF | Ethylene oxide gas sterilizer. | General Hospital | General Hospital | 510(k) | 3M" Steri-Vac" Gas Sterilizer/Aerators, Model #'s 4XL,5XL, 8XL |  |  | March 17, 2014 | 2014 | 3M Company - Health Care Business | Recently, 3M received a single report of a potential exposure to Ethylene Oxide (EO) at a hospital installation involving 3M" Steri-Vac" Gas Sterilizers. The cause was traced to a rare situation involving a failure of the EO cartridge puncture assembly. This assembly was only in production during the period November 10, 2006 through December 14, 2007 and in service during the period February 7, 20 | Device Design | 3M sent an "Urgent Medical Device Field Correction" letter dated March 6, 2014. The letter was sent to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. You are being notified because our records indicate you received a unit that contains the cartridge puncture assembly that is the subject of this field correction. 3M will contact you to schedule an appointment for a field correction on any unit in the range of serial numbers provided in the table. 3M's goal is to service the affected units with minimal disruption to your facility. Units manufactured outside of the dates November 10, 2006 to December 14, 2007 are not affected. .For Further Information contact 3M Health Care Service Line at 1-800-228-3957 or your local 3M representative. |  | 419 | 73 US, 346 OUS | Worldwide Distribution - USA (nationwide) Distribution including the states of AL, AZ, CA,CO, FL, GA, HI, IL, IN, IA, KS, KY, MD, MA, MI, MN, MS, NE, NJ, NY, NC, ND, OH, OK,OA, SC, TX, WA, WV and WI. and the countries of COLOMBIA, VENEZUELA, UAE, THAILAND, GERMANY, CHINA, KOREA, CANADA, JAPAN, RUSSIA, INDIA, POLAND, PERU, AUSTRIA, ISRAEL, INDONESIA, SPAIN, CHILE, MALAYSIA, SWITZERLAND, EL SALVADOR, ECUADOR, LEBANON, MEXICO, ARGENTINA, BULGARIA, TAIWAN, TURKEY,ROMANIA, GUATEMALA and NEW ZEALAND. | Terminated | January 22, 2015 | 311 | 419 | Computer
Z-1316-2014 | 67586 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | IMMULITE 2000; IMMULITE 2000XPi; VersaCell |  |  | February 03, 2014 | 2014 | Siemens Medical Solutions Diagnostics | There is a potential for incorrect results to be reported from processing of a sample tube. | Device Design | Siemens sent an Urgent Medical Device Correction letter dated February 3, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.The customers was told how to correct the issue if it has occurred, and how to prevent the issue from occurring. The completed form should be faxed to (302) 631-7597. Any questions can be directed to Siemens Technical Support Representative. Siemens plans to implement a correction at a later time. |  | 1006 | 1006 units | Nationwide and Foreign Distribution. | Open, Classified |  |  | 1006 | Computer
Z-1549-2014 | 68067 | No | FMT | Infant radiant warmer. | General Hospital | General Hospital | 510(k) | PandaÂż Freestanding Infant Warmers, integrated with Nellcor SpO2. |  |  | April 02, 2014 | 2014 | GE Healthcare | Medical device software may be defective and may cause user-set alarm limits to result in false positive or false negative Oxygen saturation and pulse rate alarm notifications to the caregiver. | Software design | GE Healthcare notified Biomedical/Clinical Engineering managers, Labor &amp; Delivery/ NICU Nurse managers, Respiratory Therapy Managers and Directors of Risk Management by letter 04/02/2014. End users were advised that they may continue to use the system provided they follow the GE recommended actions:1) If in use with a patient, turn off the Nellcor SpO2 function in the control panel and switch to an approved alternate form of SpO2 measurement.2) Prior to using the affected GE warmer with the Nellcor SpO2 function, or resuming the use of this function with any patient, contact your hospital Bio-Med or qualified technician to check the Nellcor SpO2 alarm functionality using either an adult Nellcor SpO2 sensor or a Nellcor SpO2 simulator following the applicable instructions in the Appendix 3) If the test activates a Low Pulse Rate alarm, the user-adjusted Nellcor SpO2 alarm limits are functioning properly and the warmer can be placed back in clinical service.End users were also advised that alarm functionality must be re-checked following every service event requiring access to the internal electronics and during annual preventative maintenance, until revised software is installed.4) If the test does NOT activate a Low Pulse Rate alarm, remove the warmer from clinical service and contact your local GE Healthcare Service Representative. |  | 21 | 21 units | Worldwide Distribution. | Open, Classified |  |  | 1110 | Computer
Z-1322-2014 | 67574 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Integrity 1.1. |  |  | March 28, 2013 | 2013 | Elekta, Inc. | The problem is that the "static tolerances" from the calibration files (loaded database) are used for beam delivery instead of the machine calculated values. | Software design | An Important Field Safety Notice (IFSN) has been released and distributed to all customers. Integrity 1.1 Incorrect Static Tolerances for Linac Geometric Axes, 200 01 502 032. The notice explains the issue and that the user should apply the static Tolerance Updater tool to check tolerances and adjust as required. Important Field Safety Modification Digital Accelerator FCO Ref: 200 02 502 031 Integrity R1.1 Static Tolerance Updater Tool was released 03-2013. Elekta has released Integrity version 1.2 which contains the permanent fix for this issue. The permanent solution fixes the error that made the digital accelerator protection functions use the value loaded from the calibration database. The solution now ensures the protection functions use the values the digital accelerator calculates. Important Field Safety Notice 200 01 502 035 has been released to inform customers the solution is available. Important Field Safety Modification 200 02 502 036 will track the installation of the software by Elekta service engineers. |  | 498 | 498 | Worldwide Distribution-USA (nationwide) including the states of AL, AZ, CT, FL, GA, KY, LA, MD, MO, NM, NY, NC, OH, OK, PA, TN, TX, and WA and the countries of Argentina, Australia, Austria, Belarus, Belgium, Brazil, Canada, China, Colombia, Croatia, Czech Republic, Denmark, Ecuador, Egypt, Estonia, France, Germany, Greece, Honduras, Hong Kong, India, Indonesia, Iraq, Ireland, Italy, Japan, Jordan, Malaysia, Mexico, Morocco, Myanmar, Netherlands, New Zealand, Nigeria, Norway, Pakistan, Panama, Peru, Poland, Romania, Russia, Saudi Arabia, Senegal, Singapore, Slovakia, South Korea, Spain, Sudan, Sweden, Switzerland, Taiwan, Thailand, Turkey, and United Kingdom. | Terminated | July 26, 2017 | 1581 | 498 | Computer
Z-1331-2014 | 67741 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | Siemens Medical Solutions, ACUSON P300 Ultrasound system with P300 PA 230E cardiac probe (Esaote product 7348). |  |  | March 07, 2014 | 2014 | Siemens Medical Solutions USA, Inc. | ACUSON P300 ultrasound systems using the PA230 transducer with a user-defined preset which includes the 100% setting for output power could cause the transducer surface temperature to exceed the maximum allowable limit. | Device Design | Siemens sent an Customer Safety Advisory Notifications on March 7, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.A new version of software resolving this error will be provided free of charge. Your local Customer Service Engineer will contact you to schedule an appointment to make the necessary repairs on your system. If you have any questions, please contact your local service support person for information regarding timelines and status. Until the necessary repairs have been completed on your system, please share this information with all personnel within your organization who need to be aware of this issue. As always, patient safety issues are a high priority. To date, no patient injury has been reported. This problem was discovered as part of our ongoing quality process.For further questions please call (650) 694-5398.We sincerely regret any inconvenience this condition may cause in your daily operations. |  | 675 | 675 | Class II Recall - Worldwide Distribution - US Distribution and the countries of Australia, Austria, Belgium, Bhutan, Bolivia, Botswana, Bulgaria, Canada, Chile, Colombia, Croatia, Dem. Rep. Congo, Denmark, Ecuador, France, Germany, Greece, India, Indonesia, Ireland, Israel, Italy, Ivory Coast, Japan, Kuwait, Malaysia, Mauritius, Mongolia, Morocco, Mozambique, Myanmar, Namibia, Netherlands, Norway, Peru, Philippines, Poland, Portugal, Republic Korea, Romania, Saudi Arabia, Serbia, Seychelles, Singapore, South Africa, Spain, Switzerland, Taiwan, Trinidad,Tobago, Turkey, U.A.E. and United Kingdom. | Terminated | November 12, 2015 | 615 | 675 | Computer
Z-1335-2014 | 67758 | No | NAY | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | 5X250 MM Curved Cannula, Arm 2 |  |  | March 13, 2014 | 2014 | Intuitive Surgical, Inc. | Intuitive Surgical has become aware of the potential for certain lots of the Single-Site 5mm Curved Cannulae to become damaged during use. The second part of this action is to update Single-Site labeling. These labeling changes instruct users how to inspect cannulae prior to use and to appropriately handle cannulae during use. | Device Design | Intuitive sent an Urgent Medical Device letter dated March 25, 2014, to all users of cannulas that have the potential to be damaged during use as well as a second letter to all single site cannula users that provided clarity to instruction for use for inspections and handling of the product.Certain lots manufactured prior to August of 2012, and users were asked to use Attachment A to help identify affected lots numbers. Phase 1 - Cannulas should then be inspected with information provided in Attachment B. Any damaged devices should be quarantined and returned to intuitive. If inspection does not find damage, the device may continue to be used. Any other cannulae and any non damaged product should be used according to Attachment C's instructions for use.Phase 2 - Once replacement cannulae are available, the undamaged cannulae should be returned.The Acknowledgement form should be completed and returned to Intuitive. Intuitive will contact consignees when replacement cannulae are available.Questions should be directed to Customer Service at 800-876-1310.Those consignees that use the 5 mm canulae but do not have affected product were provided with a letter that only includes the Single Site Cannula inspection and Use addendum. |  | 5217 | 5217 total cannulae | Worldwide Distribution - USA and the countries of Australia, Belgium, Brazil, Chile, China, Columbia, Denmark, France, Germany, India, Italy, Lebanon, Monaco, Netherlands, Norway, Romania, Saudi Arabia, South Korea, Spain, Sweden, Switzerland, Turkey and United Kingdom. | Terminated | September 03, 2014 | 174 | 5217 | Computer
Z-1528-2014 | 67947 | Yes | MRN | Nitric oxide administration apparatus. | Anesthesiology | Anesthesiology | 510(k) | INOmax DSIR, model 10007, Nitric Oxide Delivery System, with Software version 3.0.0. The INOmax DS delivery system delivers INOMAX (nitric oxide for inhalation) therapy gas and provides continuous monitoring of inspired O2, NO2, and NO, and a comprehensive alarm system. The primary targeted clinical setting is the Neonatal Intensive Care Unit (NICU) and secondary targeted clinical setting is the transport of neonates. |  |  | March 27, 2014 | 2014 | INO Therapeutics (dba Ikaria) | An anomaly impacts the monitoring system of INOmax DSIR Plus (software 3.0.0). When the integrated pneumatic backup switch is activated with low or no gas flow through the injector module, an intended NO cell failure alarm will occur. In addition, an unintended NO2 cell failure alarm may also occur. The NO/NO2 cell failure alarms are caused by an average NO/NO2 circuit concentration greater tha | Software Design Change | Ikaria sent an Urgent Field Corrective Action letter via FedEx on March 27, 2014. The letter identified the issue, and requested that all INOmax DSIR with software version 3.0.0 be returned to Ikaria. Customers with questions were instructed to contact Ikaria Customer Care at 1-877-566-9466.For questions regarding this recall call 877-566-9466. |  | 22 | 22 | Nationwide Distribution including DC and VA. | Terminated | November 21, 2014 | 239 | 22 | Computer
Z-1512-2014 | 67802 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | GemStar Docking Station, used with GemStar infusion pump, in combination with other GemStar accessories such as the GemStar Lock Box and patient bolus cord. The docking station is also used for hosting the GemStar pump for general infusions. List Numbers: 13075-XX-XX. |  |  | February 18, 2014 | 2014 | Hospira Inc. | There are two situations that may occur when using the GemStar Docking Station, List Number 13075-XX-XX, in conjunction with the GemStar infusion pump: 1) when the Docking Station is used in conjunction with a GemStar Phase 3 pump (List 13000-XX, 13100-XX, or 13150-XX) the potential exists for the GemStar Phase 3 pump to fail to power up while connected to the Docking Station, and 2) when either a GemStar Phase 3 (List 13000-XX, 13100-XX, or 13150-XX)or GemStar Phase 4 pump (List 13086-XX, 13087-XX, or 13088-XX) is used in conjunction with both a Docking Station and an External Battery Pack accessory (List 13073-XX) there exists a possibility that the GemStar pump will display error code 11/003 and give an audible alarm indicating excessive input voltage from the external sources. If the GemStar pump detects what is perceived to be more than 3.6 Volts measured on the external voltage input, the pump will stop the infusion and alarm both with an audible sound as well as a visual alarm. | Device Design | Hospira sent an URGENT MEDICAL DEVICE CORRECTION letter dated February 18, 2014, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. The letter included the following instructions: 1) To avoid failure to power up, turn the pump on first, before connecting the pump with the docking station. This will prevent the failure to power up. 2) To mitigate the potential for an 11/003 error code, remove the External Battery Pack accessory (List 13073) from the Docking Station in conjunction with an External Battery Pack accessory (List 13073), this practice should not continue. Please contact Hospira to discuss an appropriate alternative. Additional instructions included: 1) Inform potential users within the recipient's organization of the notification. 2) Should the facility experience a failure to power up or an 11/003 error code, report the issue to Hospira by calling 1-800-441-4100 (Monday - Friday, 8:00 AM - 5:00 PM CT) or e-mail ProductComplaintsPP@hospira.com. 3) Compelte the attached reply form and return it to the fax number or e-mail address on the firm, even if there is no affected product in inventory. 4) If th products were further distributed, notify the accounts that received the products identified in the notification and ask them to contact Steriycle at 866-792-5451 (Monday - Friday, 8:00 AM - 5:00 PM ET) to receive a reply form.For questions regarding this recall call 800-441-4100. |  | 26765 | 26,765 docking stations total (6866 in US) | Worldwide Distribution - USA including AK, AL, AR, AZ, CA, CO, CT, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, WY, District of Columbia, Guam, Puerto Rico, and Internationally to Australia, Austria, Bahrain, Belgium, Canada, Chile, Egypt, France, Germany, Greece, Hong Kong, Italy, Jordan, Kuwait, Lebanon, Libya, Luxemburg, Netherlands, New Zealand, Oman, Portugal, Saudi Arabia, Slovenia, South Africa, Spain, Switzerland, Turkey, United Arab Emirates, and United Kingdom. | Terminated | January 23, 2017 | 1070 | 26765 | Computer
Z-1437-2014 | 67691 | No | HWT | Template for clinical use. | Orthopedic | Orthopedic | 510(K) Exempt | GLOBALÂż Anchor Peg Glenoid Plus |  |  | March 28, 2014 | 2014 | DePuy Orthopaedics, Inc. | Urgent Device Correction notifications are being sent to USA distributors with follow-up to surgeons concerning updated surgical techniques stressing the possibility of the shoulder pin guide breaking and potentially being left in the patient. | Device Design | Beginning 3/28/12014 DePuy United States Distributors and DePuy affiliate contacts for Canada and International will be notified via email. The US Distributors will be responsible for adhering to the directions outlined in the field communication, including identifying the Surgeon users and notifying them either by mail or in person with the Surgeon Letter and providing access to revised surgical techniques. The notice emphasized technical points regarding the use of the GLOBAL APG+ 2.5mm Breakaway Guide Pin (singleuse shoulder instrument) that have been added to the surgical techniques. Risks were explained and surgical techniques revised were identified and available through the distributor. Distributors were asked to identify surgeons using the device, filling out the response form. For clinical questions from surgeons, please contact DePuy Orthopaedics, Inc.s Scientific InformationOffice at 18885542482 (MF; 8 a.m.  5 p.m. EDT).For product related questions, please contact your local DePuy Orthopaedics, Inc., Sales Consultant." For questions about recall information provided: Kim Earle, Recall Coordinator, 5743714917 (MF; 8 a.m.  5 p.m. EDT) |  | 21675 | 21,675 | USA Nationwide Distribution including Canada | Terminated | December 04, 2014 | 251 | 21675 | Computer
Z-1438-2014 | 67862 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | Philips HeartStart XL+ Defibrillator/Monitor with system |  |  | April 02, 2014 | 2014 | Philips Medical Systems, Inc. | Software communication failure may occur on the HeartStart XL+ locking the user out of clinical mode and possibly causing a delay in therapy. | Software design | Philips Healthcare sent an Urgent Medical Device Correction letter dated March 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. This Field Safety Notice is intended to inform you about:" what the problem is and under what circumstances it can occur" the actions that should be taken by the customer/user in order to prevent risks for patients or users" the actions planned by Philips to correct the problem.A software communication failure may occur on the HeartStart XL+ locking the user out of clinical mode and possibly causing a delay in therapy.Please see the attached Field Safety Notice that provides information on how to identify affected devices and instructions on actions to be taken. Please follow the ACTION TO BE TAKEN BY CUSTOMER / USER section of the notice. To correct this issue Philips will provide a software upgrade to customers with affected units free of charge.This voluntary correction has been reported to the appropriate regulatory agency.If you need any further information or support concerning this issue, please contact your local Philips representative at 1-800-722-9377.Philips apologizes for any inconveniences caused by this problem. |  | 9872 | 9872 units | Worldwide Distribution - USA (nationwide) including the countries of: ARGENTINA AUSTRALIA AUSTRIA BELGIUM CANADA CHINA CZECH REPUBLIC DENMARK FRANCE GERMANY ITALY LATVIA NETHERLANDS NEW ZEALAND NORWAY PHILIPPINES PORTUGAL SOUTH KOREA SPAIN SWITZERLAND THAILAND TURKEY and UNITED KINGDOM. | Open, Classified |  |  | 9872 | Computer
Z-1449-2014 | 67762 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Cobas c8000 Modular Analyzer Series Cobas c502, Part Number 5964067001, a clinical chemistry analyzer intended for the in vitro quantitative / qualitative determination of analytes in body fluids. |  |  | March 13, 2014 | 2014 | Roche Diagnostics Operations, Inc. | The Initial Cassette Volume Check (ICVC) is not activated on the cobas c 502 module of the cobas 8000 modular analyzer series when manually filled cobas c packs are loaded. This issue occurs with tests that use cobas c pack MULTI or empty pre-labeled cobas c packs. The ICVC feature allows the reagent probe to dive into the reagent cassette and confirm the reagent level in the bottles. This ensures | Software Manufacturing/Software Deployment | March 17, 2014 all consignees will be notified by Urgent Medical Device Correction sent via UPS Ground (receipt required). The notice identified the issue, root cause, affected assays with clinical significance. Consignees were instructed to follow the steps listed in the Urgent Medical Device Correction to correctly activate the ICVC and verify the correction. If your system is not connected to the Internet through the cobas link, and you do not have the option to run the affected reagents on an alternate system (i.e., cobas c 311 or cobas c 501 analyzer), please contact the Roche Support Network Customer Support Center at 1-800-428-2336 to have an update scheduled via a Roche Field Representative. If your system is connected to the Internet through the cobas link, follow the steps below:" Remove all reagent cobas c packs on board your cobas c 502 module for the assays listed on pages 1-2." Dispose of these reagent cobas c packs according to your local waste management ordinances." Load new cobas c packs for the affected assays, and visually verify the ICVC is performed." Complete the attached fax form (5326-00-0314) and fax it to 1-888-276-6166." Please consult with the physician or pathologist at your facility to determine specific clinical implications for your patients." File this Urgent Medical Device Correction (UMDC) for future reference. |  | 76 | 76 units | USA Nationwide Distribution including San Juan PR | Terminated |  |  | 76 | Computer
Z-1662-2014 | 68211 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | IMPAX CV Pediatric Echo Reporting, IMPAX CV Echo Measurement Import via Optical Character Recognition (OCR) Service IMPAX. For the creation of digital structured reports for cardiovascular diagnostic studies such as echocardiography. |  |  | April 30, 2014 | 2014 | AGFA Corp. | Failure to effectively install and test a hotfix to correct the known software defect within the IMPAX CV 7.8 SU3-OCR Service. | Software design | AGFA Healthcare sent an URGENT FIELD SAFETY NOTICE LETTER dated April 30, 2014, to the consignee describing the problem that had occurred, the actions being taken by Agfa and the actions Agfa has required the site to perform. As part of the correction Agfa requires the consignee to review the affected reports to confirm that the data is correct. If errors are found the consignee will need to correct the data. The consignee was asked to acknowledge, via FAX-Back to 864-421-1664 or email, that the information was received and understood by the consignee. For questions the consignee was instructed to contact their local Agfa HealthCare organization at 877-777-2432.For questions regarding this recall call 864-421-1754. |  | 1 | One | Distribution to SC only | Terminated | September 18, 2014 | 141 | 1 | Computer
Z-1510-2014 | 67901 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | Nuclear gamma cameras (ADAC Vertex Plus, ADAC Solus, ADAC Vertex V60 &amp; ADAC Vertex Classic) Philips Medical System, Cleveland, OH 44143. |  |  | March 10, 2014 | 2014 | Philips Medical Systems (Cleveland) Inc | Slippage of the radius drive belt in the relative 180 degree orientation of the superior positioned detector head will allow the detector head to drift down and potentially come in direct contact with the patient. | Device Design | On 03/10/2014, the firm sent Urgent - Medical Device Correction letters to their customers. The letter identified the affected product, the reason for recall, the hazard involved, and the actions that should be taken by customers. Customers are to cease use of the identified procedures until implementation of the Field Change Order. Additionally, customers should sign and return the attached Customer Response Form within 10 days of receipt, acknowledging that they have read and understood the Field Safety Notice and Addendum. The Field Safety Notice must be placed in customers' User Documentation. Customers are to contact their local Philips Healthcare representative or office, if they need further information or support. |  | 1043 | 1,043 cameras | Worldwide Distribution -- USA, and the countries of Australia, Austria, Belgium, Brazil, Canada, China, Denmark, Egypt, France, Germany, Hong Kong, Ireland, Italy, Korea, Lebanon, Malaysia, Netherlands, Norway, Pakistan, Portugal, Russia, Singapore, Spain, Sweden, Thailand, &amp; United Kingdom. | Terminated | August 04, 2015 | 512 | 1043 | Computer
Z-1513-2014 | 67719 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | ExacTrac 6.0 is a patient positioning and monitoring system. |  |  | February 24, 2014 | 2014 | Brainlab AG | When using multiple radiation treatment targets within a single plan, the ExacTrac v.6.0.3 might position the patient to an unintended radiation treatment point. As a result, there is a possibility that the radiation treatment dose may be delivered to the unintended target position. | Software design | A FIELD SAFETY NOTICE / PRODUCT INFORMATION letters, dated February 24, 2014, were sent to all direct accounts. The letter identified the affected product, the reason for the recall, and included instructions for direct accounts to take while a permanent fix can be deployed. The letter also advised that a software update will be made available starting the end of April 2014. Customers with questions or concerns can contact Brainlab at 800-597-5911 or via e-mail at us.support@brainlab.com. |  | 41 | 41 systems (13 systems in US; 28 systems to foreign countries) | Worldwide Distribution -- USA, including the states of CA, IN, LA, MA, OH, TX, and WI, and the countries of Australia, China, France, Germany, Japan, Netherlands, and the United Kingdom. | Open, Classified |  |  | 41 | Computer
Z-1484-2014 | 67523 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | SIGMA SPECTRUM Volumetric Infusion Pump with Master Drug Library |  |  | February 07, 2014 | 2014 | Baxter Healthcare Corp. | Baxter Healthcare Corporation has issued an Urgent Device Correction for the SIGMA SPECTRUM Volumetric Infusion Pump with Master Drug Library due to repeated System Error 322 occurrences. | Software design | On February 7, 2014, Baxter Healthcare Corporation distributed Urgent Device Correction notices and Customer Reply Forms to their customers via first-class mail. Affected customers are instructed to contact Baxter Healthcare Corporation at 1-800-356-3454 (choose option 1) Monday through Friday during the hours of 7:00 am to 7:00 pm Eastern Time. Baxter technicians would determine if an adjustment or repair is necessary. Customers are asked to complete the attached Customer Reply Form and return it to Baxter either by faxing to 1-224-270-5457 or e-mailing a scanned copy to fca@baxter.com. |  | 277096 | 277,096 units (revised) | Worldwide Distribution - USA Nationwide including PR and the country of Canada. | Terminated | May 25, 2017 | 1203 | 277096 | Computer
Z-1511-2014 | 68032 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | GE Healthcare, CardiacVX [CardiacVX (1.1.0 Ext. d) and CardiacVX Flow (1.1.0 Ext. d) analysis software running on Advantage Workstation]. Model #'s 5434389-2 and 5434384-2. |  |  | April 07, 2014 | 2014 | GE Healthcare, LLC | GE Healthcare has recently become aware of a potential safety issue due to higher than actual flow values reported on CardiacVX and CardiacVX Flow running on Advantage Workstation. | Software design | GE Healthcare sent an Urgent Medical Device Correction" letter dated August 21,2013, to all affected consignees. The letter was addressed to Hospital Administrators /Risk Managers, Radiology Department Managers and Radiologists. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction and Contact Information. Customers were instructed until the issue is corrected on their system with the affected CardiacVX or CardiacVX Flow software for phase-contrast protocols customers use the fast CINE phase-contrast sequence with the magnitude weighting user CV option disabled. For directions on how to update user CV, please refer to the MR operator manual. If customers choose to perform imaging using magnitude weighting, do not utilize CardiacVX or CardiacVX flow for flow value calculations.with the affected product. Customers with questions were instructed to contact their local service representative.For questions regarding this recall call 262-513-4122. |  | 101 | 101 (1 US , 100 OUS) | Worldwide Distribution - USA including IN and Internationally to AUSTRALIA, CANADA, CHILE, FRANCE, ITALY, JAPAN, KOREA, KUWAIT, NEW ZEALAND, PANAMA, POLAND, SAUDI ARABIA, SPAIN, SWEDEN, SWITZERLAND, UNITED KINGDOM, | Terminated | May 29, 2014 | 52 | 101 | Computer
Z-1504-2014 | 68027 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | Siemens ACUSON X700 Ultrasound Systems with software version 1.0.00, 1.0.01, 1.0.02, 1.1.00 and 1.1.02. ultrasound imaging system. |  |  | April 02, 2014 | 2014 | Siemens Medical Solutions USA, Inc. | When fourSight 4D ultrasound imaging, 3-Scrape real time 3D imaging, or Anatomical M-mode are used during an Obstetric or Gynecology study, data from the previous patient could replace the current patient's date in the report. | Software Design Change | Siemens sent a Customer Safety Advisory Notification letter on April 2, 2014 to all affected customers via certified mail. The letter identified the affected product, problem and what actions to take until software is upgraded. For questions call 1-888-826-9702. |  | 490 | 490 | Worldwide Distribution - US Nationwide and the countries of Angola, Argentina, Armenia, Australia, Austria, Bangladesh, Belgium, Bosnia Herzeg., Brunei, Chile, Colombia, Croatia, Ecuador, Egypt, France, Germany, Greece, Hungary, India, Indonesia, Ireland, Israel, Italy, Japan, Kenya, Kuwait, Latvia, Libya, Liechtenstein, Malaysia, Mauritius, Montenegro, Namibia, Nepal, Nigeria, Norway, Palestinian Ter, Philippines, Poland, Portugal, Qatar, Rep. of Yemen, Republic Korea, Romania, Saudi Arabia, Slovakia, South Africa, Spain, Switzerland, Taiwan, Turkey, Turkmenistan, U.A.E., United Kingdom. | Terminated | November 12, 2015 | 589 | 490 | Computer
Z-1495-2014 | 67868 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Philips MDC PACS R2.X / IntelliSpace PACS DCX R3.1 software, software |  |  | April 02, 2014 | 2014 | Philips Medical Systems, Inc. | Philips MDC PACS Release R2.x and Philips IntelliSpace PACS DCX R3.x, may have incorrect density measurement on Enhanced CT/MR examinations | Software design | Philips Healthcare sent a Urgent Medical Device Correction Notification/Field Safety Notice dated March 2014, to inform customers of the issue, identifies details of the units affected, gives instructions on actions to be taken by the customer and identifies what action Philips plans to take to remedy the issue.Philips will replace the software free of charge.Philips is asking customers to follow the Action to be Taken by Customer/User section of the Urgent Medical Device Correction notification/Field Safety Notice:Until you receive a software upgrade from Philips, the user must configure the MRI and CT devices to send regular CT or MR DICOM image objects (not enhanced) to PACS.. This issue does not affect images that are not enhanced image objects.Further questions please call (978) 687-1501. |  | 373 | 373 units | Worldwide Distribution - US Distribution including Puerto Rico and the states of CA, FL, IL, NJ, OK and WA., and the countries of : Argentina Australia, Austria, Azerbaijan, Bahrain, Bangladesh, Bosnia, Brazil, Bulgaria, Chile, China, Egypt, Ethiopia, France, Germany, Ghana, Greece, Guinea, India, Iraq, Israel, Italy, Latvia, Libya, Lithuania, Malawi, Malta, Mexico, Mozambique, Nepal, New Zealand, Papua, New Guinea, Paraguay, Philippines, Poland, Portugal, Romania, Russia, Slovakia, Slovenia, Solomon Islands, South Africa, Spain, Thailand, Turkey, United Arab Emirates and United Kingdom. | Open, Classified |  |  | 373 | Computer
Z-1502-2014 | 67896 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | On-Board Imager - Exact Arm Elbow Motor Shaft (kVD, kVS, and MV) with ''Type 01" elbow motors. |  |  | March 20, 2014 | 2014 | Varian Medical Systems, Inc. | Varian has identified a failure with the shaft of the Type 01 elbow motor for Exact arms used in the On-Board Imager (OBI) device due to belt over-tensioning. | Equipment maintenance | Varian sent an Urgent Medical Device Correction letters March 20, 2014 to affected custoers via trackable mail. The letter was also distributed to Varian Sales, Marketing and Service organizations to inform them of the issues. The letter identified the affected product, description of problem, details, and recommended user actions to be taken. A Varian Customer Support Service representative will arrange a site visit to check the bel tension of the Type 01 elbow motors. If over tensioned belts are detected, the elbow motor will be replace with a Type O2 elbow motor. Questions may be directed to 888-827-4265. |  | 2622 | 2622 | Worldwide Distribution. | Terminated | May 01, 2015 | 407 | 2622 | Computer
Z-1492-2014 | 67678 | No | LHN | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Proteus 235 Proton Therapy System Legacy Software version 7.2.7 and higher, models 110, 112, 113, and 116. |  |  | December 04, 2013 | 2013 | Ion Beam Applications S.A. | Internal testing of the Proteus 235 Proton Therapy System revealed incorrect dose rate calculation in case of beam line option other than Zero (0). | Software design | Ion Beam Applications ( IBA) sent a recall notification letter dated on/about December 5, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Short term action: A user notice (see annex 1) has been sent to impacted sites where the affected legacy PTS software version was delivered. This user notice explains the workaround required to be put in place in order to fully mitigate the risk and the 1BA action plan to permanently fix this issue.Corrective action:A patch on the PTS Legacy SW was requested and in implementation. In the same time, the solution was put in main code of PTS legacy software version 7 to be integrated in the next release.The customer representative signing underneath confirms that the user notice above is communicated to the appropriate employees.We apologize in advance for any inconvenience this may cause and we would like to thank you for your cooperation.The IBA Site Staff or the Operations manager is available to provide you with additional information and/or guidelines if necessary.Further questions please call : 215-972-7777 |  | 4 | 4 | US Distribution including the states of OK, IL, NJ, and WA. | Terminated | June 12, 2014 | 190 | 4 | Computer
Z-1442-2014 | 67781 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | SpaceLabs Arkon Anesthesia Delivery System with Version 2.0 Software installed. This system includes a ventilator. |  |  | March 10, 2014 | 2014 | Del Mar Reynolds Medical, Ltd. | The System may go into a controlled failed state if a series of parameters on the user interface are activated in a specific sequence. | Software in the Use Environment | Spacelabs Healthcare sent an Urgent Medical Device Correction letter, dated 11 March 2014, to consignees. The letter identified the product, the problem, and the actions to be taken by the customer. The letter stated that Spacelabs Field Service personnel will be contacting consignees to schedule a convenient time to install, at no cost, a software upgrade.Consignees with questions about this corrective action program, please contact Spacelabs Healthcare at 1-800-522-7025 and select 2 for Technical Support. |  | 16 | 16 units distributed in the US | USA Nationwide Distribution in the states of North Carolina and South Carolina | Terminated | August 20, 2014 | 163 | 16 | Computer
Z-1488-2014 | 67905 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | GE Vivid E9 Ultrasound System, a general-purpose ultrasound system, specialized for use in cardiac imaging. |  |  | March 27, 2014 | 2014 | GE Healthcare, LLC | GE Healthcare has recently become aware of a potential safety issue where the probe power surveillance settings for 4 probes (C1-5-D, C2-9-D, IC5-9-D and M5Sc-D) are incorrect. This may result in a situation where probes can overheat. | Software design | GE Healthcare sent an "Urgent Medical Device Correction" letter dated March 27, 2014. The letter was addressed to Hospital Administrators / Risk Managers, Managers of Radiology / Managers of Cardiology. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction and Contact Information (Contact your local GE Healthcare service representative). |  | 447 | 447 (81US, 366 OUS) | USA Nationwide Distribution in the states of: CA, CO EL, GA, IN, KS, KY , LA, MA, MI, MO, MT, NY, NM, NV, US: OH, TX, VA, WI.OUS: ARABIC EMIRATES, AUSTRALIA, AUSTRIA, BELGIUM, COLOMBIA, CZECH REPUBLIC, DENMARK, FINLAND, FRANCE, GERMANY, GREAT BRITAIN, INDIA,IRELAND, ITALY, JAPAN, KOREA, NETHERLAND, NORWAY, POLAND, PORTUGAL, RUSSIA, SINGAPORE, SOUTH AFRICA, SPAIN, SWEDEN, SWITZERLAND, TAIWAN. | Terminated | January 14, 2015 | 293 | 447 | Computer
Z-1464-2014 | 67763 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | Ingenuity Core 128, Computed Tomography X-ray systems, Philips Medical Systems (Cleveland), Inc. |  |  | March 06, 2014 | 2014 | Philips Medical Systems (Cleveland) Inc | A customer reported that after upgrading to software version 3.5.5 from 3.5.4, reconstructions would intermittently not complete. | Software in the Use Environment | .The firm, Phillips, sent an " URGENT ÂżÂż Medical Device Correction" letter dated2013 March 06, to its costumers . The letter disc ribes the product, problem and actions to be taken. The costumers were instructed to Verify that all planned image reconstructions are completed, and Perform off-line reconstructions if expected images are missing. The firm will release A Field Change Order (FCO) to correct the issue with software version release 3.5.7.and A Philips Field Service Engineer will contact you to schedule the software installation at your site.If you need any further information or support concerning this issue, please contact your local Philips representative or local Philips Healthcare office. For North America and Canada, contact the Customer Care Solutions Center (1-800-722-9377, Option 5: Enter Site ID or follow the prompts). |  | 116 | 116 units | Worldwide Distribution:US ( Nationwide) Including stages of MI, OR, PA, TX &amp; WV; and Countries of: Australia, Bangladesh, Belgium, Brazil, Bulgaria, China, France, Germany, India, Indonesia, Ireland, Japan, Nicaragua, Oman, Palestine, Russia, South Korea, Spain, Switzerland, Thailand, Turkey &amp; United Kingdom. | Terminated | November 05, 2014 | 244 | 116 | Computer
Z-1462-2014 | 67867 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | The OEC 9900 Elite mobile fluoroscopy system (image intensified fluoroscopic x-ray) is designed to provide fluoroscopic and spot-film images of the patient during diagnostic surgical and interventional procedure. |  |  | March 31, 2014 | 2014 | GE OEC Medical Systems, Inc | GE Healthcare has identified a potential safety issue related to the screws that attach the workstation monitor spring arm to the extension arm on certain OEC 9900 systems, causing the spring arm and monitor to fall forward. | Device Design | GE Healthcare sent an Urgent Medical Device Correction letter dated March 31, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Use caution when positioning or extending the workstation articulating arm. In addition, refrain from positioning staff or patients directly underneath or in front of a workstation monitor.Your local GE Healthcare Field Engineer will contact you to coordinate the replacement of the screws, and ensure that your system is operating safely and meets all specifications. These activities will be performed at no cost to you.Should you have any questions or concerns regarding these corrections, please do not hesitate to contact the service team for further information at 800-874-7378 option 8.Please be assured that maintaining a high level of safety and quality is our highest priority. If you have any questions, please contact us immediately.UPDATE: On June 20, 2014. GE Healthcare Surgery began distributing an Urgent Medical Device Correction Letter to customers with affected units. The letter makes clear that this field correction applies to all user sites with the listed products, regardless of whether service work was performed on the monitor extension arm since 3/31/2014. |  | 453 | 453 units | Worldwide Distribution - USA, and the countries of Australia, Argentina, Canada, China, Cambodia, Columbia, Japan, Ecuador, Japan, Korea, Hong Kong, India, Israel, Libya, Mexico, Nigeria, Peru, Russia, Thailand, Taiwan, Saudi Arabia and Singapore. | Terminated | October 26, 2015 | 574 | 453 | Computer
Z-1481-2014 | 67535 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | RaySearch Ray Station Radiation Therapy Treatment Planning System, Stand-alone software. Versions 2.0, 2.5, 3.0, 3.5 and 4.0. |  |  | February 14, 2014 | 2014 | RAYSEARCH LABORATORIES AB | RaySearch Laboratories has recalled "RaySearch Ray Station 2.0, 2.5, 3.0, 3.5 and 4.0" software due to dose miscalculations for bolus/external/fixation support structures on CT-image. | Software design | RaySearch Laboratories sent a Field Safety Notice, Medical Device Correction letter dated February 14, 2014 to their customers via email. The letter identified the affected product, problem and actions to be taken by the user. Customers with questions may contact Freddie Cardel, Director of Customer Support at 1-877-778-3849 or via email at freddie.cardel@raysearchlabs.com. |  | 251 | 251 units (148 units domestically &amp; 103 units internationally) | Worldwide Distribution - USA Nationwide in the states of: AZ, CA, CT, CO, DE, FL, IL, NJ, NY, MA, MD, MS, OH, PA, TN, VA and WA. and the countries of Australia, New Zealand, Belgium, New South Wales, Republic of Korea, Italy, Germany, China, France, Switzerland, Netherlands, Spain, Austria, United Kingdom and Norway. | Open, Classified |  |  | 251 | Computer
Z-1572-2014 | 67921 | No | LKK |  |  | General Hospital | PMA | MedStream Programmable Infusion Pump, 40 mL, Product Code: 91-4201 |  |  | November 10, 2015 | 2015 | Codman &amp; Shurtleff, Inc. | A small number of specific pumps may trigger a low battery alarm at the beginning of the implanted life. | Device Design | Codman Neuro sent an Urgent Field Safety Notice to all affected customers on April 18, 2014. The letter identified the product the problem and the action needed to be taken by the customer. The letter states the reason for recall, information about managing patients with affected pumps, and returning affected non-implanted pumps.You are receiving this voluntary recall notification because our records indicate that you are the recipient of one or more MedStream Pumps affected by this recall. Please do not implant any identified products included in this recall. Please complete the enclosed Acknowledgement Form and fax the completed form to: 001-508-977-6665. It is important that we receive this acknowledgement form, even if you have no affected product in your inventory.You may also e-mail the form to us at Codmanproductcomplaints@dpyus.jnj.comWe request that you immediately check all inventory to determine if you have affected product. Please contact your Codman Neuro sales consultant to assist you in the return process. Please use the included instructions to report your inventory status and return affected product. Note this recall applies only to |  | 34 | 34 | Worldwide Distribution including the countries of Belgium Germany, Spain, United Kingdom, Ireland, Italy, Russia and Turkey. | Terminated | November 10, 2015 | 0 | 124 | Computer
Z-1573-2014 | 67970 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | SCC SA HIS version 4.0.7.0 when used with SoftMic version 4.0.7.0. |  |  | January 29, 2014 | 2014 | SCC Soft Computer | When an isolate is resulted without a SNOMED code, isolate information in the downstream system may be incomplete or missing. As a result, there is the potential for the delay or omission of patient treatment updates. | Software design | On 2/10/2014 SCC Soft Computer used a proprietary task management system to send a Correction Communication to the consignee with the affected software. The letter identifies the product, problem and corrective actions. |  | 1 | 1 | Distributed in IL. | Terminated | November 03, 2014 | 278 | 1 | Computer
Z-1575-2014 | 68015 | No | OLO | Stereotaxic instrument. | Neurology | Orthopedic | 510(k) | NavioPFSTM System, Orthopedic Sterotaxic Instrument, model #'s NPFS-02000, NPFS-02010, NPFS 02020. |  |  | April 11, 2014 | 2014 | Blue Belt Technologies MN | Blue Belt Technologies is conducting a voluntary recall for their Navio PFS system which is used during orthopedic knee surgery. One of the electronic modules inside the computer cart assembly may have a defect in one of its components that could result in a latent failure. | Software design | Beginning April 3, 2014 consignees were hand delivered a Blue Belt Technologies, Inc. "Urgent Medical Device Recall" letter. The letter was addressed to Safety Officer / Regulatory Compliance Officer. The letter included the Description of the Issue, Description of Safety Concerns, Affected Product, Action to be taken by user, Transmission of this field safety notice and Contacts at Blue Belt Technologies. For questions in the USA contact Customer Support Blue Belt Technologies 2900 Northwest Blvd, Suite 40, Plymouth, MN 55441, USA, 1-763-452-4922 (Direct) lbrink@bluebelttech.com. In the European Union, Ross Lethbridge, European Product Manager, Blue Belt Technologies, Office number 207, 3000 Aviator Way. Manchester Business Park, Manchester, UK M22 5TG, Mobile: +44 7787 415069, rlethbridge@bluebelttech.com. |  | 11 | 11 | USA: CA, PA, OH, TX.OUS: BELGIUM, SCOTLAND, UNITED KINGDOM. | Terminated | June 10, 2014 | 60 | 11 | Computer
Z-1312-2014 | 67623 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Pinnacle3 Radiation Therapy Planning (RTP) System, software version 9.0. |  |  | February 11, 2014 | 2014 | Philips Medical Systems (Cleveland), Inc. | Philips Medical Systems have recently determined that a software nonconformance can cause incorrect beam geometry. This issue affects Pinnacle3 Radiation Therapy Planning (RTP) System, software version 9.0 manufactured from 09/29/09 to 12/01/2013. | Software design | Philips Healthcare sent an Urgent Field Safety Notice dated February 10, 2014, to all affected consignees. The letter included the Problem Description, Hazard Involved, How to Identify Affected Products, Action to be taken by Customer /User, Actions Planned by Philips and Further Information and Support. Consignees were instructed to review the following information with all members of their staff who need to be aware of the contents of this communication including clinics within their institution that are remotely connected to their Pinnacle Server. Customers were questions were instructed to contact their Philips representative. For customers in North America please contact Customer Care Solutions Center at 1-800-722-9377. For all other countries their local Philips Healthcare Office should be contacted.For questions regarding this recall call 1-800-722-9377. |  | 1016 | 1016 | Worldwide Distribution - USA including US: AK, AL, AR, AZ, CA, CO, CT, DC, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, and WV and Internationally to AUSTRALIA, AUSTRIA. BELGIUM, CANDA, CHINA, DENMARK, FRANCE, GERMANY, HONG KONG, HUNGARY, INDIA, IRELAND, ITALY, JAPAN, LUXEMBOURG, MALAYSIA, MEXICO, MONTENEGRO, NETHERLANDS, NEW ZEALAND, OMAN, RUSSIA, SOUTH ARICA, SOUTH KOREA. SPAIN,. SWITZERLAND, TAIWAN, THAILAND, TURKEY, EMIRATES, UNITED KINGDOM.. | Terminated | November 21, 2014 | 283 | 1016 | Computer
Z-1661-2014 | 68165 | Yes | LWP |  |  | Cardiovascular | PMA | Medtronic Temporary Pacemaker, Model 5348. Single Chamber External Pulse Generator (EPG) Temporary Pacemaker (AAI/VVI). |  |  | April 29, 2014 | 2014 | Medtronic Inc. Cardiac Rhythm Disease Management | Events related to a pacing rate outside of the intended setting, including events of sudden increased pacing rate up to the maximum setting of 180 pulses per minute (ppm), have been reported. This is caused by the development of high resistance on internal electrical connector contacts due to oxidation over time. | Device Design | Consignees were sent a Medtronic "Urgent Medical Device Correction" letter dated April 2014. The letter was addressed to Risk Manger, Health Care Professional. The letter included the Issue Description, Root Cause, Potentially Affected 5348 EPGs within the Serial Number Ranges, Malfunction Indications, and Recommended Actions. For questions related to EPG service policy or purchasing a replacement EPG, customers should contact your Medtronic representative. For technical questions, please call Instruments Technical Service at (1-800)638-1991. Medtronic asks that customers complete and sign the attached Customer Confirmation Certificate and fax it to Medtronic at 651-367-0612 to the attention of Customer Focused Quality or e-mail to RS.CFOFCA@Medtrnic.com. |  | 16337 | 16,337 (8,989 US; 7,348 OUS) | Worldwide Distribution -- USA, including the territories of Guam and Puerto Rico; and, the countries of Albania, Algeria, Australia, Austria, Bahamas, Bangladesh, Belarus, Belgium, Benin, Bolivia, Bosnia And Herzegovina, Brazil, Bulgaria, Burkina Faso, Chile, Colombia, Costa Rica, Cote D'Ivoire, Croatia, Cyprus, Czech Republic, Denmark, Ecuador, Egypt, El Salvador, Estonia, Faroe Islands, Finland, France, Georgia, Germany, Greece, Guadeloupe, Honduras, Hong Kong, Hungary, India, Iran, Ireland, Israel, Italy, Jamaica, Japan, Jordan, Kazakhstan, Kenya, Kuwait, Latvia, Lebanon, Lithuania, Luxembourg, Macau, Malaysia, Malta, Martinique, Mauritius, Mexico, Morocco, Nepal, Netherlands, Netherlands Antilles, New Caledonia, Norway, Oman, Pakistan, Palestine, Panama, Paraguay, Peru, Philippines, Poland, Portugal, Qatar, Reunion, Saudi Arabia, Senegal, Singapore, Slovakia, Slovenia, South Africa, South Korea, Spain, Sri Lanka, Sudan, Sweden, Switzerland, Taiwan, Thailand, Trinidad and Tobago, Tunisia, Turkey, Ukraine, United Arab Emirates, United Kingdom, Uruguay, Venezuela, and Yemen. | Terminated | October 08, 2014 | 162 | 16337 | Computer
Z-1660-2014 | 67963 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | LIFEPAK 1000 defibrillator. Affected Part Numbers: 320371500XXX |  |  | May 09, 2014 | 2014 | Physio-Control, Inc. | Physio-Control has become aware of incidents where customers have attempted to use theirLIFEPAK 1000 defibrillator and the device has shut down unexpectedly due to a very low battery.A defibrillator in this scenario has the potential to fail to deliver a shock, with the potential resultthat therapy is not delivered and a patient is not resuscitated. A software malfunction in the LIFEPAK 100 | Software Manufacturing/Software Deployment | On May 9, 2014 Physio-Control sent an Urgent Medical Device Recall Letter (dated May 2014) to all affected customers. The letter identified the product, the problem and the action needed to be taken by the customer. The customer notification letter will provide new instructions on how to check the functional and power readiness of the device and advise the customer on the actions to take based on the symbol displayed.Customers are asked to record the readiness indicator(s) shown on each device and fax the confirmation sheet back to Physio-Control at 1-866-448-9567. Customers are also directed to contact Physio-Control at 1-800-442-1142 to obtain replacement batteries, if needed. Customers are provided a Device Readiness Guide for use in interpreting the Readiness Display on the device, and are directed to refer to the Operating Instructions which were provided upon purchase.In addition to notifying Physio-Control of any potential quality problems or adverse reactions or events experienced associated to the use of these products, Physio-Control asks customers to report directly to the FDA through the MedWatch Adverse Reporting Program online, by regular mail or by fax.Physio-Control reminds customers that it is critically important in understanding what the device and battery indicators mean on yourdefibrillator and what actions you need to take as a result. At any time the battery charge can be verified by following the instruction provided on page 2-5 of the Operating Instructions. It is also important that you always carry a spare fully-charged battery, as stated in the Operating Instructions.Customers are directed to call Physio-Control at 1-800-442-1142, 6:00 a.m. to 4:00p.m. (Pacific), Monday - Friday for any further questions.*********************************************************************************************On May 18, 2015, Physio Control sent the Urgent Medical Correction-Action Required LIFEPAK 1000 DEFIBRILLATOR to U |  | 96611 | 96,611 units total (40,974 in the US and 55,637 outside the US) | Worldwide Distribution-USA (nationwide) and the countries of ALBANIA, ARGENTINA, AUSTRALIA, AUSTRIA, BAHRAIN, BANGLADESH, BARBADOS, BELGIUM, BOLIVIA, BOSNIA AND HERZOGOVENIA, BRAZIL, BRUNEI DARUSSALAM, CANADA, CANARY ISLANDS, CAYMAN ISLANDS, CHILE, China, COLOMBIA, COSTA RICA, CROATIA, Curacao, CYPRUS, CZECH REPUBLIC, DENMARK, ECUADOR, EGYPT, EspaÂża, FAROE ISLANDS, FINLAND, FRANCE, FRENCH GUIANA, FRENCH POLYNESIA, GABON, GEORGIA, GERMANY, GREECE, GUAM, GUATEMALA, HONG KONG, HUNGARY, ICELAND, INDIA, INDONESIA, IRAN, ISLAMIC REPUBLIC OF, IRAQ, IRELAND, ISRAEL, ITALY, Japan, JORDAN, KAZAKHSTAN, KENYA, KOREA, REPUBLIC OF, KUWAIT, LATVIA, LEBANON, LIBYA, LITHUANIA, MACAU, MACEDONIA, THE FORMER YUGOSL, MALAYSIA, MALDIVES, MALTA, MEXICO, MOROCCO, NAMIBIA, NETHERLANDS, NEW CALEDONIA, NEW ZEALAND, NORWAY, OMAN, PAKISTAN, PANAMA, PERU, PHILIPPINES, POLAND, PORTUGAL, QATAR, REUNION, ROMANIA, RUSSIA FEDERATION, SAN MARINO, SAUDI ARABIA, SERBIA, SINGAPORE, SLOVAKIA, SLOVENIA, SOUTH AFRICA, SPAIN, SRI LANKA, Suriname, South America, SWEDEN, SWITZERLAND, TAIWAN, THAILAND, TUNISIA, TURKEY, TURKMENISTAN, UKRAINE, UNITED ARAB EMIRATES, UNITED KINGDOM, URUGUAY, and VENEZUELA. | Open, Classified |  |  | 96611 | Computer
Z-1655-2014 | 67953 | No | JLW | Thyroid stimulating hormone test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | ADVIA Centaur Systems TSH3 Ultra 2500 tests |  |  | April 10, 2014 | 2014 | Siemens Healthcare Diagnostics, Inc | Samples containing fluorescein may show interference with the ADVIA CentaurÂż Systems TSH3 Ultra, Vitamin D and the ADVIA Centaur BRAHMS Procalcitonin assays. | Device Design | Siemens Healthcare Urgent Medical Device Correction (UMDC) was mailed to all affected Siemens Healthcare Diagnostics customers in the United States on April 11, 2014 for customer receipt on April 14, 2014. This notice informed the customers of fluorescein interference with the ADVIA Centaur Systems TSH3-Ultra and Vitamin D Total assays. This notice instructs customers that the TSH3-Ultra and Vitamin D assays can continue to be used for testing any sample that does not contain fluorescein. However, testing blood drawn from a patient, without allowing clearance of fluorescein is not recommended. This notice should be reviewed with the facility's laboratory or medical director.B. An Urgent Field Safety Notice (UFSN) was sent out for distribution to all customers outside the United States on April 10, 2014. This notice informed the customer of fluorescein interference with the ADVIA Centaur Systems TSH3-Ultra , Vitamin D Total and ADVIA Centaur BRAHMS Procalcitonin assays. This notice instructs customers that the TSH3-Ultra, Vitamin D andProcalcitonin assays can continue to be used for testing any sample that does not contain fluorescein. However, testing blood drawn from a patient, without allowing clearance of fluorescein is not recommended. This notice should be reviewed with the facilitys laboratory or medical director. (Note: The ADVIA Centaur BRAHMS Procalcitonin assay is not sold in the US.) |  | 8052 | 8,052 kits | Worldwide Distribution-USA (nationwide) and the countries of Afghanistan Algeria Armenia Austria Bahrain Belgium Bulgaria Burkina Faso Canary Islands Croatia Cyprus Czech Republic Denmark Estonia Finland France Germany Great Britain Greece Hungary Iceland Iran Iraq Ireland Israel Italy Jordan Kazakhstan Latvia Lebanon Lithuania Luxembourg Macedonia Malta Netherlands Norway Oman Pakistan Poland Portugal Qatar Romania Russian Fed. Saudi Arabia Serbia Slovakia Slovenia Spain Sweden Switzerland Turkey Unit.Arab Emir. Vatican Yemen. | Terminated |  |  | 273253 | Computer
Z-229-2014 | 67639 | No | DSQ |  |  | Cardiovascular | PMA | HeartMate II LVAD Pump and Pock Controller Kit; |  |  | March 04, 2014 | 2014 | Thoratec Corporation | Discovery of serious injuries and deaths associated with the process of changing from a primary System controller to their back-up System controller in patients using the Pocket System controller model. | Device Design | Thoratec sent an Urgent Medical Device Correction letter dated March 4, 2014 to all affected customers. The letter was sent via e-mail. All distributors outside of the USA and Canada to whom affected product has been distributed have been notified of this field correction notice and have been provided copies of the Customer Letter to forward to their customers. The letter identified the affected product , problem and actions to be taken. Customers were requested to complete and sign the attached Acknowledgement Form and return it to Thoratec via fax (925-847-8571) or e-mail a scanned copy tocustomer.service@thoratec.com. For questions call 1-800-528-2577. |  | 36 | 36 | Worldwide distribution: US (nationwide) including states of: AL, AR, AZ, CA, CO, CT, DE, FL, GA, IA, IL, IN, KY, LA, MA, MD, MI, MN, MO, MS, NC, NE, NJ, NY, OH, OK, OR, PA, SC, TN, TX, UT, VA, WA, and WI; and countries of: Austria, Belgium, Czech Republic, France, Germany, Ireland, Israel, Italy, Lebanon, Lithuania, Poland, Saudi Arabia, Singapore, Slovenia, Sweden, Switzerland, The Netherlands, Turkey, and United Kingdom. | Terminated | August 27, 2015 | 541 | 218533 | Computer
Z-1271-2014 | 67657 | No | LHN | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | MEVION S250, System, Radiation Therapy, Charged-Particle, Medical used for Proton Radiation Therapy |  |  | March 05, 2014 | 2014 | Mevion Medical Systems, Inc. | Geometric positioning error can occur after a rotational correction has been applied to one or more treatment fields | Software design | Mervion issued an Important Safety Notice letter on March 5, 2014 by visit to sites to the single customer using the product to prevent the use of multiple isocenters treatment modalities. The letter identified the affected product, problem and actions to be taken. For questions call 978-540-1555. |  | 1 | 1 unit | USA Nationwide Distribution in the state of MO | Terminated | April 03, 2014 | 29 | 1 | Computer
Z-1640-2014 | 67620 | No | KOI | Electrical peripheral nerve stimulator. | Anesthesiology | Anesthesiology | 510(k) | GE Healthcare, NeuroMuscular Transmission Module, E-NMT-01 |  |  | April 09, 2014 | 2014 | GE Healthcare, LLC | When the E-NMT-01 module is used in conjunction with the ElectroSensor, the Neuromuscular Transmission (NMT) values may indicate a deeper level of muscle relaxation than the actual level of muscle relaxation. In the clinical situation visual movements of the hand are seen after TOF (Train of Four) stimulation, but the patient monitor shows no counts, or counts are not corresponding to the actual a | Software design | GE Healthcare sent an "Urgent Medical Device Correction" letter (Ref 36099) dated April 9, 2014 to affected customers. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction, and Contact Information (contact Technical Support at 1-800-558-7044 or your local Service Representative).UPDATE: On May 11, 2015 a second GE Healthcare letter FMI 36099-2 was sent to OUS consignees. The updated report was for additional affected modules due to FRU installation. |  | 1794 | 1,794 (471 units US, 1,323 units OUS) | US Nationwide Distribution in the states of: AR, CA, CO,DC, FL, GA, IA, IL, IN, KS, KY, LA, MA, MD, MI, MN, MI, MN, MO, NC, NV, NY, OH, OK, TN, TX, UT, WA, WI, WY.OUS: ARGENTINA. AUSTRALIA, AUSTRIA, BELGIUM, BOLIVIA, CANADA, CHILE, COLOMBIA, COSTA RICA, CYPRUS, CZECH REPUBLIC, DENMARK, ECUADOR, EGYPT, ESTONIA, FINLAND, FRANCE, GERMANY, GREECE, HONG KONG, HUNGRY, INDIA, IRA, ISRAEL, ITALY, JAPAN, KAZAKTSTAN, KUWAIT, LATVIA, LITHUANIA, MALAYSIA, MEXICO, NETHERLANDS, NEW ZEALAND, NORWAY, OMAN, PANAMA, POLAND, QATAR, ROMANIA, RUSSIAN, FÂżDÂżRATION, SAUDI ARABIA, SINGAPORE, SLOVAKIA, SPAIN, SWEDEN, SWITZERLAND, TURKEY, TAIWAN, UNITED ARAB EMIRATES, UNITED KINGDOM, URUGUAY, VENEZUELA. | Terminated | September 28, 2015 | 537 | 1794 | Computer
Z-1645-2014 | 67957 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | 550 TxT Treatment Table of the Digital Linear Accelerator (LINAC) |  |  | April 01, 2014 | 2014 | Siemens Medical Solutions USA, Inc | Table may lose calibration during patient treatment, which may cause the possibility of mistreatment. | Software design | Siemens sent an Medical Device Correction, Field Safety Notice of Improvement to End users in March 2014, informing them that a Siemens representative would install a software update.By installation of this software update the issues were solved by the following actions:1. A plausibility check was introduced for comparing the two independent encoder systems of each table axis.2. The time between two encoder comparisons was reduced to 50ms and an error code was introduced.Please include this Field Safety Notice in your Digital Linear Accelerator System Owner Manual chapter "Safety Advisory Letters" where it should remain.The relevant National Competent Authority has been informed of this update.We regret any inconvenience that this may cause, and we thank you in advance for your understanding. Further questions please call (610) 219-6300. |  | 120 | 120 | USA ( nationwide) Distribution. | Terminated | December 11, 2014 | 254 | 120 | Computer
Z-1272-2014 | 67556 | No | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | AXIOM Artis Systems |  |  | February 07, 2014 | 2014 | Siemens Medical Solutions USA, Inc | Certain units of the AXIOM Artis Systems are using a flat detector cooling unit with inappropriate connectors that may possibly leak, causing a system failure and possible drop out of the gantry system. | Device Design | Siemens sent a recall notification letter, dated February 7, 2014, to all affected customers. The letter identified the affected product, problem and actions to be taken. For questions call +49 9191 18 0. |  | 37 | 37 | USA Nationwide Distribution in the states of WI, KY, OH, NJ, NC, FL, PA, MA, IA, TX, NY, IL, MI, AZ, CA, and LA. | Terminated | December 30, 2014 | 326 | 37 | Computer
Z-1646-2014 | 68262 | No | OUG | Medical device data system. | General Hospital | General Hospital | 510(K) Exempt | Philips XRE, the dictation editor system included in the Philips XIRIS 8.1.50 |  |  | May 05, 2014 | 2014 | Philips Healthcare Informatics, Inc. | When using XIRIS Rio version (8.1.50.0) in conjuncture with iSite PACS a synchronization error can occur and if not noticed, information for the incorrect patient may be displayed | Software design | Philips Healthcare sent an Urgent Field Safety Notice dated May 5, 2014. to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Customers were advised of the synchronization problem and request that a copy of this notification be kept with the equipment instructions for use. The document goes on to outline the affected product, the problems and conditions when it can occur. A photo of the label for easy identity the actions to be taken are provided. Philips is informing users of the issue via the notification and will deploy a software correction to all customers when available.Questions may be directed to Philips reps or Care at 1-866-946-4701 or OneEMSProd@philips.com |  | 5 | 5 units | Worldwide Distribution - US including the states of FL and TX and the countries Canada and South Africa. | Terminated | April 28, 2015 | 358 | 5 | Computer
Z-1647-2014 | 67787 | No | MKW | Cyclosporine test system. | Clinical Chemistry | Toxicology | 510(k) | Dimension VistaÂż CSA and CSAE |  |  | March 19, 2014 | 2014 | Siemens Healthcare Diagnostics, Inc. | A complaint was received of Dimension VistaÂż CSA and CSAE under-recovery of cyclosporine when patient samples are run from the Small Sample Container. The under-recovery that has been observed is 18 to 44%. In March 2014, Siemens recalled due to the complaint received of Dimension VistaÂż CSA and CSAE reporting under-recovery of cyclosporine when patient samples are run from theSmall Sample C | Device Design | Siemens sent an Urgent Medical Device Correction letter dated March 2014. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed not to use the SSC with patient samples when running CSE or CSAE.Siemens also recommended a look back at previously generated CSA or CSAE results if the patient samples have been run from the SSC and that the content of this letter should be discussed with the Medical Director. Customers were asked to complete the attached form and fax it to (302) 631-8467.For questions regarding this recall call 800-441-9250. |  | 4175 | 4175 | Worldwide Distribution - USA (nationwide) and Internationally to Canada, Belgium, Switzerland, Germany, Spain, France, Great Britain, Italy, Saudi Arabia, Australia, and Japan. | Terminated | August 10, 2015 | 509 | 4175 | Computer
Z-1275-2014 | 67615 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | DICOM Export Format for Sum Dose with Velocity v3.0.0. |  |  | July 25, 2013 | 2013 | Velocity Medical Solutions, LLC | An anomaly was detected such that, under certain conditions, a summed dose created by Velocity replaces the calculated dose from an approved plan in Eclipse when importing the dose from Velocity. | Software design | Velocity Medical Solutions sent a Urgent Field Safety Notice dated July 25th, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Velocity advises customers creating Dose volumes in Velocity to refrain from exporting these dose files to Eclipse until the customers install has been upgraded to Velocity 3.0.1, which will be available August 5th, 2013.A new version of Velocity that corrects this error, Velocity v3.0.1, will be available on August 5th, 2013 for upgrade at no cost. Please contact your Velocity Medical representative to obtain this upgrade.Service representative will upgrade the application at your earliest convenience.Please advise the appropriate personnel working in your medical department of the content of this letter.We sincerely apologize for any inconvenience and thank you in advance for your co-operation.Contact Information. Helpdesk: (404) 920-1966 or (855) 857-0802 and press 2. |  | 54 | 54 installs | Worldwide Distribution - US (nationwide) including the states of AL, CA, CO, CT, DC, FL, GA, IL, IN, LA, MA, MD, MI, MS, NC, NY, PA, SC, TN, TX, WI, and the countries of Belarus, Italy and Switzerland. | Terminated | September 11, 2014 | 413 | 54 | Computer
Z-1539-2014 | 68011 | No | BZO | Ventilator tubing. | Anesthesiology | Anesthesiology | 510(K) Exempt | Infant Nasal CPAP (continuous positive airway pressure) Prong, Nasal. |  |  | April 15, 2014 | 2014 | Fisher &amp; Paykel Healthcare, Ltd. | The Infant Nasal CPAP Prong has the potential to detach from the Nasal Tubing and therapy is likely to be interrupted. | Device Design | Fisher &amp; Paykel Healthcare sent an Urgent Medical Device Recall letter on April 11, 2014, to all affected customers. The letter informed them of the problems identified and the actions to be taken. Customers were instructed to complete Section A 'Inspection of Stock' on the Urgent Medical Device Recall Response Form and return it to their FPH Representative. Destroy the Affected Prongs by cutting them in half. Customers were also instructed to contact their FRH Representative for replacement product. For questions regarding this recall call 949-453-4000, ext 1423. |  | 62063 | 62,063 units total (14,420 units USA) | Worldwide Distribution - USA (nationwide) and Internationally to:ARGENTINAAUSTRALIABANGLADESHBELGIUMCANADACHILECHINAFRANCEGERMANYGREAT BRITAINGREECEHUNGARYINDIAINDONESIAIRAQISRAELITALYJAPANLIBYAMALAYSIANEPALNETHERLANDSNEW ZEALANDPERUPOLANDPORTUGALQATARROMANIARUSSIARWANDASAUDI ARABIASOUTH AFRICASPAINSRI LANKASWITZERLANDTAIWANTHAILANDTRINIDAD &amp; TOBAGOTURKEYUNITED ARAB EMIRATES | Terminated | September 30, 2014 | 168 | 62063 | Computer
Z-1297-2014 | 67672 | No | MIH |  |  | Cardiovascular | PMA | Ovation Prime Abdominal Stent Graft System; |  |  | February 27, 2014 | 2014 | Trivascular, Inc | Communication to European physician users of the Ovation Prime Abdominal Stent Graft system to notify them of changes to Instructions for Use. | Device Design | Trivascular sent an Important Medical Device Information letter dated February 27, 2014, to all European physicians who have scheduled cases through March 31, 2014 or until IFU;s are translated and available for European physicians. The letter provided notification of an update that will appear in a future revision of the product Instructions for Use (IFU). The IFU will be available in their country when the required translations and product registration/notification have been completed. This IFU does not require rework or return of the affected product. Customers with questions were instructed to contact their local TriVascular representative.For questions regarding this recall call 707-543-8805. |  | 45 | 45 | Worldwide Distribution including Austria, Belgium, Czech Republic, France, Germany, Greece, Hungary, Italy, Luxembourg, Norway, Poland, Spain, and Sweden. | Terminated | April 11, 2014 | 43 | 45 | Computer
Z-1638-2014 | 68224 | No | MUE | Full-field digital mammography system. | Radiology | Radiology | 510(k) | MAMMOMAT Inspiration. |  |  | April 25, 2014 | 2014 | Siemens Medical Solutions USA, Inc | There is a potential and possible hazard to the user when using the MAMMOMAT Inspiration PC monitor at the control desk, in that the holder of the PC monitor can break causing an unstable monitor to fall causing possible serious injury. | Device Design | A safety advisory notice, dated April 25, 2014, was sent to direct accounts to notify them of the issue and provide instructions to avoid the potential risk of injury. |  | 5 | 5 | Distributed in the states of GA, FL, NY, NV, and MA. | Terminated | December 15, 2014 | 234 | 5 | Computer
Z-1202-2014 | 67016 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Siemens Artis Q and Artis Q biplane systems |  |  | December 05, 2013 | 2013 | Siemens Medical Solutions USA, Inc | There exists the possibility that during clinical operation, a breakdown of the x-ray tube could result in loss of x-ray imaging for the Artis Q and Artis Q biplane systems with a flat detector as 40 HDR (large 30x40). | Device Design | Siemens sent an Important Customer Safety Notice letter dated November 22, 2013, to affected customers. The letter identifed the affected product, problem and actions to be taken. Customers were instructed to discontinue use until the affected product has been successfully replaced by a Siemen's representative.at no cost. |  | 19 | 19 | USA Nationwide Distribution in the states of TN, MA, CA, IL, OH, FL, UT, WA, GA, IA, NY, MO, IN, and CO. | Terminated | November 06, 2014 | 336 | 19 | Computer
Z-1307-2014 | 67504 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | RayStation Radiation Therapy Treatment Planning System, models 2.5, 3.0, 3.5 and 4.0; a software system designed for treatment planning and analysis of radiation therapy. |  |  | February 06, 2014 | 2014 | RAYSEARCH LABORATORIES AB | Potential for dose errors due to software program errors. | Software Design Change | RaySearch Laboratories sent an Field Safety Notice, Medical Device Correction letter dated February 4, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customers.ACTIONS TO BE TAKEN BY THE USERA workaround that eliminates the described error is to select the discard rotations option in the Image Registration Module before creating a rigid registration intended to be used for hybrid deformable registration. Another option is to use the ROI based deformable registration when applicable, as this functionality is not affected by the error.SOLUTIONThis issue will be resolved in a patch release of RayStation, scheduled for market release February 28, 2014. Release schedule is subject to regulatory approval for some markets. In the meantime, this field safety notice is being distributed to all customers.CONTACT INFORMATIONFor further information, please contactFor the Americas:Customer Support, RaySearch Americas, at +1 877 778 3849 For Europe, Asia and rest of the world:RaySearch Laboratories AB, at +46 8 5450 6130 Thank you for your cooperation, and we apologize for any inconvenience |  | 151 | 151 units Total (88 units domestically &amp; 63 units internationally) | Worldwide Distribution - US including the states of NY, OH, MD, CT, CO, TN, IL, AZ, MS, NJ, PA, DE, FL, CA and WA., and the countries of New Zealand, Australia, Republic of Korea, China, Italy, Germany, Norway, Belgium, Switzerland, United Kingdom and Spain. | Completed |  |  | 151 | Computer
Z-1310-2014 | 67628 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | TDWorkstation version 11.01.A or higher. TDWorkstation Evolution. |  |  | February 17, 2014 | 2014 | Technidata S.A. | If a conflict occurs between a Demography Query and a Result Transmission by the instrument, the TDWorkstation could send an incorrect patient demography to an instrument. | Software design | An email dated 2/17/14 was sent to their customers to inform them about concerns only if their site uses TDWorkstation and TD-Workstation Evolution products that meet all the listed criterias. The criterias are listed in the customer notification letter and it informs the customers of the problems identified and the actions to be taken. Customers are instructed to complete the reply form and return it by fax (514) 270-2770 with in 15 days from receiving the customer notification form. Customers that do not respond the reply form, Technidata will consider those customers do not wish to install the correction. |  | 85 | 85 units total (1 unit in US) | Distributed in the state of PA. | Terminated | March 31, 2014 | 42 | 85 | Computer
Z-1567-2014 | 68066 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Alaris Pump model 8100 with software version 9.1.18 and software upgrade kits with "Delay Until" Option and "Multidose" Feature |  |  | April 23, 2014 | 2014 | CareFusion 303, Inc. | CareFusion is recalling the Alaris Pump model 8100 version 9.1.18 because it may have a software issue that results in situation where the pump module will not properly delay an infusion when the "Delay Until" option or "Multidose" feature is used. | Software design | CareFusion sent an Urgent Medical Device Recall Notification letter dated April 23, 2014, to all affected customers. The letter informed the customers of the problems identified and the actions to be taken. Customers were informed that if the infusion starts earlier or later than intended and is not immediately detected and stopped by the clinician this could result in serious injury or death. Customers are informed although no adverse events or deaths have been reported there is a potential for this risk. If customers experience the issue while using the Alaris Pump module then they are instructed to contact CareFusion Customer Advocacy at (888) 812-3266. Customers are instructed to promptly complete and return the enclosed Customer Response Card to expedite the corrective action process.For questions regarding this recall call 858-617-4000. |  | 57219 | 57,219 units total (55,756 units in US) | Worldwide Distribution-USA (nationwide) and the countries of Canada, AE, AU, and KW. | Terminated | March 01, 2017 | 1043 | 57219 | Computer
Z-1617-2014 | 68208 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | McKesson Radiology-PACS. |  |  | April 17, 2014 | 2014 | Mckesson Medical Immaging | High priority alert icon was not displayed in the Study List when there are no Stat studies present. However, the High priority alert icon is visible when studies marked as Stat priority exist in the Study List. | Software Design Change | On April 17, 2014, McKesson Medical Imaging Company initiated a process for applying a software solution to all McKesson Radiology customer systems that may be affected to prevent future occurrences of this issue. The systems with the potentially affected software versions will be corrected without removing the systems from their location. |  | 45 | 45 | Worldwide Distribution-USA (nationwide) including the states of AL, CA, GA, ID, IL, IA, MD, MS, MO, NE, NV, NH, NY, NC, PA, TN, TX, VT, WA, WV, WY, and the countries of Canada and France. | Terminated | February 02, 2016 | 656 | 45 | Computer
Z-1309-2014 | 67558 | No | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | Artis Zee Angiographic x-ray systems. |  |  | February 07, 2014 | 2014 | Siemens Medical Solutions USA, Inc | An inappropriate cable connection was assembled for certain Artis zee systems. which may cause a system failure and possibly lead to a restricted maneuverability of the gantry system. | Device Design | A recall notification letter dated February 7, 2014, was sent to End Users. |  | 9 | 9 | MN, WI, PA, MI, AL, OH, NC, AR, and NJ. | Terminated | March 16, 2015 | 402 | 9 | Computer
Z-1619-2014 | 68042 | Yes | KDI | High permeability hemodialysis system. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Revision K Actuator Test Boards sold as replacement parts for use with Fresenius 2008-Series Hemodialysis Machines (2008K, 2008K2, 2008K@home and 2008T). |  |  | April 25, 2014 | 2014 | Fresenius Medical Care Holdings, Inc. | 2008-Series Hemodialysis Machines with Revision K Actuator Test Boards may fail to detect a disconnected Ultrafiltration (UF) Pump | Device Design | Fresenius Medical issued Urgent Recall letter issued on 4/25/14 by certified mail with signature confirmation and fax-back form. Customers instructed to examine their inventory to determine whether they have any of the affected 2008T machines or if they have installed any of the affected replacement boards. If customers have the affected machines or replacement boards, they will be instructed to contact FMC-RTG Technical Service to have the Actuator Test Boards replaced. Contact the FMCNA Technical Service Team at 1-800-227-2572. |  | 39 | 39 units | USA (nationwide) and the country of Canada.* | Open, Classified |  |  | 118 | Computer
Z-1308-2014 | 67511 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | RaySearch Ray Station, Radiation Therapy Treatment Planning System, Stand-alone software |  |  | February 06, 2014 | 2014 | RAYSEARCH LABORATORIES AB | RaySearch Laboratories has recalled RaySearch Ray Station 4.0 software due to clinical dose calculation errors during radiation therapy. | Software Design Change | On February 6, 2014, RaySearch Laboratories distributed Urgent Medical Device Correction notices to their customers via email regarding the affected product. To determine if the version you are using is affected, bring up the About box in the RayStation application and check if the number reported is "4.0.2.9". If so, the notice applies to your version. Customers are advised to inform all planning staff and all users about the instructions listed on the notice. Customers are advised to inspect their system and identify all installed units with the referenced software version numbers listed on the notice. Corrective labeling is distributed to the customers in the form of the safety notification attached to this report and the software problem will be resolved in RayStation patch 4.0.3 scheduled for market release end of February 2014 and in all future releases. For further information and return of the reply form, please contact: Freddie Cardel, Director of Customer Support, RaySearch Americas, (877) 778-3849 or freddie.cardel@raysearchlabs.com. |  | 19 | 19 units (16 units domestically &amp; 3 units internationally) | OH, MD, NJ, WA, IL, MA, FL, Spain, United Kingdom, Italy, | Open, Classified |  |  | 19 | Computer
Z-1613-2014 | 67960 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Syngo Dynamics v9.5 system. A Picture Archiving and Communication System (PACS) intended for acceptance, transfer, display, storage, archive and manipulation of digital medical images, including quantification and report generation. |  |  | April 03, 2014 | 2014 | Siemens Medical Solutions USA, Inc | Systems with the affected serial number are experiencing a software error which may lead to incorrect patient demographics display. | Software design | Siemens sent a Recall Notification letter dated March 26, 2014 to affected customers. The letter identified the affected product, problem and actions to be taken. Siemens is preparing a Service Patch modification that would be performed to resolve the software issue. |  | 50 | 50 | US Nationwide Distribution in the states of AR, CA, CT, GA, IA, IL, KS, MA, MI, MN, NJ, NM, NY, OH, OR, PA, TX, TN, TX, WA, WI, and WV. | Terminated | December 15, 2014 | 256 | 50 | Computer
Z-1553-2014 | 67603 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | ADVIA Centaur XP Immunoassay system, an automated in vitro diagnostic analyzer. Catalog # 078-A011-03 - ADVIA Centaur XP, Catalog # 078-A011R03 - ADVIA Centaur XP refurb. |  |  | February 18, 2014 | 2014 | Siemens Healthcare Diagnostics | Instrument Cover Gas Spring failures. Over time, the gas spring may lose its effectiveness and fail to support the cover in its full or partially open position. This may lead to the cover falling. | Device Design | On February 18, 2014 Siemens Healthcare Diagnostics distributed Urgent Medical Device Correction notices to their customers via courier service (Urgent Medical Device Correction/). Upon receipt of the communication, customers are asked to follow "Actions to be Taken by the Customer". If the cover is unable to stay partially raised, immediately contact your Siemens technical support representative to arrange for a service. Complete and return the Field Correction Effectiveness Check attached to this letter within thirty (30) days by faxing to (302) 631-7597. Customers with questions should contact their Siemens technical support representative. |  | 5900 | 5,900 units total (1,679 domestically &amp; 4,221 internationally) | Nationwide and internationally | Open, Classified |  |  | 5900 | Computer
Z-1311-2014 | 67749 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | Zoll Medical X Series, PropaqMD, PropaqM Defibrillator/Pacemaker/Monitor With System Software Version 02.10.02.00 or Higher, automated external defibrillator. |  |  | March 05, 2014 | 2014 | ZOLL Medical Corporation | Defibrillator/Pacemaker/Monitor May Latch in a Continuous Device Reset Loop and cannot be used | Software design | Zoll Medical issued a Notification letter daetd March 5, 2014, to all affected consignees. The notification letter provides information on how to avoid the issue and instructs customers to contact ZOLL to get affected devices updated. Customers will be given the option to have a software upgrade kit sent to them to update their devices or return devices to ZOLL Service department in Chelmsford Massachusetts to have the device upgraded with the latest software release (Software version 02.16.04.00). Customers with questions were instructed to contact the 24/7 technical support numbers 1 (800) 348 9011 or +1 (978) 421 9460 to assist users with any aspect of this notice. |  | 3169 | 3,169 units | Worldwide Distribution - USA (nationwide) and Internationally to Canada, AE, Afganistian, Australia, Belgium, Czech Republic , China, Chile, Croatia, Denmark, France, Great Britain, Hungary, Israel, Italy, Japan, Korea, Kuwait, Russia, Saudi Arabia, Singapore, South Africa,Thailand, and Turkey. | Open, Classified |  |  | 3169 | Computer
Z-1616-2014 | 67961 | No | NBW | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | ACCU-CHEK Compact Plus Test Strips |  |  | April 14, 2014 | 2014 | Roche Diagnostics Operations, Inc. | Roche Diabetes Care has become aware the ACCU-CHEK Compact Plus test strips may produce erroneously low blood glucose readings in patients undergoing Ceftriaxone therapy. The interference is not described in the product labeling. | Device Design | On 4/14/14, Roche Diagnostics Operations issued URGENT MEDICAL DEVICE CORRECTION notification to consignees which included description of the affected product and issue, clinical significance, and actions required. Consignees are directed to call the ACCU-CHEKÂż Customer Care Service Center at 1-800-858-8072 for any questions about the information contained in the notification. |  | 590136 | 590,136 boxes | Nationwide Distribution. | Terminated | September 18, 2014 | 157 | 590136 | Computer
Z-1566-2014 | 67884 | No | NEP | Intravascular administration set. | General Hospital | General Hospital | 510(K) Exempt | ABACUS TPN Calculation Software, Product Codes: 8300-0045 ABACUS CE (Calculator Edition), 8300-0046 ABACUS SE (Single-Workstation Edition), 8300-0047 ABACUS ME (Multi-Workstation Edition), 8300-0156 ABACUS 3.0 CE, 8300-0157 ABACUS 3.0 SE, 8300-0158 ABACUS 3.0 ME, 8300-0166 ABACUS 3.1 DE, 8300-0167 ABACUS 3.1 CE, 8300-0168 ABACUS 3.1 SE, 8300-0169 ABACUS 3.1 ME. ABACUS is a Windows - based order entry TPN software application for comprehensive calculations and label printing. The primary use of Abacus is for Total Parenteral Nutrition (TPN) order calculations. |  |  | April 02, 2014 | 2014 | Baxter Corporation Englewood | 1. ABACUS v3.1 may calculate quantities of electrolytes that are double the expected values during the creation of TPN orders.2. ABACUS v3.1 may automatically add additional sterile water to a formula equal to the volume of a premixresulting in an over dilution.3. All software versions of ABACUS software display the calcium phosphate curve points for Premasolincorrectly.4. All software v | Software design | Baxter sent an Urgent Device Correction letter dated April 21, 2014, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. . 1. Follow safe compounding practice, which includes 1) using of the Summary button to verify the order against the calculated amounts prior to completing the order, 2) verifying that the ordered ingredients and quantities displayed in the software and printed on the Bag label and the Solution Formula label match the PN prescription prior to preparation, and 2) using a filter for administration of a PN bag. Use of the ABACUS Software must comply with your facility's policies, procedures, and training, including verification that all information entered into the ABACUS Software is accurate and clinically appropriate.2. To determine if any of these Issues are present in your current ABACUSconfiguration, please note the following immediate actions should be taken:"Double Salts: Baxter contacted all facilities installed by Baxter with ABACUS v3.1 to correct the Double Salts Issue via telephone between April 4, and April 16, 2014.To avoid the Double Salts Issue, a pharmacy installed with ABACUS v3.1 should 1) ensure the "Include non-electrolyte ion contributions in balance" checkbox is checked for all salt-based templates, and 2) compare the ordered amount of salts to the delivered amount on the Ions tab in ABACUS."Sterile Water Added to TPN with Pre-Mix: Customer should ensure that all premix not intended for use on the compounder are marked as "Pre-Mix PN Solution" type in the Formulary. Customer should then check the template, under the "Infusion" tab, to ensure the "Infusion Mode" is set to "Pre-Mix (Cyclic)"."Premasol CaP04 Curve Data Points: Review Attachment 1 for information relevant to the use of calcium phosphate curves for Premasol 3% Dex 25 |  | 2034 | 2034 | Worldwide Distribution - USA nationwide, Canada, Colombia, Dominican Republic, Panama, Austria, Bosnia Herzegovina, Bulgaria, Bahrain, Estonia, France, Germany, Georgia, Hungary, Ireland, Israel, Jordan, Kuwait, Latvia, Morocco, Oman, Qatar, Saudi Arabia, Slovenia, Slovakia, South Africa, Switzerland, United Arab Emirates, United Kingdom, Hong Kong, and Taiwan. | Terminated | January 20, 2016 | 658 | 2034 | Computer
Z-1870-2014 | 68531 | No | NAY | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | IS4000 da Vinci Xi - Endoscope accessory; |  |  | June 11, 2014 | 2014 | Intuitive Surgical, Inc. | In some procedures when using the endoscope in conjunction with the IS4000 (da Vinci Xi) System, the surgical view has an orange-yellow tint which can make it hard to differentiate between tissue | Device Design | Medical Device Recall notifications were sent to the 4 affected sites on June 11, 2014. Letters describe the problem and the potential health risk. The two model numbers of the endoscope are listed and the statement that ISI will perform a software upgrade to the da Vinci System and replace the endoscopes.Letters request that the following actions take place:- Ensure that all appropriate personnel are fully informed of the information provided in the notification letter.- Complete and return the acknowledgement form using instructions provided.- Retain a copy of The letter for your records.Intuitive representatives will schedule a site visit to provide the re-calibrated endoscope, update the software and ensure that it is properly integrated with the da Vinci Xi Surgical System.Questions or concerns may be directed to Intuitive Customer Service at 800-876-1310. |  | 20 | 20 units | Distributed in the states of FL, TX, and NY. | Terminated | July 02, 2014 | 21 | 39 | Computer
Z-1148-2014 | 66119 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | GE Healthcare, Model Number/ Description: |  |  | March 25, 2011 | 2011 | GE Healthcare, LLC | When performing a secondary capture generated from GSI Viewer, the header information is correcton the images. However, when changing the ROI (region of interest) and zooming in, the DICOMheader does not change to reflect the new image, thus the line measurements may be inaccurate. | Software design | No consignee letter was sent out. FMI No. 25431 shows where the Field Engineer had to communicate with the hospital (Instead of them sending a letter).GE Healthcare has updated all of the affected Advantage Workstation systems affected by the GSI viewer issue. At this time, deployment of the recall is considered complete. As an effectiveness check, they have tracked all service requests for the field action to completion.These service requests record the repair of the affected units. Exceptions to completion of the service request have been recorded as well. Records of field action completion will be retained per their processes and procedures. |  | 283 | 283 (59 US, 224 OUS) | Worldwide Distribution: US (nationwide): AL, AZ, CA, CT, FL, IL, IN, IA, KY, MD, MA, MI, MS, NV, NH, NJ, NY, PA, WA, and WI; and countries of: AUSTRALIA, BELGIUM, BRAZIL, CANADA, CHINA, DENMARK, FINLAND, FRANCE, GERMANY, GREECE, HONG KONG, INDIA, INDONESIA, ISRAEL, ITALY, JAPAN, KUWAIT, LATVIA, NEW ZEALAND, POLAND, SAUDI ARABIA, SINGAPORE, SOUTH AFRICA, SOUTH KOREA, SWEDEN,, SWITZERLAND, TAIWAN, and UNITED KINGDOM. | Terminated | March 13, 2014 | 1084 | 283 | Computer
Z-0910-2014 | 67152 | No | NBW | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | MooreMedical Supporting Health &amp; Care...MooreBrand Blood Glucose Monitor for Self-Testing &amp; Multi-Patient Use TRUEtrack. Manufactured by: NIPRO DIAGNOSTICS, Fort Lauderdale, FL 33309 U.S.A. |  |  | October 02, 2013 | 2013 | Nipro Diagnostics, Inc. | Manufactured with an incorrect factory set unit of measure. | Software design (manufacturing process) | Nipro Diagnostics sent an Urgent Medical Device Voluntary Recall letter dated December 31, 2013 to all affected customers. The customers were instructed to do the following: 1) Screen their inventory immediately for any of the serial numbers listed on the back of the products. 2) If they have any of the affected products, return the product by contacting Stericycle at 1-866-236-4518. 3) Complete and return the enclosed response form via fax at 1-888-358-7294. Customers with questions were instructed to call Stericycle at 1-866-236-4518.For questions regarding this recall call 954-677-9201. |  | 105 | 105 | Worldwide Distribution - USA (nationwide, HI and Puerto Rico and Internationally to Canada, European Union, Jamaica, Suriname, Trinidad and Tobago, Kazakhastan, Russian, Turkey, Australia, Bangladesh, Pakistan, Cambodia, Laos, Myanmar, Thailand, Vietnam, and India. | Terminated | January 28, 2016 | 848 | 606 | Computer
Z-2234-2014 | 68802 | No | MWI | Cardiac monitor (including cardiotachometer and rate alarm). | Cardiovascular | Cardiovascular | 510(k) | Vital Sync" VPMP &amp; IM, software, V2.4. (DVD), packaged on a CD or other electronically controlled device. CDs are packaged in clam shell jeweled cases within a carton box. |  |  | July 17, 2014 | 2014 | Nellcor Puritan Bennett Inc. (dba Covidien LP) | Covidien is conducting a field correction of the Vital Sync Virtual Patient Monitoring Platform and Informatics Manager due to a mismatch between the alarm priority on the medical devices and the Vital Sync system. | Software design | The firm, Covidien, sent an "URGENT FIELD CORRECTIVE ACTION NOTICE" letter dated 7/16/14 to its customers on 7/17/14 via FedEx. The letter informs users that a representative will be in contact to schedule a software update.The letter also described the product, problem and actions to be taken. The customers were instructed to disregard the audible and visual alarm priority indicators and only use the alarm description text on the patient tile and on the patient detail page for alarm priority information, until the mismatch is corrected. The Vital Sync" Virtual Patient Monitoring Platform and Informatics Manager or Vital Sync" Remote Monitoring System may still be used until the software update is available.If you have questions about this letter or to report any issues with the Vital Sync" Virtual Patient Monitoring Platform and Informatics Manager or Vital Sync" Remote Monitoring System, contact our Technical Support Department at (800)635-5267, option 3, then option 1. |  | 3 | 3 | US distribution to states of: LA, IL, MS, CO, IN, GA, CA and TX. | Terminated | October 21, 2014 | 96 | 12 | Computer
Z-2298-2014 | 68817 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Persona (TM) The Personalized Knee System |  |  | June 10, 2013 | 2013 | Zimmer, Inc. | Zimmer had received complaints that reported breakage of the Persona Knee Trial Articular Surface Provisionals (TASPs). This is a retrospective recall which covers Zimmer's corrective actions taken earlier this year concerning the TASPs. | Device Design | On February 10, 2014, Zimmer sent letter to surgeons who had filed complaints regarding the affected TASP fractures. The letter provided the surgeons with Zimmer's investigational findings of the issue. Surgeons were directed to send questions or any additional informtion to Zimmer.PER@Zimmer.com. The Zimmer Weekly Wrap-Up, dated February 14, 2014, was issued electronically to sales force. They were notified to engage in mandatory training to improve the usage of the TASPs and prevent potential breakage. Zimmer has taken additional corrective measures by initiating another recall in July 2014 whereby they notify the distributors as well as all users of the enhanced surgical technique. |  | 70986 | 70,986 distribution events | Worldwide Distribution-USA (nationwide) including DC and Puerto Rico and the states of AK, AL, AZ, CA, CO, CT, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, MA, MD, ME, MI, MN, MO, MS, NC, ND, NE, NJ, NM, NV, NY, OH, OK, OR, PA, SD, TN, TX, UT, VA, WA, WI, WY, and the countries of Australia, Canada, Germany, India, Japan, Korea, and Singapore. | Terminated | September 11, 2014 | 458 | 70986 | Computer
Z-0949-2014 | 65214 | No | LKM | Automated cell counter. | Hematology | Pathology | 510(k) | The UF-1000i is a fully automated urine particle analyzer intended for vitro diagnostic use in urinalysis. The instrument can perform screening for abnormal samples with a high degree of accuracy, so it increases automation and efficiency in laboratories. |  |  | March 31, 2013 | 2013 | Sysmex America, Inc. | The current Century Break patch did not contain the same setting for the Century Break year for both the Red Hat Linux and GUI files. Due to the mismatch in file settings, the software could not properly calculate a purge date for the database files based on the customer data retention setting in the application. This resulted in the database files meeting its maximum capacity and causing system p | Software design | A Product Notification letter dated July 2012, was sent to Siemens Healthcare Diagnostics. The letter suggested that an attached procedure should be followed for an server that requires the Century Break parameter change. If Siemens Healthcare Diagnostics has any questions about the procedure, the firm's representatives can contact the Sysmex Technical Assistance Center at 1-888-879-7639. A response form was also included with the letter requesting the Siemens Healthcare Diagnostics complete and return it. A second Product Notification letter dated September 2012, was sent to Siemens Healthcare Diagnostics. The letter contained largely the same instructions as the July 2012, letter. A third Product Notification letter dated March 2013, was sent to Siemens Healthcare Diagnostics. The letter include substantially similar instructions as the pervious letters; however, the procedure for changing the Century Break parameter was expanded. |  | 214 | 214 units | Worldwide Distribution - USA (nationwide) Distribution including Puerto Rico and the states of AL, AR, CA, CO, CT, FL, GA, HI, IA, IS, IL, IN, KY, LA, MA, MD, MI, MN, MO, MS, MT, NC, NE, NH, NJ, NM, NV, NY, OH, OK, PA, SC, SD, TN, TX, VA, WA and WI and the country of Canada. | Open, Classified |  |  | 214 | Computer
Z-2148-2014 | 68071 | No | MQN | Bone plate. | Dental | Dental | 510(k) | DePuy Synthes Craniomaxillofacial Distraction System (AB Distractor Bodies and BC Distractor Bodies) |  |  | April 16, 2014 | 2014 | Synthes, Inc. | DePuy Synthes is initiating a recall of certain lots of the Craniomaxillofacial Distraction System (AB Distractor Bodies and BC Distractor Bodies) because they may reverse post-operatively. | Device Design | The firm, DePuy Synthes, sent "Urgent Notice: Medical Device Recall" letters dated April 16, 2014 to affected customers. The letter identified the affected product, reason for recall, potential patient impact, mitigation steps for patients with distractors currently implanted and actions to be taken. Customers were instructed to review inventory, immediately remove the affected lots from stock, complete verification section of the letter and return the affected product. For questions call 610-719-5450 or contact your DePuy Synthes Sales consultant. |  | 1499 | 1499 | US Nationwide Distribution | Terminated | September 09, 2015 | 511 | 1499 | Computer
Z-2288-2014 | 68653 | No | FSM | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | PERSONA (TM) The Personalized Knee System Tibial General Instrument Tray, non-sterile |  |  | July 01, 2014 | 2014 | Zimmer, Inc. | The firm discovered that the current Personal Tibial/General Fixed Bracket Trays may not have the specific geometry to accommodate the tibial drill and stop guide with proper position/orientation. The drill bracket and drill guide are not the correct geometry to allow for a proper fit in the Persona Fixed Case Kit which could lead to improper sterilization of the tibial drill and stop guide. | Device Design | On July 1, 2014, Zimmer, Inc. issued recall notification to their consignees via e-mail and/or letter. Notification included affected product issue and instructions for returning the affected product. All distributors were notified via electronic mail. Hospital risk managers, as well as distributors with affected inventory were also notified via courier. Distributors were sent a letter identifying the issue and their responsibilities. These responsibilities include locating and removing the affected product in their territory. Distributors are responsible for returning on-hand affected products to Zimmer and ensuring all of their products are accounted for using the form provided in the letter. Hospital risk managers were provided with a letter identifying the issue and their responsibilities. These responsibilities include locating the affected product, quarantining the product, and returning it to their Zimmer sales representative. |  | 190 | 190 units | Worldwide Distribution - USA including AZ, CA, CO, FL, IL, IN, KS, MI, MN, NC, NJ, NY, OH, OR, PA, TN, TX &amp; VA and Internationally to AUSTRALIA, FRANCE, INDIA, ITALY, and GERMANY. | Terminated | June 09, 2015 | 343 | 190 | Computer
Z-0827-2014 | 67264 | No | LZG | Infusion pump. | General Hospital | General Hospital | 510(k) | Product trade name t:slim" Insulin Delivery System |  |  | January 10, 2014 | 2014 | Tandem Diabetes Care Inc | Tandem Diabetes CareÂż, Inc. announced that it is initiating a voluntary recall of specific lots of insulin cartridges that are used with the t:slimÂż Insulin Pump. The affected cartridges may be at risk for leaking. The cause of the recall was identified during Tandems internal product testing, and has not been associated with any complaints or adverse events reported by customers. However, a c | Equipment maintenance | Tandem Diabetes Care sent an Urgent Medical Device Voluntary Recall letter dated January 10, 2014, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed check the lot numbers of their cartridge supply against the llist of affected lot numbers. Do Not use any cartridges from the lots affected by this recall. If customers don't have useable cartridges, discontinue using the pump and revert to their backup plan until new supplies arrive. Customers were instructed to call Tandem Technical Support at 1-877-801-6901 to receive replacement cartridges at no charge. Customers were asked to return the enclosed return response form via email, fax, or mail. UPDATED: Tandem Diabetes Care expanded the recall om January 20, 2014. Tandem Diabetes CareÂż, Inc. (NASDAQ: TNDM), or Tandem, today announced that it is expanding a voluntary recall of specific lots of insulin cartridges that are used with the t:slimÂż Insulin Pump.For questions regarding this recall calll 858-366-6900, ext 6963. |  | 12 | 12, 807 boxes (10packs) | USA (nationwide) | Terminated | September 23, 2014 | 256 | 12 | Computer
Z-2267-2014 | 68719 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | IntelePACS software versions 3-8-1 to 4-9-1, |  |  | July 03, 2014 | 2014 | Intelerad Medical Systems | In specific situation a partially truncated file may be written to the IntelePACS server. This partially truncated file still can be viewed by InteleViewer and InteleViewer will report to the user that the file in question is complete. This situation only occurs if a file is partially truncated. | Software design | Intelerad sent an Urgent Recall for Product Correction letter to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers were informed that a new software release will be installed on the system. Afterwards, an audit will be run on the the entire PACS to identify all image data that may have been truncated. Customers were requested to complete and return the attached Client Response Form. |  | 230 | 230 | Worldwide Distribution - USA Nationwide and the countries of Scotland, Ireland, United Kingdom, Canada, Australia, and New Zealand. | Terminated | January 13, 2016 | 559 | 230 | Computer
Z-0965-2014 | 66947 | No | LCP | Glycosylated hemoglobin assay. | Hematology | Hematology | 510(k) | D-10 Dual Program, which contains Update Kit D-10 A1c program diskette. |  |  | November 22, 2013 | 2013 | Bio-Rad Laboratories, Inc. | D-10 A1c Dual Program Floppy was programmed incorrectly and cartridge injections counts are not being updated correctly when switching between the D-10 Dual Short Program and Extended Program on the D-10 Hemoglobin Testing System. | Software design | BIO-RAD sent an Urgent Medical Device Correction letter dated November 22, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to continue to use the affected update kit disk(s) included with the reorder pack(s), but manually track the number of injections on the cartridge, reviewing QC results closely throughout the life of the reorder pack or destroy the affected disk(s) included with these reorder pack(s) and replace with the new Update Kit disk(s) The new Update Kit disks will elminate the problem and may be used with all reorder packs. The replacement disks do not contain reagent or calibrator information; after installing teh new Update Kit disk, they ust manually update the component information using instructions provided with the disk.For questions regarding this recall call 510-741-6265. |  | 2906 | 2906 kits (318 Kits in US; 2588 kits internationally) | Worldwide Distribution - USA (nationwide) and Internationally to Latin America, Asia Pacific, Canada, Australia, New Zealand and India. | Terminated | March 20, 2014 | 118 | 2906 | Computer
Z-0971-2014 | 67437 | No | LIT | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | Boston Scientific, Coyote" MONORAIL" PTA Balloon Dilatation Catheter, UPN # H74939185201510, Catalogue # 39185-20151. |  |  | February 04, 2014 | 2014 | Boston Scientific Corporation | Boston Scientific is initiating this Recall Removal of one lot of the Coyote" OVER-THE-WIRE PTA Balloon Dilatation Catheter and one lot of the Coyote" MONORAIL" PTA Balloon Dilatation Catheter. Boston Scientific has become aware that the balloons of the devices from these lots may not deflate as expected during a procedure. To date, Boston Scientific has received three complaints for this issue. T | Device Design | Consignees were sent a Boston Scientific " Urgent Medical Device Recall Removal" letter on 2/4/2014. The letter was addressed to Materials Manager and Field Action Contact. The letter described the product involved in the recall and the problem. Advised consignees to segregate the affected lots and return to Boston Scientific in accordance to the enclosed recall instructions. The instructions also requested consignees to complete and return the Account Reply Verification Tracking Form (RVTF). |  | 35 | 35 for both products | US: AL, CA, FL, IL, IN, MI, MO, NY, OH, PA, TX.OUS: ITALY, GERMANY, FRANCE, GREAT BRITAIN, NETHERLANDS, WHITE RUSSIA, INDIA, JAPAN. | Terminated | August 04, 2014 | 181 | 35 | Computer
Z-0979-2014 | 67305 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | GEMINI TF Big Bore PET/CT, Computer Tomography System, Philips Healthcare System, Cleveland, OH |  |  | December 05, 2013 | 2013 | Philips Medical Systems (Cleveland) Inc | The TF Big Bore patient table experienced mechanical binding during horizontal table motion resulting in an automatic Emergency Stop (E-stop). An E-Stop will interrupt a current scan, resulting in an incomplete study. | Software design | ON 12/5/2013 the firm sent Medical Device Correction Letters to their customers. |  | 47 | 47 Units | Product was shipped to the following states: AR, CA, CO, FL, IL, LA, MA, MD, MI, MN, MO, NC, NE, NJ, NY, OH, PA, TX, VA, VT &amp; WI.Product was also shipped to the following countries: Belgium, Canada, China, Denmark, France, Germany, Italy, Japan, Netherlands, Saudi Arabia &amp; Venezuela. | Terminated | August 25, 2015 | 628 | 47 | Computer
Z-2252-2014 | 68979 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Philips IntelliSpace PACS 4.4 image management systems performs digital image processing, measurement and communication and storage. |  |  | July 24, 2014 | 2014 | Philips Healthcare Informatics, Inc. | A problem was detected in the Philips IntelliSpace PAC 4.4 software where if patient information is changed in the DICOM header and the study is exported before the image is updated in the PACS, the unchanged data may export causing the potential for diagnosis on incomplete data. | Software design | Urgent Medical Device Correction was sent to all affected sites on July 24, 2014 by certified mail. |  | 109 | 109 units | US and Australia, Belgium, Canada, Egypt, France, Germany, Netherlands, Oman, Saudi Arabia, Spain, Switzerland, United Arab Emirates, United Kingdom. | Terminated | February 11, 2016 | 567 | 109 | Computer
Z-0976-2014 | 67374 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Horizon Medical Imaging, Picture Archive Communication System (PACS). |  |  | December 23, 2013 | 2013 | Mckesson Information Solutions LLC | Images may be archived to an incorrect or invalid archive location resulting in incorrect images or no images being displayed for a study following retrieval from archive. | Software design | McKesson sent a field notice to customers with the potentially affected software versions on 12/23/2013. |  | 618 | 618 units | U.S. Nationwide (Puerto Rico), Australia, Canada, France, GU, IE, and The United Kingdom. | Terminated | April 25, 2017 | 1219 | 618 | Computer
Z-2239-2014 | 68948 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | TomoTherapy Treatment System, Model: Hi-Art, Catalog/Part Number: H-0000-0003, software versions 1.2.0, 1.2.1 and 1.2.2 (Hi-ArtÂż 4.2.0, 4.2.1 and 4.2.2). |  |  | July 17, 2014 | 2014 | TomoTherapy Incorporated | Accuray is voluntarily recalling TomoTherapy H Series software versions 1.2.0, 1.2.1 and 1.2.2 (Hi-ArtÂż 4.2.0, 4.2.1 and 4.2.2). Accuray has identified potential safety issues (anomalies) with these software versions. | Software design | Consignees were sent an Accuray "Urgent Device Correction" letter dated July 16, 2014. The letter was addressed to TomoTherapy System Medical Physicist. The letter described the Purpose of this letter, Description of the Potential Problems, Safety Instructions and software upgrade information. For questions or concerns they can contact Accuray Customer Support at 1-877-668-8667 (USA) or 1-408-716-4700(non-USA) or customersupport@accuray.com |  | 322 | 322 | Worldwide Distribution-USA (nationwide) including DC and Puerto Rico and the states of US:; AL, AK, AZ, CA, CO, CT, FL, GA, HI, IL IN, IA, KS,KY, LA, MD, MA, MI, MO, MT, NE, NV, NJ, NM, NY, NC,ND, OH,OK, OR, PA, SC, SD, TX, UT, VA, WV, and WI, and the countries of AUSTRALIA, BELGIUM CANDA, CHINA, EL SALVADOR, ENGLAND, FRANCE, GERMANY, HONG KONG, INDIA, ITALY, JAPAN, KOREA, MALAYSIA, MEXICO, NETHERLANDS, PHILIPPINE, POLAND, RUSSIA, SAUDI ARABIA, SINGAPORE, SPAIN, SWEDEN, TAIWAN, THAILAND, TURKEY, and UNITED ARAB EMIRATES. | Terminated | September 08, 2015 | 418 | 322 | Computer
Z-2240-2014 | 68960 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | TomoTherapy Treatment System, Model: Hi-Art, Catalog/Part Number: H-0000-0003, software versions 2.0.0 and 2.0.1 (Hi-ArtÂż 5.0.0 and 5.0.1). |  |  | July 16, 2014 | 2014 | TomoTherapy Incorporated | Accuray is voluntarily recalling TomoTherapy H Series software versions 2.0.0 and 2.0.1 (Hi-ArtÂż 5.0.0 and 5.0.1). Accuray has identified potential safety issues (anomalies) with these software versions. | Software design | Accuray sent an "Urgent Device Correction" letter dated July 16, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customers. These anomalies will be corrected with a new software version.You will be contacted by Accuray Customer Support to schedule a software upgrade to your system.If you have any questions or concerns regarding this issue, please do not hesitate to contact Accuray Customer Support at +1.877.668.8667 (USA) or +1.408.716.4700 (non-USA) or customersupport@accuray.com. |  | 168 | 168 (75 US; 93 OUS) | Worldwide Distribution - Nationwide including the states of AZ, AR, CA, FL, GA, IL, IN, KS, KY, MD, MN, NM, NY,NC, OK, OR, SD, TN, TX, UT, WA, WY, WI and WV., and the countries of AUSTRALIA, BELGIUM, CANADA, CHINA, ENGLAND, FRANCE, GERMANY, HONG KONG, ITALY, JAPAN, KOREA, MEXICO, NETHERLANDS, PHILIPPINES, RUSSIA, SPAIN, TAIWAN, TURKEY, MARTINIQUE, ROMANIA, SRI LANKA and SWITZERLAND. | Terminated | July 21, 2015 | 370 | 168 | Computer
Z-2237-2014 | 68953 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | IntelliSpace Portal (ISP) DX/HX/EX, IntelliSpace Portal (ISP) IX, and IntelliSpace Portal (ISP) LX SPECT |  |  | July 27, 2014 | 2014 | Philips Medical Systems (Cleveland) Inc | The following software issues have been identified in the affected products.Problem 1: When reopening a bookmark generated from processing a MUGA (Multi-Gated Acquisition) scan within the NM Cardiac Application, the ejection fraction (EF) displayed should be the same as the ejection fraction (EF) originally displayed when the bookmark was first created. In some instances, the ejection fraction ( | Software design | Philips Healthcare issued an Urgent Medical Device Correction letter dated July 28, 2014, via certified mail to all affected consignees. The letter identified the product, the problem, and the action to be taken by the consignee. Philips Key Market representative would distribute the FSN internationally. Additionally, a Field Service Engineer would contact each site to schedule a time to install the software update.Customers with questions were instructed to contact their local Philips representataive or local Philips Healthcare office. For North America and Canada customers were instructed to contact the Customer Care Solutions Center at 1-800-722-9377, option 1.For questions regarding this recall call 978-659-7703. |  | 1099 | 1099 Units | Worldwide Distribution - USA including AK, AL, AR, AZ, CA, CO, FL, GA, IL, IN, KS, LA, MA, MD, MI, MN, MO, MS, NC, ND, NJ, NM, NV, NY, OH, OK, OR, PA, SC, TN, TX, VT, WA, WI, WV, WY,Washington D.C. and Internationally to the following countries: Algeria, Argentina, Australia, Austria, Bangladesh, Belgium, Brazil, Bulgaria, Canada, Cayman Islands, Chile, China, Colombia, Congo, Costa Rica, Czech Republic, Denmark, Ecuador, Egypt, Ethiopia, Finland, France, Gabon, Germany, Greece, Guatemala, Hong Kong, Hungary, India, Indonesia, Iraq, Ireland, Israel, Italy, Japan, Jordan, Kazakhstan, Kenya, Latvia, Lebanon, Libyan Arab Jamahiriya, Lithuania, Malaysia, Mauritius, Mexico, Mongolia, Morocco, Myanmar, Nepal, Netherlands, New Zealand, Nicaragua, Norway, Oman, Pakistan, Palestine, Panama, Peru, Philippines, Poland, Portugal, Puerto Rico, Qatar, Russia, Saudi Arabia, Singapore, Slovakia, South Africa, South Korea, Spain, Sweden, Switzerland, Taiwan, Thailand, Tunisia, Turkey, Ukraine, United Arab Emirates, United Kingdom, Viet Nam, and Yemen. | Terminated | February 02, 2016 | 555 | 1099 | Computer
Z-0941-2014 | 66467 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | enGen Track System with enGen Select v5.0 with autoverification enabled |  |  | October 08, 2013 | 2013 | Ortho-Clinical Diagnostics | Ortho Clinical Diagnostics issued an Urgent Product Correction Notification involving the enGen Select V5.0 &amp; enGen Custom v3.2.2 due to an anomaly that affects the autoverification feature and custom rules that depend upon data elements used by the referenced enGen Systems. | Software design | Ortho Clinical Diagnostics (OCD), issued an Urgent Product Correction Notification letter dated October 7, 2013 and on December 17, 2013 to their customers via Federal Express or via US Postal Service. The letter identified the affected product, problem and actions. The notices advise the customers to ensure that all required data elements are entered and in the correct format if using the autoverification or custom rules for the referenced enGen Laboratory Automation Systems. OCD trained service representatives will contact all affected customers. Customers with questions may call the Customer Technical Service at 1-800-421-3311. |  | 8 | 8 units | Worldwide Distribution - USA Nationwide in the states and countries of: AZ, FL, WV, TX, PA, DE, MI, OH, TN, IL, NC, Chile, India, Italy, Portugal, Spain, Sweden, UK | Terminated | April 11, 2017 | 1281 | 28 | Computer
Z-2222-2014 | 68656 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Double Offset Rasp Handle, Right Hand |  |  | July 23, 2014 | 2014 | Zimmer, Inc. | The firm received complaints for failure of the handle to remain assembled to the rasp during use due to various failure modes, including fracture of the leaf spring. | Device Design | On July 23, 2014, Zimmer issued recall notification to their consignees. Distributors were notified via electronic mail and courier. Recall notification included a description of the issue, affected product, possible risks to patient, and instructions for responding to the recall. Questions or concerns regarding the recall are directed to Zimmer customer call center @ 1-800-348-2759 M-F, 8am-8pm ET. |  | 412 | 412 units | Worldwide Distribution-USA (nationwide) including the states of AK, AL, AZ, CA, CO, FL, GA, IN, KS, KY, MI, MN, NC, NJ, NY, OH, OK, OR, PA, SD, TN, TX, UT, and VA, and the countries of Australia, France, Germany, Italy, Switzerland, and Canada. | Terminated | September 15, 2015 | 419 | 412 | Computer
Z-0981-2014 | 67343 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | GE Healthcare, PET Discovery 610, Discovery 710, Discovery 600 and Discovery 690. |  |  | November 20, 2013 | 2013 | GE Healthcare, LLC | GE Healthcare has recently become aware of a potential safety issue due to scatter overcorrection associated with the List Mode Replay feature of your Discovery 600, 610, 710 and Discovery 690. | Software Manufacturing/Software Deployment | GE Healthcare sent an Urgent Medical Device Correction letter dated November 20, 2013, and December 4, 2013, to all affected consignees. The letters described the Safety Issue, Safety Instructions, Affected Products Details, Product Correction and Contact Information ( In the US Call 1-800-437-1171. For other countries contact your local GE Healthcare Service Representative). Customers were given the following instructions:To avoid potential issues, replay the list mode data prior to another acquisition.In the event another acquisition is made before replaying the list data, or following reboot, do the following steps:Âż Perform a PET only acquisition using the same tracer that is intended for the List Mode Replay.Âż Then, proceed to perform the List Mode Replay and subsequent reconstructions.Âż Make sure to perform a quality check on any images created after any replay of data.AffectedProduct DetailsPET Discovery 610 and PET Discovery 710, with software versions pet_coreload.44 and pet_mfk.44. The software version is visible on the homepage of the service desktop.To determine the software version currently on your system, do the following:1. On the right monitor upper left corner select the Service Icon.2. Verify that the Service window is under Home.3. Locate the SW Build Version listed in the center of the page.For questions regarding this recall call 800-437-1171. |  | 376 | 376 (101 US, 275 OUS) | Worldwide Distribution - USA including AL, AZ, AR, CA, CO, FL, GA, IL, IN, IA, KY, LA, MN, MO, NE, NJ, NM, NY, NC, OH, OK, PA, SD, TN, TX, UT, VA, WA, WV, WI and Internationally to:AUSTRALIA, AUSTRIA BELGIUM BOSNIA AND HERZEGOVINA, BRAZIL, CANADA, CHILE, COLOMBIA, CHINA, CROATIA, CZECH REPUBLIC, DENMARK, EGYPT, FRANCE, FINLAND, GEORGIA, GERMANY, HONG KONG, IRAN, ISRAEL INDIA,M ITALY, JAPAN, JORDAN, KUWAIT, LEBANON, MALAYSIA, MEXICO, MOROCCO, NEW ZEALAND, NORWAY, RUSSIA, REPUBLIC OF KOREA, POLAND, ROMANIA, SAUDI ARABIA, SPAIN, SWEDEN SWITZERLAND, TAIWAN, TUNISIA, THAILAND, TURKEY, UNITED KINGDOM. | Terminated | July 29, 2014 | 251 | 376 | Computer
Z-2213-2014 | 68845 | No | HAW | Stereotaxic instrument. | Neurology | Neurology | 510(k) | Navigation Spine &amp; Trauma 3D Version 2.0 and 2.1 |  |  | May 21, 2014 | 2014 | Brainlab AG | Brainlab Navigation Software Spine &amp; Trauma 3D 2.0/2.1 offers automatic registration of intraoperatively acquired CT image data sets. To enable automatic registration, the software requires the gantry position of the scanner. The gantry position can either be entered manually or submitted automatically from the CT scanner. If automatic gantry communication is available, the so-called &quot; | Software in the Use Environment | The firm, Brainlab AG, sent "FIELD SAFETY NOTICE / PRODUCT NOTIFICATION" letters dated May 21, 2014 to all consignees. The letters included instructions for consignees to: 1) not use Automatic Registration of iCT Data in combination with Siemens SOMATOM Definition AS CT scan (with Software Somaris SOM 7 or higher and Automatic Gantry Communication (available for Spine &amp; Trauma 3D 2.0 and 2.1 only); 2) always make sure to perform careful accuracy verification according to the instructions in the Software User Guide; 3) Brainlab will perform corresponding service visits to change the gantry communication from automatic to manual to enable a safe and effective usage of this device combination, and 4) advise the appropriate personnel working in your department of the content of the letter.Customers with any questions can contact Brainlab via telephone at 800-597-5911 or e-mail at us.support@brainlab.com. |  | 17 | 17 systems | Worldwide distribution: US (nationwide) including states of: NC, NY, OH, WI;and countries of: Austria, China, Germany, Italy, Japan, Malaysia, and Saudi Arabia. | Open, Classified |  |  | 17 | Computer
Z-0982-2014 | 67370 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | GE Healthcare, Dash 3000, Dash 4000 and Dash 5000 Patient Monitors, Dash Port2 Docking Station. |  |  | November 27, 2013 | 2013 | GE Healthcare, LLC | GE has become aware of multiple issues affecting the Dash 3000/4000/5000 Patient Monitor and Dash Port 2 Docking Station:DASH: 1. Dash stops communicating on the Unity Network resulting in Loss of Monitoring at the CARESCAPE" CIC Pro Central Station*.2. Dash displays distorted ECG waveforms in combination monitoring mode. 3. Dash incorrectly becomes Time Master on the Unity NetworkÂż.4. | Software design | Consignees were sent a GE Healthcare "Urgent Medical Device Correction" letter dated November 27,2013. The letter was addressed to Healthcare Administrator / Risk Manager, Chief of Nursing and Director of Biomedical Engineering. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction and Contact Information (Contact Technical Support at 1-800-558-7044 or your local Service Representative). |  | 178711 | 178,711 (92,123 US, 86,588 OUS) | Worldwide Distribution: USA (nationwide) including DC and PR; and Internationally to: ALGERIA, ANGOLA, ARGENTINA, AUSTRALIA, AUSTRIA, BAHAMAS, BAHRAIN, BANGLADESH, BELARUS, BELGIUM, BOLIVIA, BOSNIA-HERZEGOVINA, BRAZIL, BRUNEI, BULGARIA, CAMBODIA, CAMEROON, CANDA, CARIBBEAN, CHILE, CHINA, COLOMBIA, COSTA RICA, CROATIA, CYPRUS, CZECH REPUBLIC, DENMARK, DJIBOUTI, DOMINICAN REPUBLIC, ECUADOR, EGYPT, EL SALVADOR, ESTONIA, ETHIOPIA, FINLAND, FRANCE, GABON, GEORGIA, GERMANY, GHANA, GREECE, GUATEMALA, HONDURAS, HONG KONG, HUNGARY, ICELAND, INDIA, INDONESIA, IRAN, IRAQ, IRELAND, ISRAEL, ITALY, IVORY COAST, JAMAICA, JAPAN, JORDAN, KAZAKHSTAN, KENYA, KOREA, KUWAIT, LATVIA, LEBANON, LIBYA, LITHUANIA, MACEDONIA, MALAYSIA, MALTA, MAURITIUS, MEXICO, MOLDOVIA, MOROCCO, NEPAL, NETHERLANDS, NEW ZEALAND, NICARAGUA, NIGERIA, NORWAY, OMAN, PAKISTAN, PALESTINE, PANAMA, PARAGUAY, PERU, PHILIPPINES, POLAND, PORTUGAL, QATAR, ROMANIA, RUSSIA, SAUDI ARABIA, SENEGAL, SERBIA MONTENEGRO, SINGAPORE, SLOVAKIA, SLOVENIA, SOUTH AFRICA, SPAIN, SRI LANKA, SUDAN, SWEDEN, SWITZERLAND, SYRIA, TAIWAN, TAJIKISTAN, TANZANIA, THAILAND, TRINIDAD TOBAGO, TUNISIA, TURKEY, TURKMENISTAN, UGANDA, UKRAINE, UNITED ARAB EMIRATES, UNITED KINGDOM, URUGUAY, VENEZUELA, VIETNAM, YEMEN, and ZIMBABWE. | Open, Classified |  |  | 178711 | Computer
Z-2199-2014 | 68788 | Yes | KDI | High permeability hemodialysis system. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Prisma M100 PRE pump infusion set. Product number 107791. Each set is packaged in a plastic pouch, 4 pouches per carton box |  |  | July 18, 2014 | 2014 | Gambro Renal Products, Incorporated | Gambro initiated a field action on the Prisma and Prismaflex sets due to breakages of connectors resulting in leakage of liquid or blood. | Device Design | Gambro sent an Urgent Field Safety Notice dated July 18, 2014 to all end-users. A Customer Reply Form was attached to the notice for customers to complete and return. |  | 401451 | 401,451 | Worldwide Distribution. | Terminated | February 05, 2015 | 202 | 401451 | Computer
Z-2184-2014 | 68837 | No | JXG | Central nervous system fluid shunt and components. | Neurology | Neurology | 510(k) | Codman Certas - In Line Valve Only; Product Code: 82-8800 |  |  | July 03, 2014 | 2014 | Codman &amp; Shurtleff, Inc. | To clarify the CODMAN CERTAS Valve virtual off (setting 8) pressure specification outlined in the product's Instructions for Use (IFU). | Device Design | Codman Neuro sent an Urgent Medical Device Safety Notification dated July 17, 2014, to clinicians who use the device. The notification clarifies the CODMAN CERTAS Valve virtual off (setting 8) pressure specification outlined in the product's Instructions for Use (IFU). Consignees are also asked to complete the enclosed Acknowledgement Form and fax the completed form to 1-888-239-1305. For questions or concerns regarding this notification, consignees are instructed to contact their local Codman Neuro Representative or Scientific and Medical Affairs at SciMedAffairs@its.jnj.com or (866) 685-7325.For questions regarding this recall call 508-828-2726. |  | 9498 | 9,498 devices total | Worldwide Distribution - USA including AL, AZ, CA, CO, CT, DC, DE, FL, GA, IA, ID, IL, IN, KY, LA, MA, MD, ME, MI, MN, MO, MT, NC, ND, NE, NJ, NV, NY, OH, OK, OR, PA, RI, SC, TN, TX, UT, VA, VT, WA, WI, WV, WY; and Internationally to: Canada, Australia, Austria, Belgium, Brazil, Czech Republic, Denmark, Egypt, Finland, France, Germany, Great Britain, Greece, Hungary, Ireland, Israel, Italy, Kuwait, Lebanon, Luxembourg, Netherlands, Norway, Portugal, Russia, Saudi Arabia, Slovak Republic, Slovenia, South Africa, Spain, Sweden, Switzerland, Turkey. | Open, Classified |  |  | 9498 | Computer
Z-0988-2014 | 67241 | No | LEZ |  |  | Ear Nose &amp; Throat | Enforcement Discretion | Siemens Prisma 2K Hearing Aid |  |  | July 01, 2005 | 2005 | Siemens Hearing Instruments, Inc | Siemens became aware of an issue where an infant or child could open the tamper resistant battery compartment of the Siemens Prisma 2K hearing aid under certain unintended use conditions. | Device Design | Siemens sent a notice letter to affected customers in 7/2005 and 12/2005. The letter identified the affected product, problem and actions to be taken. Siemens will install an additional security measure to affected product at no charge. For questions contact your Sales Representative at 800-766-4500. |  | 738 | 738 units | USA Nationwide Distribution | Terminated | July 06, 2017 | 4388 | 738 | Computer
Z-2180-2014 | 68634 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Siemens Syngo RT Therapist and Syngo RT Oncologist intended use of the SIEMENS branded Syngo RT Therapist and Syngo RT Oncologist software applications used with SIEMENS family of linear accelerator systems is to deliver X-Ray photon and electron radiation for the therapeutic treatment of cancer. |  |  | June 17, 2014 | 2014 | Siemens Medical Solutions USA, Inc | Siemens became aware of a potential issue with the syngo RT Therapist 4.3.SP1 and syngo RT Oncologist 4.3SP1 software applications running on Siemens brand LINAC systems that may experience sporadic system crashes that will cause the inability to unload patient data. | Software design | In June, 2014, an urgent medical device correction field safety notice of improvement was sent to end users informing them of the product, problem, and action to be taken to avoid the potential safety issue. |  | 56 | 56 | UT, WI, NY | Terminated | December 15, 2014 | 181 | 56 | Computer
Z-2179-2014 | 68888 | No | MDS |  |  | Clinical Chemistry | PMA | Medtronic MiniMed Guardian Monitor, Model No. CSS7100, CSS7100K |  |  | July 21, 2014 | 2014 | Medtronic MiniMed Inc. | Medtronic MiniMed is recalling the Guardian REAL-Time monitors because of a manufacturing issue that has the potential to compromise the water resistance of the affected Guardian monitors. | Device Design | Medical Device Recall letters dated July 21, 2014 andJuly 22, 2014 were sent to distributors, patients, and healthcare providers who purchased the Guardian REAL time monitors. The letters inform the customers of the problems identified and the actions to be taken. Distributors with questions are instructed to email rs.nrdlrinbox@medtronic.com. If healthcare professionals want to speak to Medtronic's medical officers then they are instructed to call (818) 576-4211, ext. 64511. Customers and healthcare professionals' patients are instructed to call Guardian Recall Hotline at (800) 646-4633, ext. 21090, 8am-6pm, CST. |  | 86 | 86 units | Worldwide:Distribution - US Nationwide and the countries of Croatia, Denmark, Finland, Germany, Netherlands, Norway, Sweden, Switzerland. | Terminated | October 08, 2014 | 79 | 86 | Computer
Z-2177-2014 | 68727 | No | KPR | Stationary x-ray system. | Radiology | Radiology | 510(k) | Kodak DirectView DR 7500 Diagnostic X-Ray System |  |  | July 09, 2014 | 2014 | Carestream Health, Inc. | An on-site evaluation of the equipment parts that attach the Beta Assembly (x-ray tube and collimator) to the Over Head Tube Crane (OTC) Telescope was completed by Carestream Health (CSH) Service Engineering. Inspection of these parts revealed that the tab on the lock washer was not bent into the lock nut recess and because the Beta Assembly detached from the OTC telescope, the lock nut was not sufficiently tightened. | Device Design | Carestream will notify all DR 7500 customers of the issue and schedule a Field Engineer to inspect and repair the system if required. Specific instructions (Service Modification) for the inspection and repair will followed by the Carestream Field Engineer in order to complete the correction. Documented evidence of the Field Engineer visit will be noted in the Carestream service engineering record for each affected system. Carestream estimates that the plan of correction will be completed within 2-4 weeks |  | 589 | 589 systems | Worldwide Distribution - US Nationwide and the countries of (United Kingdom, United Arab Emirates, Turkey, Taiwan, Switzerland, Sweden, Spain, South Africa, Portugal, Poland, Oman, New Zealand, Netherlands, Kuwait, Italy, Israel, Hong Kong, Germany, France, Finland, Denmark, Dubai, Czech Republic, China, Korea, Belgium, Australia, and Amsterdam. | Terminated | May 12, 2017 | 1038 | 589 | Computer
Z-2175-2014 | 68752 | No | KPR | Stationary x-ray system. | Radiology | Radiology | 510(k) | DX-D 600 - DXD Imaging Package |  |  | July 02, 2014 | 2014 | AGFA Healthcare Corp. | Some customers reported that while using the DX-D600 in room or exam preparation, the Overhead Tube Crane (OTC) showed increased velocity and uncontrolled longitudinal movement. | Device Design | AGFA HealthCare sent an URGENT FIELD SAFETY NOTICE letter dated July 2, 2014 to affected customers. The letter described the affected product, problem and actions to be taken. Customers were instructed to complete and return the Acknowledgment form, via FAX-Back or email. A Mandatory Service Bulletin DIS012.14E/No. 32, was released June 30, 2014, to implement the mandatory upgrade of full automatic DX-D600 systems to version 3.6 to prevent unexpected system movements. This Mandatory Service Bulletin was issued to Agfa DR Specialists to perform the service. A 2nd release will be implemented to upgrade semi-automatic systems. For questions contact your local AGFA HealthCare organization 1-877-777-2432. |  | 17 | 17 units | Worldwide Distribution - US Natinwide in the state of GA, IN, KS, KY, MA, NJ, SC, VA, DC and the country of Canada | Terminated | June 04, 2015 | 337 | 17 | Computer
Z-2167-2014 | 68648 | No | JOP | Ultrasonic transducer. | Cardiovascular | Cardiovascular | 510(k) | Esaote PA230E Transducer |  |  | June 26, 2014 | 2014 | Biosound Esaote, Inc. | Using the PA230E transducer with a user-defined setting with the output power to 100% (Max) in 2D-TEI mode focusing at maximum depth, could cause the transducer surface temperature to exceed the maximum allowable limit. | Software design | On 06/26/2014, Esaote North America sent an advisory letter via FEDEX (with signature required) to each customer or consignee affected by this voluntary recall. The packet contained the advisory letter, including actions to be taken by the user, information about a pending software upgrade, and a business reply card to confirm they received the letter and have the equipment noted on the advisory letter. Customers with questions were instructed to call 800-428-7378.For questions regarding this recall call 1-317-813-6005 |  | 1301 | 1,301 units | Nationwide-Product was shipped to the following states: AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, MT, NC, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, PR, SC, SD, TN, TX, UT, VA, WA, WI &amp; WV. International**Product was also shipped to consignees in Canada | Terminated | July 17, 2015 | 386 | 1301 | Computer
Z-2135-2014 | 68596 | No | LMH |  |  | General &amp; Plastic Surgery | PMA | Enhancement Medical expression. model # EXP-150. Injectable Gel 1.5cc Syringe with hyaluronic acid concentration of 20 mg/ml. |  |  | August 07, 2013 | 2013 | Enhancement Medical, LLC | Manufacturer made a change in the production process that resulted in a change in final gel weight. RECALL EXPANDED 7/8/2014 Firm expanded their recall to include all lots of product. | Device Design | Enhancement Medical, LLC called all Independent Sales Representatives and then provided each representative a list of consignees with impacted product and instructed the Representatives to disclose low gel weight issue and potential for common treatment responses. Representatives were instructed to exchange the affected product. EXPANDED Enhancement Medical sent an Urgent Medical Device Recall letter beginning July 8, 2014 to all consignees whom received product. The letter identified all lots as being recalled, the reason for the recall, recommended that Expression not be injected subcutaneously, and asked for product to no longer be used and be returned to the recalling firm. An acknowledgement Form was asked to be mailed, e-mailed or faxed back to the firm.Customers with questions were instructed to call 414-918-4280. |  | 17875 | 17,875 | Worldwide Distribution - USA (nationwide including Puerto Rico) and Internationally to Canada. | Open, Classified |  |  | 17875 | Computer
Z-2535-2014 | 68983 | No | DQK | Programmable diagnostic computer. | Cardiovascular | Cardiovascular | 510(k) | Philips IntelliSpace ECG Management System, with software option C61 needed to import ECGs from the Quinton Stress ECG System. |  |  | August 07, 2014 | 2014 | Philips Medical Systems, Inc. | Philips IntelliSpace ECG (IECG) Management Systems that are importing Stress ECG records from a Quinton Stress ECG system have the potential for the patient record to contain multiple and therefore inaccurate patient identifiers such as: patient name, patient ID (PID), Medical record number (MRN). This creates the potential for misdiagnosis and incorrect therapy. | Software design | Philips Healthcare sent an Urgent Medical Device Correction Notification / Field Safety notice dated August 7, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. This Field Safety Notice is intended to inform you about:" what the issue is and under what conditions it can occur" the actions that should be taken by the customer/user in order to prevent risks for patients" the corrective action planned by Philips to address the following issue1)Previously Imported Stress Records:Institutions that have imported Quinton Stress ECG records into IECG should review each previously imported record to make sure that multiple patient ID data is not present. Institutions that imported stress records at the end of each stress exam or did not batch process the records should not find any records with multiple Patient IDs in the record. However, we do recommend that you perform the review to confirm that multiple Patient IDs do not exist.2) Continued Use: Software is now available to address the issue described in this letter. If you have not already been contacted by Philips please call to arrange for immediate installation of the software upgrade.The correction will consist of free of charge software upgrade to A.01.06.. A Philips Healthcare representative will contact customers with affected devices to arrange for service. Further information or support concerning this issue, please contact your local Philips representative or call us at 1-800-722-9377 |  | 11 | 11 units | Worldwide Distribution - US including the states of CA, GA, NC, NH, NY and OR., and the countries of Canada and Spain. | Terminated | October 21, 2015 | 440 | 11 | Computer
Z-2543-2014 | 68801 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Zimmer PERSONA Distal Valgus Alignment Guide, Nonsterile. orthopedic surgical instrument. |  |  | July 30, 2014 | 2014 | Zimmer, Inc. | Persona Distal Valgus Alignment guide collet locks are seizing in the 'unlocked' position. | Device Design | An Urgent Medical Device Correction letter dated, July 31, 2014, was issued to consignees. The letter identified the affected product, the problem, the risks, and responsibilities. Risk managers and surgeons are to review the notification and the updated surgical technique provided in Attachment 1; as well as, ensure that all affected personnel are aware of the contents. Questions or concerns should be directed to Zimmer customer call center at 1-877-946-2761. |  | 2248 | 2,248 units | Worldwide Distribution -- United States, including Puerto Rico; and, to the countries of Australia, Austria, Japan, Belgium, Italy, Switzerland, India, Germany, Spain, France, Israel, Netherlands, Saudi Arabia, South Africa &amp; Canada. | Terminated | March 06, 2015 | 219 | 2248 | Computer
Z-1084-2014 | 67184 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Agility. |  |  | November 25, 2013 | 2013 | Elekta, Inc. | During the treatment table, gantry, and collimator calibration procedures, it is possible to type incorrect values or invert a value. | Software design | Corrective Action #1 - Release of Important Field Safety Notice 200 01 502 043 - "Recommended quality assurance after a calibration procedure" Released 25 November, 2013. Corrective Action #2 - Permanent Solution - All users of Integrity 3.1 software will be informed or a mandatory upgrade to Integrity 3.2 upon its release which will prevent calibration from being inverted. Target Release for Integrity 3.2 will be available at the end of February 2014. |  | 300 | 300 | Worldwide Distribution-USA (nationwide) including DC and Puerto Rico and the states of AL, AK, AZ, CA, CT, FL, GA, IL, KY, LA, MD, MA, MI, MO, NJ, NM, NY, NC, ND, OH, OK, OR, PA, RI, SD, TN, TX, VT, VA, WA, WV, WI, and WY, and the countries of Algeria, Argentina, Australia, Austria, Bahrain, Belarus, Belgium, Brazil, Bulgaria, Canada, China, Colombia, Czech Republic, Cuba, Denmark, Ecuador, Estonia, Finland, France, Germany, Greece, Honduras, Hong Kong, Hungary, India, Indonesia, Iraq, Ireland, Israel, Italy, Japan, Jordan, Malaysia, Mexico, Morocco, Netherlands, New Zealand, Nigeria, Norway, Pakistan , Panama, Peru, Poland, Portugal, Romania, Russia, Saudi Arabia, Singapore, Slovakia, Slovenia, South Africa, South Korea, Spain, Sudan, Sweden, Switzerland, Taiwan, Thailand, Turkey, United Kingdom, Ukraine, Venezuela, and Vietnam. | Open, Classified |  |  | 889 | Computer
Z-2586-2014 | 69125 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | ARIATELE TELEMETRY TRANSMITTER, Model 96281, with SpO2 Option C. Physiological patient monitor. Used in a hospital environment. |  |  | July 02, 2014 | 2014 | Spacelabs Healthcare Inc | The transmitter display and the Telemetry Central Station display may show a SpO2 value when the sensor is not connected to the patient or in certain sensor failed conditions.This is an expansion of recall Z-0871-2014 because additional serial numbers were identified as being part of this recall action as part of the firm's retrospective analysis of all recalls since 2011. | Software design | Spacelabs notified both consignees via email of the expanded recall July 2, 2014. A f/u letter was sent by overnight delivery to the single new consignee the same day. Spacelabs sent an announcement of the expanded recall to both consignees identified in the expanded recall August 25, 2014. A field service technician from Spacelabs will visit each consignee and upgrade the software as necessary.If you have any questions about this corrective action program, please contact SpaceLabs Healthcare at 1-800- 522-7025, and select 2 for Technical Support. |  | 8 | 8 units | US Distribution to TX. | Terminated | October 16, 2014 | 106 | 8 | Computer
Z-2637-2014 | 69086 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | ProVue Instrument Driver for JResultNet (ProVue.jrm) v. 1.5.3 and below |  |  | August 15, 2014 | 2014 | Data Innovations, LLC | Potential issues with the ProVue driver have been identified which may cause data to be misassociated with patient results under a given set of circumstances. | Software design | Data Innovations LLC sent an Important Medical Device Information letter on August 15, 2014, to consignees for the JResultNet using ProVue Instrument Driver JResultNet (provue.jrm) v. 1.5.3 and below. Potential issues with the ProVue driver have been identified which may cause data to be mis-associated with patient results under a given set of circumstances. The user on the ProVue transmits the same results twice back to back by clicking the Send to Host button twice. The ProVue sends two ASTM messages back to back before ending transmission (sending an EOT character) The solution was to disconnect JResultNet from the ProVue instrument and apply the patch to driver provue.jrm when it is available. If you have further questions regarding this notice, please feel free to contact us at northamerica-support@datainnovations.com or 239-931-6004 Toll-Free (from the US and Canada): 1-800-332-0499. |  | 84 | 84 | Nationwide Distribution including AL, AZ, CA, CT, DC, FL, GA, IN, LA, MI, MO, NC, NH, NJ, NY, OH, OK, OR, PA, TN, TX, UT, VA, and WA. | Terminated | March 04, 2015 | 201 | 84 | Computer
Z-0907-2014 | 67138 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Siemens SOMATOM Definition, Definition AS, Definition Flash, and Definition Edge systems. |  |  | November 15, 2013 | 2013 | Siemens Medical Solutions USA, Inc | During clinical operation, a malfunction may occur causing a system display freeze with SOMATOM Definition, Definition AS, and definition Edge while running software version syngo CT2012 B and syngo CT2013B. | Software design | A recall notification letter, dated November 15, 2013, was sent to Users. |  | 147 | 147 | Nationwide Distribution. | Terminated | March 10, 2015 | 480 | 147 | Computer
Z-2635-2014 | 69038 | No | LEJ | Sunlamp products and ultraviolet lamps intended for use in sunlamp products. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | SUN DASH RADIUS 252 PN 8000809705. |  |  | July 25, 2014 | 2014 | JK Products &amp; Services, Inc | The integrity of the wires inside main electrical chord become compromised from excessive flexing (during the door open/close process). | Device Design | The owners of the affected device will be notified via certified letter. After the recalling firm confirms ownership, a technician will be dispatched to install a door clamp and door stop, and inspect the condition of the main electrical chord. |  | 1948 | 1948 units | Nationwide Distribution. | Open, Classified |  |  | 1948 | Computer
Z-2626-2014 | 68908 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Stryker Aria Implant Inserter, REF 48758500, |  |  | July 07, 2014 | 2014 | Stryker Spine | Stryker has initiated the recall of its Aria Standard and Specialty Implant Inserters due to complaints of product fracturing. | Device Design | Stryker Spine sent an "Urgent Medical Device Removal Notification/Customer Response" form dated July 7, 2014 to all their affected customers via Fedex. Customers were instructed to examine their inventory and hospital locations to identify the product. Fax a copy of the Customer Response Form to 855-632-9049 or email it to Spine-RegulatoryAction@Stryker.com.For questions regarding this recall call 201-760-8298. |  | 64 | 64 units | Worldwide Distribution - USA including OR, AZ, TN, MA, MD, GA, FL, NY, IA, IN, TX, MI, PA, UT, CA, MN, MO, OH, KS, OK, MT, FL, AL, CO, TN, and WI and Internationally to Australia. | Open, Classified |  |  | 121 | Computer
Z-2292-2014 | 68785 | No | JAA | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Siemens Luminos dRF, Ysio or Uroskop Omnia systems. |  |  | July 07, 2014 | 2014 | Siemens Medical Solutions USA, Inc | It was discovered that during a RAD examination using Siemens Luminos dRF, Ysio or Uroskop Omnia systems with software version VB10C to VB10F and automatic exposure control, a highly unlikely malfunction can result in longer radiation exposure than required. This may also result in an overexposed image that is not of diagnostic quality. As a result the taken examination needs to be repeated. Thi | Device Design | Siemens had informed all affected customers about this issue via Customer Safety Advisory Notice XP023/13/S (dated June 3, 2013) reported under C&R; # 2240869-05/16/13-0013-C. A corresponding software update to fix the issue was originally planned to be rolled out by May 2014. The originally scheduled roll out date could not be met and to ensure that this new software version works accurately Siemens incorporated a comprehensive monitoring phase. Siemens has now scheduled the roll out for the software version VB10J to be completed by end of October 2014. A letter, dated July 7, 2014, with important safety information was sent to end users to notify them that the new software version VB10J will be introduced to solve the potential malfunction of the generator. |  | 192 | 192 | Nationwide Distribution. | Terminated | February 24, 2015 | 232 | 192 | Computer
Z-0908-2014 | 67242 | No | LWR |  |  | Cardiovascular | PMA | Product Name: Model 7305- Mosaic Aortic Obturator |  |  | January 09, 2014 | 2014 | Medtronic Cardiovascular Surgery-the Heart Valve Division | Medtronic, Inc. initiated a recall of Medtronic Mosaic Aortic Obturator Model 7305 used with the Mosaic Porcine Aortic Bioprosthesis, Model Number 305,because of higher-than-expected transvalvular gradients occurring post implant. | Device Design | Medtronic sent an Customer Notification letter dated January 9, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customers. Medtronic recommends surgeons continue to remain diligent in selecting the appropriate size of Mosaic Aortic Bioprosthesis. Adoption of the updated Mosaic Aortic Bioprosthesis sizing chart (iEOA chart) and Mosaic Obturators/Sizers will facilitate optimal aortic valve size selection.Your Medtronic Field Representative will be contacting you about receiving a new Mosaic Aortic Obturator/Sizer set and sizing chart. If you would prefer to order the new Mosaic Aortic Obturator/Sizer set directly through Medtronic Customer Service, contact them at 1-800-848-9300.Medtronic will notify all applicable regulatory agencies, as required, about this matter. Please share this notification with others in your organization as appropriate. We appreciate your review of this notification and apologize for the inconvenience that it may cause. If you have any questions, please contact your Medtronic Field Representative or Lifeline Technical Services at 1-877-526-7890. |  | 8276 | 8276 units | Worldwide Distribution - USA (nationwide) | Terminated | May 30, 2014 | 141 | 8276 | Computer
Z-2616-2014 | 69036 | No | NKB | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | Streamline MIS Spinal Fixation System Fixed Rod Holder |  |  | June 14, 2013 | 2013 | PIONEER SURGICAL TECHNOLOGY, INC. | The affected fixed rod holders did not allow the instrument to lock down on the spinal rod. | Device Design | Pioneer Surgical sent a Product Field Action letter on July 7, 2013, to all affected customers. The letter described the affected product and issue, and provided instructions for return of the recalled product. All 9 affected instruments were returned to Pioneer Surgical, reconditioned and sent back to the field. Customers with questions were instructed to call 906-226-4489.For questions regarding this recall call 906-226-4812. |  | 9 | 9 | Nationwide Distribution including TX, CA, IL, and FL. | Terminated | September 30, 2014 | 473 | 9 | Computer
Z-2611-2014 | 68959 | No | PFF | High throughput genomic sequence analyzer for clinical use. | Immunology | Toxicology | 510(K) Exempt | Illumina Worklist Manager (IWM) (software v1.0.15), a component of Illumina MiSeqDx Platform. |  |  | July 28, 2014 | 2014 | Illumina Inc | A software component of the MiSeqDx instrument, called the Illumina Worklist Manager(IWM) v1.0.15, fails to perform as intended | Device Design | Illumina notified affected customers via phone, email and Product Quality Notification letter dated July 28, 2014 to describe the software anomaly. The letter identified the affected product, problem and actions to be taken. For questions contact your local support team or Illumina Technical Support at techsupport@illumina.com. |  | 35 | 35 | Worldwide Distribution - US (nationwide) and the countries of: France, Germany, Italy, Norway, UK and United Arab Emirates. | Terminated | November 12, 2014 | 107 | 35 | Computer
Z-0911-2014 | 67358 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | GE Healthcare Precision 500D, Precision MPI &amp; Precision RXI X-ray imaging systems. |  |  | December 12, 2013 | 2013 | GE Healthcare, LLC | GE Healthcare has recently become aware of a potential safety issue involving the Precision 500D, Precision MPI &amp; Precision RXI X-ray imaging systems overhead LCD Video Monitor Suspension (Single monitor &amp; Dual monitor).There is a potential for the locking nut to loosen out of the LCD Video Monitor Suspension yoke. This will not lead to an immediate fall of the suspended monitor since it will | Device Design | The firm, GE Healthcare, sent an "Urgent Medical Device Correction" letter dated December 11, 2013 to its consignees. The letter was addressed to Hospital Administrators / Risk Managers, Managers of Radiology/ Cardiology and Radiologists / Cardiologists. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction and Contact Information. The consignees were instructed to immediately contact your local GE Healthcare Service Representative if you observe any inappropriate performance or excessive movement of the monitor suspension yoke.GE Healthcare will correct all affected systems at no cost to you. A GE Healthcare service representative will contact you to arrange for this correction.If you have any questions or concerns regarding this notification, please call one of the following numbers: United States: 800-437-1171, Japan: 0120-055-919, for other countries contact your local GE Healthcare Service Representative. |  | 816 | 816 (722 US, 94 OUS). | Worldwide distribution: US (nationwide) including Guam and DC but excluding CT, DE, HI and RI; and internationally to: AUSTRALIA, BULGARIA, CANADA, CHILE, DOMINICAN REPUBLIC, GERMANY, HONDURAS, IRELAND, ITALY, KUWAIT, LEBANON, MALTA, NAMIBIA, PORTUGAL, ROMANIA, RUSSIA, SAUDI ARABIA, SLOVAKIA, SPAIN, SWITZERLAND, TAIWAN, TURKEY, UNITED KINGDOM, VENEZUELA, and YEMEN. | Terminated | September 11, 2014 | 273 | 816 | Computer
Z-2609-2014 | 68821 | No | DOP | Mass spectrometer for clinical use. | Clinical Chemistry | Toxicology | 510(K) Exempt | AB SCIEX API 3200MD" LC/MS/MS System with software: |  |  | July 15, 2014 | 2014 | Ab Sciex | Software- AB SCIEX API 3200MD LC/MS/MS SYSTEM with MultiQuantMD 3.0 software may provide incorrect data for patient diagnosis | Software design | AB SIEX issued letter on 7/15/14, notifying customers to implement this field correction, a temporary fix to be implemented immediately.Additionally, a software update is in development and is expected to be available in approximately 8 weeks. Upon availability, AB SCIEX will send a DVD with instructions on how to install the new software update. At that time, outlined temporary actions will no longer be necessary.If you have any questions regarding this recall, potential hazards or actions to be taken please feel free to contact AB SCI EX at +1 289 982 2531.Please confirm receipt of this letter by signing and faxing back the attached Acknowledgement Form. |  | 6 | 6 units | FL, NC, OK, PA, TX | Terminated | October 01, 2014 | 78 | 6 | Computer
Z-2599-2014 | 69058 | No | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | CLAW PLATE 20mm INTERAXIS CHARLOTTE F&A; SYSTEM, REF 40S20020, 1 EACH, Rx ONLY, NON STERILE |  |  | July 30, 2014 | 2014 | Wright Medical Technology, Inc. | Plates manufactured from an incorrect raw material. | Device Design | Wright Medical sent an Urgent Medical Device Recall letter dated July 25, 2014, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to examine their inventory for the affected product and immediately quarantine all affected product and return to Wright Medical as soon as possible. For return instructions customers were instructed to contact Customer Service at 800-238-7117. Customers were also asked to complete the attached response form, confirming receipt of the notice and the status of the product(s) on hand and return by fax to 901-867-7401.Customers with questions were instructed to call 901-290-5662.For questions regarding this recall call 800-238-7117. |  |  |  | Worldwide Distribution - USA (Nationwide) and Internationally to Canada Australia, Great Britain, Germany, Spain, Netherlands, France. | Open, Classified |  |  |  | Computer
Z-2597-2014 | 68982 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | RADIANCE v. 2.4X, 2.5X and 2.6X |  |  | July 28, 2014 | 2014 | Radiometer America Inc |  | Software design | Radiometer sent an Product Notification letter dated July 31, 2014 to affected customers via FedEx. The letter identified the affected product, problem and actions to be taken. The letter informed customers that local RADIOMETER IT specialist will contact customers in order to schedule the upgrade to v. 2.74 or 3.5.2. All RADIANCE versions below 2.7X will be upgraded to v. 2.74 or 3.5.2.For questions regarding this letter or the Recall, please contact Radiometer America Technical Support at 1-800-736-0600 opt.4. |  |  |  | Worldwide Distribution - US (nationwide) in the states of AZ, CA, MO, MS, TX and UT. and the countries of: Austrialia, Austria, France, UK, Italy, Japan, Spain, India, and the Netherlands. | Terminated | October 15, 2014 | 79 |  | Computer
Z-2583-2014 | 68655 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Natural-Knee System Patella Bushings. Orthopedic surgical instrument. Part Number: 6290-00-702. |  |  | July 10, 2014 | 2014 | Zimmer, Inc. | Zimmer received a trend of complaints indicating corrosion of product. | Device Design | Zimmer issued an Urgent Medical Device Recall-Lot Specific notification via e-mail/letter dated July 19, 2014, to all affected customers. The notification identified the product the problem and the action needed to be taken by the customer. All distributors were notified via electronic mail. Hospital risk managers and surgeons, as well as distributors with affected inventory were notified via courier. Distributors were sent a letter identifying the issue and their responsibilities. These responsibilities include locating and removing the affected product in their territory as well as identifying hospitals and surgeons who have previously used the affected product. Distributors are to return on-hand affected products to Zimmer, and ensure all of their products are accounted for using the form provided in the letter. Hospital risk managers and surgeons were provided with a letter identifying the issue and their responsibilities. These responsibilities include locating the affected product, quarantining the product, and returning it to their Zimmer sales representative. If you have questions or concerns regarding this recall please call the customer call center at 1-877-946-2761. |  | 8638 | 8638 distribution events | Worldwide Distribution - US Distribution and the countries of : Argentina, Australia, Brazil, France, Germany, India, Japan, Spain, Switzerland, and Venezuela. | Terminated | August 20, 2015 | 406 | 8638 | Computer
Z-2576-2014 | 68878 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | PERSONA The Personalized Knee System Constrained Posterior Stabilized Tibial Articular Surface Provisional TOP with the following: L 3-5 CD; L 6-9 CD; L 3-5 EF; L 6-9 EF; L 10-11 EF; L 6-9 GH; L 10-12 GH; L 10-12 J; R 3-5 CD; R 6-9 CD; R 3-5 EF; R 6-9 EF; R 10-11 EF; R 6-9 GH; R 10-12 GH; and R 10-12 J. |  |  | August 07, 2014 | 2014 | Zimmer, Inc. | Reports of breakage of the Persona Knee Trial Articular Surface Provisionals (TASPs). The Persona primary knee surgical technique has been updated. This field action correction is to ensure that all distributors and users of this instrument are aware of the enhanced surgical techniques. | Device Design | On 8/7/2014, Zimmer notified all distributors via electronic mail. Distributors that have received affected inventory are also notified via courier. Hospital risk managers and surgeons are also notified via courier.Hospital risk managers and surgeons were provided with a notification identifying the issue and their responsibilities. These responsibilities include ensuring the affected personnel are aware of the contents of the notification. Distributors were sent a notification identifying the issue and to provide a copy of the notification sent to surgeons and risk managers. The distributors will also be instructed as needed to assist with identifying a comprehensive list of surgeons. |  | 77737 | 77,737 units total | Worldwide Distribution: US (nationwide) including states of: AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, WA, WI, WV, and WY; and countries of: AUSTRALIA, INDIA, JAPAN, KOREA, SINGAPORE, AUSTRIA, BELGIUM, FRANCE, GERMANY, ISRAEL, ITALY, LUXEMBOURG, NETHERLANDS, SAUDI ARABIA, SOUTH AFRICA, SPAIN, SWITZERLAND, UNITED KINGDOM, and UNITED ARAB EMIRATES. | Terminated | January 27, 2015 | 173 | 77737 | Computer
Z-2544-2014 | 68277 | No | OYC |  |  | Clinical Chemistry | PMA | Medtronic MiniMed Paradigm Insulin Infusion Pumps, |  |  | March 13, 2014 | 2014 | Medtronic MiniMed Inc. | Medtronic is recalling the Medtronic MiniMed Paradigm Insulin Infusion Pumps because they have received reports that users have accidentally programmed the pump to deliver the maximum bolus amount. | Device Design | The firm, Medtronic, sent an "IMPORTANT MEDICAL DEVICE SAFETY INFORMATION" letter dated March 13, 2014 to its customers. The letter described the product, problem and actions to be taken. The letter informs the customers that because accidental button pressing errors may occur it is important that customers always confirm the insulin dose flashing on the display is correct before pressing ACT to start delivery. The letter informs the customers of what safety settings will minimize the risk of an accidental programming error and what actions to be taken. Should you have any questions about this letter or concerns with your keypad, please do not hesitate to contact the 24 hour helpline at (866) 222-7304 for a replacement. |  | 559374 | 559,374 units total (444,374 units in US) | Worldwide Distribution-Nationwide in US | Terminated | April 07, 2015 | 390 | 559374 | Computer
Z-2581-2014 | 68985 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Monaco |  |  | August 18, 2014 | 2014 | Elekta, Inc. | Incorrect Delivery of Composite VMAT Fields - When DICOM is exporting an arc plan (VMAT, Dynamic Conformal or 3D Static Arcs) and using the Composite Field Sequencing (CFS) feature, the DICOM exported plan will not match the approved treatment plan. | Software design | The firm, Elekta, issued an "URGENT IMPORTANT FIELD SAFETY NOTICE (IFSN 382-01-MON-001)" to all affected customers on August 18, 2014. The notice describes the product, problems and actions taken. Confirmation of receipt is to be completed by October 31, 2014. The problem will be resolved in patches to the following releases - Monaco 3.30 and Monaco 5.00. Development is underway with a release date of September 2014 with confirmation of installation of the patch no later than 6 months from release of the patch. An updated notice from the manufacturer is also being sent to affected customers (FCO: 382-01-MON-001 VID:2.0). If you have any queries about this Notice, please contact your local Elekta office or call 770-300-9725 or email: Elekta_Global_PMS@elekta.com. |  | 928 | 928 units | Worldwide distribution: US (nationwide) including states of: AL, AK, AZ, CA, CO, CT, FL, IL, IN, MD, MA, MI, MN, MO, MT, NH, NJ, NM, NY, NC, ND, OH, OK, OR, PA, SD, TN, TX, VA, WA, WI, Puerto Rico; and countries of: Algeria, Angola, Antigua &amp; Barbuda, Australia, Austria, Bahrain, Belarus, Belgium, Brazil, Canada, Chile, China, Colombia, Cyprus, Czech Republic, Denmark, Ecuador, Egypt, Estonia, Finland, France, Germany, Greece, Hong Kong, India, Iraq, Ireland, Israel, Italy, Japan, Kazakhstan, Libya, Lithuania, Malaysia, Malta, Mexico, Morocco, Myanmar, Namibia, Netherlands, New Zealand, Norway, Pakistan, Peru, Philippines, Poland, Portugal, Romania, Russia, Saudi Arabia, Singapore, Slovenia, South Africa, South Korea, Spain, Sir Lanka, Sweden, Switzerland, Taiwan, Thailand, Turkey, United Kingdom, and Ukraine. | Open, Classified |  |  | 928 | Computer
Z-0926-2014 | 67248 | No | IKZ | Powered patient rotation bed. | Physical Medicine | Physical Medicine | 510(K) Exempt | The Rotoprone therapy System is an advanced patient care system for the treatment and prevention of pulmonary complications. The Rotoprone Therapy System provides up to 62 degrees in both prone and supine positions as well as Trendelenburg/Reverse Trendelenburg positioning. |  |  | September 11, 2013 | 2013 | ArjoHuntleigh Inc. | The lock pin can become stuck in the "in" position during the manual rotation, resulting in the inability to prone a patient | Device Design | ArjoHuntleigh Inc. sent a Technical Bulletin package to all affected service centers and an Urgent Field Safety Notice package to all affected customers. For rental units, a training plan was established for employees. For sold units, the notice contain a revised user manual and label explaining what to do if the lock pin becomes stuck in the "in" position. Customers that purchased the affected product were required to complete and return a Customer Response form to ArjoHuntleigh to confirm receipt and understanding of the field safety notice, and that the required actions were completed. For questions call 630-785-4885. |  | 231 | 231 units | US Nationwide and Middle East | Terminated | January 08, 2015 | 484 | 231 | Computer
Z-0932-2014 | 67167 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | Diagnostic imaging systems (GEMINI LXL Model #882412, TruFlight Select Model #882438, GEMINI TF PET/CT16 Model #882470, GEMINI TF PET/CT 64 Model #882471, GEMINI TF Ready Model #882473 &amp; GEMINI TF Big Bore Model #882476) |  |  | December 04, 2013 | 2013 | Philips Medical Systems (Cleveland) Inc | The PET Reconstruction Server (PRS) database may randomly lock up after the low dose CT acquisition is complete and will not allow the PET acquisition to begin, resulting in an unexpected incomplete study. | Software design | On 12/4/2013, the firm sent Field Safety Noticies to their customers. |  | 386 | 386 units | Product was shipped to the following states: AL, AR, AZ, CA, CO, DC, FL, GA, HI, IA, IL, IN, KY, LA, MA, MD, MI, MN, MO, MS, NC, ND, NE, NJ, NV, NY, NY , OH, OR, PA, TN, TX, VA, VT, WA &amp; WI.Product was also shipped to the following countries: Australia, Belgium, Brazil, Bulgaria, Canada, China, Croatia, Cuba, Denmark, Egypt, France, Germany, Hong Kong, Hungary, India, Indonesia, Italy, Japan, Kuwait, Lebanon, Malta, Mexico, Netherlands, Pakistan, Philippines, Poland, Portugal, Republic of Korea, Russia, Russian Federation, Saudi Arabia, Singapore, Slovakia, South Korea, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, Ukraine, Ukraine, United Kingdom, Venezuela &amp; Vietnam. | Terminated | February 02, 2016 | 790 | 386 | Computer
Z-2574-2014 | 68984 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | FOCAL Sim Workstation, Radiation Treatment Planning System. |  |  | August 18, 2014 | 2014 | Elekta, Inc. | When DICOM is exporting an arc plan (VMAT, Dynamic Conformal of 3D Static Arcs) and using the Composite Field Sequencing (CFS) feature, the DICOM exported plan will not match the approved treatment plan. | Software design | Elekta, Inc. issued an Urgent Important Field Safety Notice letter (IFSN 382-01-MON-001) to all affected customers.. Confirmation of receipt is to be completed by October 31, 2014. The problem will be resolved in a patch to Focal 4.80. Development is underway with a release date of September 2014 with confirmation of installation of the patch for the fix no later than 6 months from release of the patch. An updated notice from the manufacturer is also being sent to affected customers (FCO: 382-01-MON-001 VID:2.0). For questions contact your local Elekta office. |  | 554 | 554 units | Worldwide Distribution - US (nationwide)in the states of AK, CA, FL, MI, MN, MO, NJ, NY, NC, OH, PA, TN, VA,WA and Puerto Rico, Algeria, Antigua &amp; Barbuda, Australia, Austria, Bahrain, Belarus, Brazil, Bulgaria, Canada, Chile, China, Colombia, Cyprus, Czech republic, Ecuador, Egypt, Estonia, France, Germany, Greece, India, Iraq, Israel, Italy, Japan, Kazakhstan, Malaysia, Malta, Mexico, Morocco, Netherlands, New Zealand, Pakistan Portugal, Romania, Russia, Saudi Arabia, Singapore, Solvenia, South Africa, South Korea, Spain, Switzerland, Thailand, Turkey, and United Kingdom | Terminated | August 08, 2017 | 1086 | 554 | Computer
Z-0928-2014 | 67224 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | XiO RTP System. |  |  | November 25, 2013 | 2013 | Elekta, Inc. | When bolus is present, the Effective Depth (with bolus) to the weight point returned on the Source Data Report is incorrect. | Software design | Corrective Action 1 - An important safety notice, LINXIO0018 - Incorrect Effective Depth for Bolused Beams is in the process of being distributed to affected customers. (Release date November 25, 2013). Corrective Action 2 - Develop a fix to the defect in an expeditious timeframe and make available in Patch Release 4.0.03 and Release 5.00.01. |  | 1424 | 1424 | Worldwide Distribution-USA (nationwide) including DC and Puerto Rico and the states of AL, AK, AZ, CA, CO, CT, DE, FL, GA, IL, IN, KS, KY, LA, ME, MD, MI, MN, MS, MO, MT, NE, NV, NJ, NM, NY, NC, ND, OH, OK, OR, PA, RI, SC, SD, TN, TX, QA, WA, WV, WI, and WY, and the countries of Albania, Algeria, Argentina, Austria, Australia, Bosnia, Bulgaria, Belgium, Bahrain, Brazil, Belarus, Botswana, Bangladesh, Canada, China, Columbia, Costa Rica, Croatia, Cuba, Cyprus, Chile, Czech Republic, Ecuador, Egypt, Estonia, France, Finland, Gabon, Germany, Great Britain, Greece, Guatemala, Hungary, India, Indonesia, Italy, Ireland, Israel, Iraq, Jordan, Japan, Kosovo, Kuwait, Lithuania, Libya, Latvia, Morocco, Malaysia, Malta, Mexico, Myanmar, Netherland, New Zealand, Pakistan, Peru, Philippines, Poland, Portugal, Romania, Russia, Saudi Arabia, Serbia, Singapore, Slovenia, Slovakia, South Africa, South Korea, Spain, Sudan, Syria, Taiwan, Thailand, Turkey, Turkmenistan, Ukraine, Venezuela, and Vietnam. | Terminated | December 02, 2015 | 737 | 1424 | Computer
Z-2564-2014 | 68758 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | Hitachi Echelon Oval MRI system is a diagnostic imaging device (one unit per package) and is intended to provide the physician with physiological and clinical information, obtained non-invasively and without the use of ionizing radiation. |  |  | August 18, 2014 | 2014 | Hitachi Medical Systems America Inc | The Gradient Coil was found to have a failure mode that allowed it to overheat and become a burn hazard. | Device Design | Letters will be sent to customers. Hatachi Serivce will be sent to each site to exchange the Gradient Coil Assembly. |  | 39 | 39 systems | Worldwide Distribution: US (nationwide) including states of: CA, DE, FL, ID, KS, KY, MD, MS, NY, OH, and WY; countries of: Japan, Brazil, France, Germany, and Kyrgyzstan. | Terminated | September 14, 2015 | 392 | 39 | Computer
Z-2563-2014 | 68912 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Roche cobas IT 1000 |  |  | July 31, 2014 | 2014 | Roche Diagnostics Operations, Inc. | Potential that test results would be assigned to the wrong Patient ID if cobas IT 1000 receives an ADT ADMIT, TRANSFER or MERGE event message that contains a genuine Patient ID (e.g. 999999), which is the same as what has been configured as the Generic Patient ID (999999 by default). | Software design | Roche sent an Urgent Medical Device Correction letter dated July 31, 2014 via UPS Ground (receipt required) to affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to follow the steps provided in the letter to prevent assigning patient results to an incorrect generic patient ID. Please contact the Roche Support Network Customer Support Center at 1-800-440-3638, prompt 5, 24 hours a day, seven days a week if you have questions about the information contained in this UMDC. |  | 148 | 148 Units | US Nationwide Distribution in the states of: AK, AL, AR, AZ, CA, CT, DE, FL, GA, HI, IA, IL, IN, KS, KY, LA, MD, ME, MI, MO, NC, NE, NH, NJ, NY, OH, OR, PA, PR, SC, TN, TX, VA, VT, WA, WI, and WV. | Terminated | June 16, 2015 | 320 | 148 | Computer
Z-2561-2014 | 68965 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Siemens EasyLink Informatics System V4.0/V5 |  |  | August 13, 2013 | 2013 | Siemens Healthcare Diagnostics, Inc. | May not perform as intended under certain conditions, causing the release of results to the Laboratory Information System (LIS) that should have been held for manual review due to auto-verification rules or the delay/omission of resulttransmission to the LIS. | Software design | Siemens issued an Urgent Medical Device Correction (UMDC) letter on August 13, 2013 to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to review this letter in its entirety and ensure that all operators understand the information presented. The issues described are specific to the features listed in the letter.. If the laboratory is not using a particular feature, then they are not affected by the issues described for that feature. For questions contact your Siemens Customer Care Center or your local Siemens technical supportrepresentative. |  | 1347 | 1347 | US Nationwide Distribution | Terminated | December 02, 2014 | 476 | 1347 | Computer
Z-0933-2014 | 66861 | No | DWJ | Thermal regulating system. | Cardiovascular | Cardiovascular | 510(k) | The Gaymar Medi-Therm device provides a means of regulating patient temperature by supplying temperature-controlled water through a connector hose to accessory Gaymar Hyper/Hypothermia blanket(s)/body wrap(s). The blanket/body wrap provides an interface for heating or cooling the patient. Accessory YSI 400 series patient probe interfaces between the |  |  | December 27, 2013 | 2013 | Stryker Medical Division of Stryker Corporation | Stryker Medical has identified that on some units of the MTA7900, there is a lack of a water temperature display while the unit is in Auto mode. While in Auto mode, only the patient temperature is displayed along with the set point. The water display will be blank during Auto mode. If the user wants to see the water temperature displayed and switches out of Auto mode to Manual mode, then back to | Device Design | On Dec. 27, 2013, Stryker Medical notified their consignees, identifying the product, reason for the recall and associated risks. Consignees were asked . Locate the units listed attach enclosed warning tags to the units listed. For units that do not have the water temperature displayed, advise staff not to switch to manual mode in order to display the water temperature.Stryker will make arrangements to update the software so that the water temperature will be displayed in both Auto and Manual modes. Fax (269-488-8691) or e-mail (ProductFieldAction@stryker.com ) the enclosedacknowledgment to confirm receipt of this notification. Please forward a copy of this letter to the new users and advise us of their new location and let Stryker know if any units have been discontinued. Questions or concerns, should be addressed to Stryker Customer Service at (1-800-327-0770). Monday-Friday 8 a.m-5 p.m. (ET). |  | 1002 | 1002 units | nationwide and worldwide | Terminated | November 04, 2014 | 312 | 1002 | Computer
Z-0937-2014 | 67237 | No | GFO | Partial thromboplastin time tests. | Hematology | Hematology | 510(k) | Hemochron Jr. Citrate APTT Cuvette |  |  | December 19, 2013 | 2013 | ITC-Nexus Dx | ITC has determined that some Citrate APTT cuvettes may recover higher than expected results in normal individuals. | Device Design | The firm, ITC The Point of Care, sent an "URGENT MEDICAL DEVICE RECALL" letter dated December 19, 2013 a long with an attached Health Professional Account Tracking Form via UPS to its customers. The letter describes the product, problem and actions to be taken. The customers were instructed to forward this communication to all those within your organization who need to be aware of this matter including any facility that you forwarded any of the affected product to; check all inventory to determine if you have any of the affected lots listed; if you have any affected lots, stop using them, remove from inventory and return to ITC; complete the attached Health Professional Account Tracking Form and return it within 10 days via fax to: 866-429-3132, email: RA@itcmed.com or mail to: ITC Technical Support, 20 Corporate Place South - Piscataway, NJ 08854 . It is important that you complete the form whether or not you have remaining inventory of these lots. In addition, contact ITC Technical Support immediately if you become aware of an unreported adverse event related to the use of affected J103C cuvettes.ITC will review your response form and will contact you to arrange for anyinventory to be returned to ITC and ITC will replace all unused, returned product.If you have questions, please contact ITC Technical Support: Telephone 800-579-2255 or 732-548-5700, Extension 4011 or E-mail Techsupport@itcmed.com. |  | 25965 | 25,965 curvettes (7,425 cuvettes US; 18,540 cuvettes OUS) | Worldwide Distribution: US (nationwide) and countries of: Austria, Belgium, Canada, Serbia, Germany, Estonia, Italy, Spain and Saudia Arabia. | Terminated | October 30, 2014 | 315 | 25965 | Computer
Z-2243-2014 | 68914 | No | MMX | Embolectomy catheter. | Cardiovascular | Cardiovascular | 510(k) | CloverSnare 4-Loop Vascular Retrieval Snare. |  |  | July 17, 2014 | 2014 | Cook Inc. | This remedial action is the result of six product complaints associated with separation of the snare from the distal tip of the wire. | Device Design | Cook Medical sent an Urgent Medical Device Recall letter dated July 17, 2014, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to review the attached list of affected products and lot numbers shipped to their account, and quarantine any affected product that remains unused. Immediately collect and return all unused affected products to Cook Medical as soon as possible for credit. Complete the attached Recalled Product Reply Form and return to Cook Medical either with the product or separately. Customers were asked to report any Adverse Event to Cook Medical Customer Relations at 800-457-4500 or 1-812-339-2235. |  | 696 | 696 devices total distribution (671 nationwide) | Worldwide Distribution - USA including AK, AL, AZ, CA, CT, DC, FL, GA, HI, IA, ID, IL, IN, KY, LA, MD, ME, MI, MN, MO, MS, NC, NE, NH, NJ, NV, NY, OH, OK, PA, SC, TX, UT, VA, WA, WI, WV and Internationally to:Canada, Austria, Belgium, Denmark, Germany, Great Britain, Ireland, Italy, Spain, Sweden, and Switzerland. | Terminated | June 16, 2015 | 334 | 696 | Computer
Z-0940-2014 | 67027 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Siemens ADVIA Chemistry Systems models 1200, 1650, 1800 and 2400 |  |  | November 06, 2013 | 2013 | Siemens Healthcare Diagnostics, Inc. | The ADVIA Chemistry software systems using certain software versions allows the system to automatically run assay calibrations and quality control (QC). | Software design | Siemens sent an Urgent Medical Device Correct letter, dated November 6, 2013, to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to retain the recall notification letter with their laboratory records and forward to those who may have received the affected product. |  | 3759 | 3759 | Worldwide Distribution and USA Nationwide. | Terminated | August 10, 2015 | 642 | 3759 | Computer
Z-2553-2014 | 68854 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Siemens Dimension EXL Integrated Chemistry System in vitro diagnostic devices intended to measure a variety of analytes in human body fluids. All Serial #s of material #s: 10636929, 10636928, 10486890, 10472176, 10639337, and 10639338. |  |  | July 16, 2014 | 2014 | Siemens Healthcare Diagnostics, Inc. | The Reagent Lid hinge may lose its effectiveness and slowly shift downward during maintenance procedures, potentially resulting in physical contact and injury to the operator. | Device Design | Siemens notified end users of a potential safety issue with the Reagent Lid hinge on the Dimension EXL instrument via an Urgent Medical Device Correction letter, dated July 2014. The letter indicated Customer Service Engineers will inspect and adjust the tension on the reagent lid hinges during every visit and replace hinges which no longer properly adjust. Users were urged to use caution when the Reagent Lid is open and notify Siemens the Reagent Lid is shifting down. |  | 2757 | 2757 | Nationwide and internationally, including Canada and Mexico | Terminated | December 16, 2014 | 153 | 2757 | Computer
Z-2166-2014 | 68670 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | PerkinElmer Specimen Gate Screening Center software, version 1.4 and 1.6; Product Code: 5002-0500 |  |  | June 27, 2014 | 2014 | PerkinElmer Health Sciences, Inc. | Potential for demographic information and test result to be mismatch. | Software design | PerkinElmer sent an Urgent Medical Device Recall letter on July 1, 2014, to all affected customers. The letter informed customers of the recall and provided instructions. Customers were asked to complete the enclosed response form and fax it to 1-330-825-8520. For additional information, customers were instructed to contact their local PerkinElmer representative or email specimen.gate.support@perkinelmer.com.Customers with questions were instructed to call 1-800-321-9632. option 1. |  | 13 | 13 devices | Worldwide Distribution - USA including MA, GA, FL, NV, and Internationally to Canada, Ireland, Italy, Saudi Arabia, Belgium, and UK | Terminated | April 06, 2015 | 283 | 13 | Computer
Z-1061-2014 | 67230 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Synergy XVI |  |  | November 25, 2013 | 2013 | Elekta, Inc. | Potential collision risk when using XVI and external beam shaping devices. | Software design | Elekta sent an Important Field Safety Notice dated November 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customers. Corrective Action #1 - Potential collision risk when using XVI and external beam shaping devices. (Only software version R4.2.1 from this notice are applicable to this recall). Recommendations to Users - Monitor the Actual column of the Table Move Assistant dialog box. If the Actual is blank, do not press the ASU button on the FKP. To correct the error - (1) Restart XVI. (2) Put the patient in their initial setup position. (3) Scan the patient again, and continue the usual workflow. Always monitor the patient when do RATM. When you do Volume View acquisitions with the Apex collimator attached, make sure that the stop angle of the Volume View parameters is in the lower half of the gantry rotation. This decreases the risk of a collision with the patient. Read and follow the advice in the Important Field Safety Notice (200 01 507 081) and acknowledge receipt to your Elekta representative within 30 days. Corrective Action #2 -Permanent Solution - XVI Risk control measure - changes to the software to prevent table movement if the start position has not been recorded correctly. Recommendations for changes to product are - (1) Table zero should default to NIHII not zero. (2) Check that the table zero is valid because calculating absolute table positions. (3) Check prescribed table position is not more than 2.5cm from actual table position when permitting ASU from outside room. Target Release XVI R.4.2.2 estimated for release April 2014.For further questions please call (770) 300-9725. |  | 312 | 312 | Worldwide Distribution: US (nationwide) including Puerto Rico and the states of: AL, AK, AZ, CA, CO, CT, FL, GA, IL, IA, KY, MD, MA, MI, MN, MS, MT, NE, NV, NJ, NM, NY, NC, OH, OK, OR, PA, SC, TN, TX, VT, VA, WA, WV, WI, DC, and Internationally to: Algeria, Australia, Austria, Belarus, Belgium, Brazil, Canada, Chile, China, Cyprus, Czech Republic, France, Germany, Greece, Hong Kong, India, Israel, Italy, Japan, Malaysia, Mexico, Pakistan, Portugal, Russia, Singapore, Slovakia, South Korea, Spain, Taiwan, Thailand, and United Kingdom. | Open, Classified |  |  | 312 | Computer
Z-2147-2014 | 68700 | Yes | LWS |  |  | Cardiovascular | PMA | Boston Scientific ENDOTAK RELIANCEÂż SG, transvenous defibrillation lead, Models 0180 &amp; 0292, Sterile EO. |  |  | June 19, 2014 | 2014 | Boston Scientific CRM Corp | A review of manufacturing test records for Boston Scientific ENDOTAK RELIANCE implantable leads revealed a suspected test data recording error. Specifically, some test results were recorded as "failed" without any other indication of failure. There are no reported injuries from the devices. | Software Manufacturing/Software Deployment | Sales representatives hand delivered a Boston Scientific "Medical Device Retrieval" Letter dated June 19, 2014 or June 20, 2014 when they were retrieving the devices from the hospital shelves. The letter was addressed to Hospital Administrator. The letter described the problem, product being recalled and the retrieval of the device by Boston scientific sales representative. For further questions they can contact their local sales representative or Boston Scientific International Technical Services.Physicians were contacted via telephone starting on July 6, 2014. A Boston Scientific letter dated July 8, 2014 followed the telephone conversation. For further information physicians can contact their local Boston Scientific representative or Technical Services at 1-800-227-3422. |  | 21 | 21 | Worldwide Distribution -US (nationwide) in the states of OH, MN and country of France | Terminated | December 03, 2014 | 167 | 21 | Computer
Z-1980-2014 | 68234 | No | HKI | Ophthalmic camera. | Ophthalmic | Ophthalmic | 510(k) | Non-Mydriatic Auto Fundus Camera. Model AFC-330. An ophthalmic camera for use in capturing images of the retina and anterior segment of the eye. |  |  | June 16, 2014 | 2014 | Nidek Inc | Image taken by AFC-330 has a white spot which may affect diagnosis or evaluation of image. | Device Design | Field Correction notifications were sent to all affected users on June 16, 2014 by trackable method. |  | 185 | 185 cameras | Nationwide Distribution. | Terminated | March 03, 2015 | 260 | 185 | Computer
Z-2029-2014 | 68024 | No | KFM | Nonroller-type blood pump. | Cardiovascular | Cardiovascular | 510(k) | Sarns Disposable Centrifugal Pump with X-Coating, 164275X, is a sterile, single use device. |  |  | April 14, 2014 | 2014 | Terumo Cardiovascular Systems Corporation | During set-up and priming of the bypass circuit, leaks were detected on the centrifugal pumphead. | Device Design | Urgent Medical Device Recall letters, dated April 11, 2014, were sent to customers. The letters identified the affected product, as well as the reason for the correction, the potential hazard, the correction, the affected population, and instructions for customers. Customers are to review the notice and ensure that all users have received a copy. Additionally, customers are to confirm receipt of the communication by returning the attached Customer Response Form as instructed; and, return all effected centrifugal pumps. Terumo CVS will issue a Returned Goods Authorization upon receipt of the Response Form. Questions should be directed to Terumo CVS Customer Service at 800-521-2818. |  |  |  | Nationwide Distribution | Open, Classified |  |  |  | Computer
Z-2037-2014 | 68643 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | GEMINI TF PET/CT16: 882470, GEMINI TF PET/CT 64: 882471, GEMINI TF Big Bore: 882476. A diagnostic imaging system for fixed or mobile installations that combines Positron Emission Tomography (PET) and X-ray Computed Tomography (CT) systems. |  |  | June 18, 2014 | 2014 | Philips Medical Systems (Cleveland) Inc | When beginning a procedure to start the gantry and CT Host, a Close Estop message may appear to the operator, but disappears before action can be taken. If this occurs it results in an inability to close E-Stop and the system is not operational for clinical use. This problem occurs randomly. | Software design | Philips Healthcare is informing the consignee of this issue through a Field Safety Notice (FSN) 88200487-488. The US consignees will receive this notice via certified mail and Philips Key Market representatives will distribute the FSN internationally. Additionally, a Field Service Engineer will contact each site to schedule a time to install the software update. |  | 33 | 33 | Distributed in the states of CA, CO, IL, OK, OR and WI. | Terminated | September 16, 2015 | 455 | 33 | Computer
Z-2036-2014 | 68696 | No | MEC | Medical washer-disinfector. | General Hospital | General Hospital | 510(K) Exempt | Reliance 1227 Cart and Utensil Washer/Disinfector |  |  | June 30, 2014 | 2014 | Steris Corporation | STERIS has identified that deformation of the washers door handle pin may occur when the door handle is rotated with sufficient force. Should deformation and/or breakage of the door handle pin occur there is a potential for the washer door to become inoperable. | Device Design | Customers will be notified via onsite visits by STERIS Field Service Technicians to each affected Customer site. These visits will be documented and monitored until completion. |  | 37 | 37 | US Nationwide Distribution - in the states of AL, AR, CA, DE, FL, GA, IL, IA, LA, MA, MI, NJ, NY, OH, OK, OR, SC, TX, VA. and in the country of Canada. | Terminated | July 01, 2015 | 366 | 37 | Computer
Z-1110-2014 | 67427 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | Philips HeartStart MRx Monitor/Defibrillator used |  |  | February 04, 2014 | 2014 | Philips Medical Systems, Inc. | HeartStart MRx Defibrillator Monitor therapy connection could experience accelerated wear causing risk for patients and/or caregivers | Device Design | Philips Medical North America issued on 2/4/14 the Urgent Medical Device Correction Notification/Field Safety Notice informs customers of the issue, identifies details of the units affected, gives instructions on actions to be taken by the customer and identifies what action Philips plans to take to remedy the issue.The correction will consist of free of charge service support for the MRx therapy connection, including inspection of therapy connection, and installation of a stabilizing collar if not already present. If wear is detected, therapy components (including therapy pads/cables, CPR therapy cable, external paddles cable, and internal paddles cable/adapter) within their service life will be replaced.A Philips Healthcare representative will contact customers with affected devices to arrange for service. Philips is asking customers to follow the Action to be Taken by Customer/User section of the Urgent Medical Device Correction notification/Field Safety Notice:You can continue to use your MRx prior to inspection and installation of upgrade hardware by a Philips Authorized Service Provider provided that upon receipt of this notification: Ensure you are performing ongoing therapy connection inspections on all of your MRx devices to detect wear. The HeartStart MRx Instructions for Use, 453564396411 Ed 1, 453564307761 Ed 2, or 453564174011, a previously circulated Therapy Connection Maintenance and Inspection When Used in Transport Addendum, Edition 6, describe how users can identify wear. If wear is detected, remove affected devices immediately from use and contact local Philips representative to arrange service or call 1-800-722-9377. |  | 58466 | 58466 units | Worldwide Distribution: USA (nationwide) and to countries of: AFGHANISTAN ALBANIA ARGENTINA ARUBA AUSTRALIA AUSTRIA AZERBAIJAN BAHRAIN BANGLADESH BELGIUM BOLIVIA BOSNIA AND HERZEGOWINA BRAZIL BRUNEI DARUSSALAM BULGARIA CAMBODIA CHILE CHINA COLOMBIA COSTA RICA CROATIA (local name: Hrvatska) CYPRUS CZECH REPUBLIC DENMARK ECUADOR EGYPT EL SALVADOR ESTONIA ETHIOPIA FINLAND France GERMANY GHANA GREECE GUADELOUPE HONG KONG ICELAND INDIA INDONESIA IRAQ IRELAND ISRAEL ITALY JAPAN JORDAN KAZAKHSTAN KENYA KOREA, REPUBLIC OF KUWAIT LAOSÂż LATVIA LEBANON LIBYAN ARAB JAMAHIRIYA LITHUANIA LUXEMBOURG MACEDONIA MALAYSIA MEXICO MONGOLIA MOROCCO MOZAMBIQUE MYANMAR (Burma)Âż NAMIBIA NEPAL NETHERLANDS NETHERLANDS ANTILLES NEW CALEDONIA NEW ZEALAND NICARAGUA NORWAY OMAN PAKISTAN PALESTINE PANAMA PAPUA NEW GUINEA PARAGUAY PERU PHILIPPINES POLAND PORTUGAL PUERTO RICO QATAR REUNION ROMANIA RUSSIAN FEDERATION SAUDI ARABIA SERBIA SINGAPORE SLOVENIA SOUTH AFRICA SPAIN SRI LANKA SWEDEN SWITZERLAND TAIWAN, PROVINCE OF CHINA TANZANIA, UNITED REPUBLIC OF THAILAND TRINIDAD AND TOBAGO TUNISIA TURKEY TURKMENISTAN UGANDA UKRAINE UNITED ARAB EMIRATES UNITED KINGDOM UZBEKISTAN VENEZUELA VIET NAM YEMEN | Open, Classified |  |  | 58466 | Computer
Z-1120-2014 | 67438 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | HeartStart MRx Defibrillator/Monitor M3538A Lithium Ion Batteries |  |  | February 18, 2014 | 2014 | Philips Medical Systems, Inc. | Philips HeartStart MRx Monitor/Defibrillator, when operating on battery power only, may experience an unexpected shutdown if exposed to elevated levels of electromagnetic interference from RF energy source | Device Design | The firm, Philips Healthcare, sent an "URGENT- Medical Device Correction" letter dated February 2014 and issued a Field Safety Notice on February 18/ 2014 to its customers. The letter and field safety notice informs customers of the issue, identifies details of the units affected, gives instructions on actions to be taken by the customer and identifies what action Philips plans to take to remedy the issue. Philips will replace the affected batteries free of charge.Philips instructed customers to follow the Action to be Taken by Customer/User section of the Field Safety Notice. Fixed, portable, and mobile radio frequency communications equipment can affect the performance of medical equipment. See the HeartStart MRx Instructions for Use for the minimum recommended separation distance between RF communications equipment and the HeartStart MRx. The condition can be avoided by operating the device on AC or external DC power or by the following actions when operating on battery.If you need any further information or support concerning this issue, please contact your local Philips representative or call us at 1-800-722-9377. |  | 43792 | 43,792 batteries | Worldwide Distribution: US (nationwide) and Internationally to: ALBANIA, ARGENTINA, ARUBA, AUSTRALIA, AUSTRIA, AZERBAIJAN, BAHRAIN, BANGLADESH, BELGUIM, BOLIVIA, BRAZIL, CANADA, CHINA, COLOMBIA, CONGO, CROATIA, CZECH REPUBLIC, DENMARK, EGYPT, EL SALVADOR, ESTONIA, ETHIOPIA, FINLAND, FRANCE, GERMANY, HONG KONG, ICELAND, INDIA, INDONESIA, IRELAND, ISRAEL, ITALY, JAPAN, JORDAN, KAZAKHSTAN, KENYA, KOREA REPUBLIC, KUWAIT, LATIVA, LEBANON, LIBYAN ARAB, LITHUANIA, MALAYSIA, MALTA, MEXICO, MONGOLIA, MOROCCO, MOZAMBIQUE, MYANMAR, NEPAL, NETHERLANDS, NEW CALEDONIA, NEW ZEALAND, NICARAGUA, NIGERIA, NORWAY, OMAN, PAKISTAN, PARAGUAY, PERU, PHILIPPINES, POLAND, PORTUGAL, QATAR, ROMANIA, RUSSIAN FEDERAL, SAUDIA ARABIA, SERBIA, SINGAPORE, SLOVENIA, SOUTH AFRICA, SPAIN, SWEDEN, SWITZERLAND, TAIWAN PROVIDENCE, THAILAND, TRINIDAND AND TABAGO, TUNISIA, TURKEY, TURKMENISTANIA, UKRAINE, UNITED ARAB EMIRATES, UNITED KINGDOM, UZBEKISTAN, VENEZUELA, and VIETNAM. | Open, Classified |  |  | 43792 | Computer
Z-2023-2014 | 68399 | No | INN | Mechanical chair. | Physical Medicine | Physical Medicine | 510(K) Exempt | Aquatec Fixed Backrest - accessory for Mobile Shower and Toilet commode Aquatec Ocean, Ocean VIP and Ocean E-VIP family. |  |  | June 25, 2014 | 2014 | Invacare Corporation | The AquatecÂż Fixed Backrest withstands a load of only 250 newtons at uneven load. This may result in the Fixed Backrest becoming detached from the frame. | Device Design | The recall notification will be sent to the two consignees informing them of the recall and to notify their customers. The letter will include a provider letter, customer letter, provider response card and impacted units listing. |  | 465 | 465 units | Distributed in PA and Canada. | Terminated | April 15, 2015 | 294 | 465 | Computer
Z-1093-2014 | 67380 | No | NYQ |  |  | Immunology | PMA | Product Name: HER2 CISH pharmDx Kit |  |  | January 21, 2014 | 2014 | Dako North America Inc. | Dako has initiated a recall of HER2 CISH pharmDxTM Kit because weak red signals have been observed. | Device Design | Dako sent a Recall Notification letter dated January 21, 2014to affected customers via mail. The letter identified the affected product, description of the problem, actions to be taken by the user, communication, and Dako contact information. The letter was accompanied by a "Recall Form (Product Recall Form to customers)". For questions call 805-566-5464. |  | 49 | 49 | Worldwide Distribution - USA Nationwide and in the country of Chile. | Terminated | June 04, 2014 | 134 | 49 | Computer
Z-2026-2014 | 68678 | No | FQO | Operating tables and accessories and operating chairs and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | The STERIS 4085 General Surgical Table |  |  | June 27, 2014 | 2014 | Steris Corporation | STERIS engineering analysis has determined the D1 pressure switch in the hydraulic column manifold does not adequately detect pressure level. If one or more of the table floor lock feet become unlocked due to the loss of hydraulic pressure, there is a potential for the table hand control to incorrectly indicate that the floor lock feet are locked. | Device Design | STERIS sent an Urgent Voluntary Field Correction Noticel letter dated June 27, 2014 to affected Customers.via FedEx with tracking numbers for delivery confirmation. The letter identified the description of the product, description of the problem and the actions to be taken. For questions contact Steris Field Service Dispatch at 1-800-333-8828. |  | 3689 | 3689 | Worldwide Distribution - US (nationwide in the states of AL, AK, AZ, AR, CA,CO, CT, DE, FL, GA, HI, ID, IL, IN, IA, KS, KY, LA, ME, MD, MA, MI, MN, MS MO, MT, NE, NV, NH, NJ, NM, NY, NC, ND, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VT, VA, WA, WV, WI, WY in the countries of Canada, India, Vietnam, China, India, Korea, Mexico, costa Rica, Brazil, Taiwan, Australia, New Zealand, El Salvador, Indonesia, Hong Kong, Singapore, Nepal. | Terminated | September 16, 2015 | 446 | 3689 | Computer
Z-1096-2014 | 67424 | No | MWI | Cardiac monitor (including cardiotachometer and rate alarm). | Cardiovascular | Cardiovascular | 510(k) | Invivo Corporation Expression IP5 Information Portal |  |  | October 03, 2013 | 2013 | Invivo Corporation | Invivo Corporation located at 12151 Research Parkway, Orlando, FL 32826 is recalling their Expression Information Portal (Model IP5). Three different IP5 devices lost audio during use. | Software design | Philips Healthcare sent an "URGENT - MEDICAL DEVICE RECALL" letter dated October 3, 2013 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Customers were instructed to refer to the "Action to be Taken by Customer/User section on the attached pages. A Philips representative will contact the customer regarding their affected devices. All affected devices will have a software update installed in order to correct the problem free of charge. Contact your local Philips representative for further information and support. |  | 271 | 271 | Worldwide Distribution - USA and the countries of Australia, Austria, Denmark, France, Germany, Netherlands, Poland, Singapore, Sweden, Switzerland, United Kingdom. | Open, Classified |  |  | 271 | Computer
Z-1995-2014 | 68580 | No | LDS | Medical washer. | General Hospital | General Hospital | 510(K) Exempt | HLD System, Model 610, Medical Device Cleaning and High Level Disinfection Washer/Pasteurizer. |  |  | January 24, 2013 | 2013 | Cenorin, LLC | The temperature sensor/control system in the HLD Systems Model 610 may provide a false temperature reading. This caused the water to remain at ambient temperature throughout the pasteurization cycle rather than using hot water. | Device Design | Cenorin sent a Notification letter dated December 18, 2013 to all affected customers. The letter identified the affected product, problem and actions to be taken. The letter stated: Cenorin has developed an upgrade to mitigate this issue as well as provide additional process control and system monitoring benefits. Cenorin included a Return Response form with the recall letter to be completed and return. For questions contact HLD Systems Sales Representative for questions. |  | 25 | 25 HLD systems | Worldwide Distribution - US (nationwide): CA, DE, FL, MA, MI, NY, PA, TX, UT, WA, and WI and the countries of Uruguay and Taiwan. | Terminated | February 10, 2015 | 747 | 25 | Computer
Z-1994-2014 | 68593 | No | NAY | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | Patient Side Cart Battery Box used in conjunction with the da Vinci A surgical system. |  |  | April 27, 2012 | 2012 | Intuitive Surgical, Inc. | Patient Side Cart battery boxes may overheat in rare charging conditions, causing battery bulging. | Device Design | No notification was sent. Issue is being addressed during routine maintenance or through complaint process. |  | 2606 | 2606 | Worldwide Distribution. | Terminated | July 07, 2014 | 801 | 2606 | Computer
Z-1113-2014 | 66527 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | TSX-101A Aquilion 32/64 ONE Computed Tomography X-Ray System; TXS-101/R Aquilion RXL Computed Tomography X-Ray System; TSX-301A Aquilion One Computed Tomography X-Ray System; TSX-301B Aquilion Premium Computed Tomography X-Ray System; TSX-302A Aquilion Prime Computed Tomography X-Ray System. |  |  | October 01, 2013 | 2013 | Toshiba American Medical Systems Inc | Toshiba intitiated this recall because their investigation revealed that due to a problem of the X-ray output control software, the X-ray output conditions may not be set correctly. As a result, X-ray exposure under the scan conditions used for contrast study may not start. | Software design | Toshiba sent an Urgent Medical Device Correction letter on October 1, 2013, via USPS return receipt mail and accompanied by a response form. The recall notification letter provided the Device Model, serial number, SID, problem, corrective action, request to customers, and contact information (800) 421-1968 or (800) 521-1968. Toshiba designed and test an upgraded ROM for the SC2 PWB to prevent recurrence of the fault. Service technicians will replace the ROM on site at no cost to the system owners. All repairs are expected to be completed by April 2014 and progress reports will be provided on the effectiveness of the plan.For questions regarding this recall call 714-730-5000. |  | 338 | 338 units | Nationwide Distribution. | Terminated | March 03, 2015 | 518 | 338 | Computer
Z-1124-2014 | 67458 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | MEDITECH Microbiology - Calculator/data processing module, for clinical use. |  |  | January 17, 2014 | 2014 | Medical Information Technology, Inc. | Potential for patient results being removed from LAB/EMR. | Device Design | Beginning on January 17, 2014 MEDITECH distributed a notification to MAGIC and Client Server customers who have the Microbiology product in TEST and/or LIVE only for the releases noted in this correction report. This notification was made via e-mailed task updates that can be printed by the customer. Task updates sent via e-mail are immediately transmitted to the customer.Instructions provided to customers: Task Description: CS LAB DTS 11696/MG LAB DTS 16703 - MIC: Inactive procedure prompt not displaying previous results. We are setting up this task to inform you of a potential patient safety issue that may result in patient results being removed from LAB/EMR. This task will continue to be updated to provide you with the status on this issue and details on the deployment of changes to address. This issue has been determined to impact both Client Server and MAGIC 5.6 and higher. Until the changes to address above are delivered, please refrain from inactivating procedure prompts in the MIC Procedure Dictionary. Although patient results are not permanently deleted from the record, because of the potential impact on a patient should this issue occur and go undetected, we felt it necessary to inform you of the issue. We strongly recommend delivery of the changesto address the issue.Âż We have engaged all necessary resources within MEDITECH to identify and resolve this issue as expediently as possible.Âż CS LAB DTS 11696 - MIC: Inactive procedure prompt not displaying previous resultsÂż MG LAB DTS 16703 - MIC: Inactive procedure prompt not displaying previous resultsÂż Inactivating Procedure Prompts in the MIC Procedure Dictionary causes data to be removed from existing specimens. These results are also removed from the EMR. However, once the prompts are re-activated, the results are again present on the specimens.Âż We strongly recommend delivery of this change to your system to alleviate this potential patient safety concern. Please confirm via this task th |  | 737 | 737 | Worldwide Distribution: US (nationwide) and to countries of: Canada, Ireland, and United Kingdom. | Open, Classified |  |  | 737 | Computer
Z-1960-2014 | 65562 | No | LPS | Urinary homocystine (nonquantitative) test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | ADVIA Centaur HCY Diluent - 10mL Bottle Catalog number 09877493 (124533), SMN 10310370, and Ancillary Pack Catalog number 03302138 (124581), SMN 10378677. Used on the ADVIA Centaur, ADVIA Centaur XP and ADVIA Centaur CP platforms. For in vitro diagnostic use in the quantitative determination of total homocysteine (HCY) in serum or plasma. |  |  | June 25, 2013 | 2013 | Siemens Healthcare Diagnostics, Inc | ADVIA Centaur Homocysteine dilution recovery on the diluent identified the percent recovery for the 1:10 patient sample dilutions lower than reported in the Instructions for Use (IFU). | Device Design | Siemens Healthcare sent an Urgent Field Safety Notice that was sent to all affected Siemens Healthcare Diagnostics customers in the United States on June 25, 2013. The Urgent Field Safety Notice was sent out for distribution to all customers outside the United States on June 24, 2013. These notices inform the customer that the 1:10 diluted patient recovery is lower than reported in the Instructions for Use (IFU). Customers should discontinue use of both manual and on board 1:10 dilutions of samples with the ADVIA Centaur HCY assay. Customers can continue to use a manual or on board 1 :2 dilution for samples with results outside of the assay range. Siemens has confirmed theperformance at the 1 :2 dilution level meets the IFU ranges.Please retain this letter with your laboratory records, and forward this letter to those who may have received this product.We apologize for the inconvenience this situation has caused. If you have any questions, please contact your Siemens Technical Solutions Center or your local Siemens technical support representative.For further questions please call (508) 668-5000. |  | 51 | 51 units | Worldwide Distribution - USA (nationwide) and the countries of : Austria, Belgium, Bulgaria, Croatia, Cyprus, Czech Republic, Denmark, Finland, France, Germany, Great Britain, Greece, Hungary, Ireland, Italy Kazakhstan, Latvia, Liechtenstein, Luxembourg, Netherlands, Norway, Poland, Portugal, Romania, Saudi Arabia, Serbia, Slovakia, Spain, Sweden, Switzerland ,Turkey, Unit.Arab Emir and Vatican | Open, Classified |  |  | 51 | Computer
Z-2146-2014 | 68741 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | ExacTrac is intended to be used to place patients at an accurately defined point within the treatment beam of a medical accelerator for stereotactic radiosurgery or radiotherapy procedures, in order to treat lesions, tumors and conditions anywhere in the body when radiation treatment is indicated. ExacTrac may also be used to monitor the patient position during the treatment. |  |  | June 24, 2014 | 2014 | Brainlab AG | When using multiple isocenters (radiation treatment targets) within a single plan, in certain workflow conditions ExacTrac v.6.0.4 might move the patient to an unintended isocenter position, despite displaying the green "OK" icon. If this anomaly occurs and is not detected by the user, the radiation treatment dose at the linear accelerator may be delivered to the unintended target position. | Software design | Brainlab sent a FIELD SAFETY NOTICE / PRODUCT NOTIFICATION letter dated June 24, 2014, to all affected customers. The letter included instructions for customers to: 1) Avoid using plans with multiple isocenters in ExacTrac 6.0.4 whenever possible, and 2) If plans with multiple isocenters cannot be avoided and such a plan is used in ExacTrac v.6.0.4, then always select "Yes" in the dialog displayed to "keep the previous correction shift" when switching between isocenters and always verify all final treatment positions. Brainlab will provide a software update with this issue solved to affected customers. Brainlab will actively contact you to schedule the update within the next 6 months. Customers with questions can contact Brainlab at 800-597-5911 or via e-mail at us.support@brainlab.com. |  | 15 | 15 systems | Worldwide Distribution - USA including CA, CT, and IL; Internationally to France, Germany, Japan, Netherlands, Russia, and Slovenia. | Open, Classified |  |  | 15 | Computer
Z-1989-2014 | 68528 | No | KPR | Stationary x-ray system. | Radiology | Radiology | 510(k) | GE Healthcare, Revolution XR/d. |  |  | June 04, 2014 | 2014 | GE Healthcare, LLC | GE Healthcare has become aware of a potential safety issue due to a collimator installation error during a service maintenance activity involving the Proteus XR/a and Revolution XR/d X-ray imaging systems. . There was a reported incident of a collimator fall which caused a serious patient injury. | Equipment maintenance | GE Healthcare sent an Urgent Medical Device Correction letter dated June 4, 2014, to all affected consignees. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction and Contact Information. For questions in regard to this recall the consignees in the United States were advised to call 800-437-1171. For other countries contact your local GE Healthcare Service Representative.For questions regarding this recall call 800-437-1171. |  | 1053 | 1053 | Worldwide Distribution - USA including PR. DC, GUAM, all states except VT. Internationally to the following countries:ALBANIA, ALGERIA, ARGENTINA, AUSTRALIA, AUSTRIA, BAHRAIN, BENIN, BOSNIA AND HERZEGOVINA, BOTSWANA, BRAZIL, BULGARIA, BURKINA FASO, CANADA, CHILE, CHINA, COLOMBIA, CZECH REPUBLIC, DOMINICAN REPUBLIC, ECUADOR, EGYPT, ESTONIA, FINLAND, FRANCE, FRENCH POLYNESIA, GERMANY, GREECE, GUATEMALA, HONDURAS, HONG KONG, HUNGRY, ICELAND, INDONESIA, IRAQ, IRELAND, ISRAEL, ITALY, JAMAICA, JAPAN, JORDAN, KAZAKHSTAN, REPUBLIC OF KOREA, KUWAIT, LATVIA, LEBANON, LIBYAN ARAB JAMHIRIYA, LITHUANIA, MALAYSIA, MALTA, MARTINIQUE, MEXICO, MOROCCO, NETHERLANDS, NEW ZEALAND, NICARAGUA, NIGERIA, NORWAY, PAKISTAN, PANAMA, PHILIPPINES, POLAND, PORTUGAL, ROMANIA, RUSSIA, SAUDI ARABIA, SINGAPORE, SLOVAKIA, SLOVENIA, SOUTH AFRICA, SPAIN, SUDAN, SWEDEN, SWITZERLAND, TAIWAN, TANZANIA, THAILAND, TUNISIA, TURKEY, UGANDA, UKRAINE, UNITED ARAB EMIRATES, UNITED KINGDOM, URUGUAY, VENEZUELA, VIETNAM YEMEN. | Terminated | May 06, 2015 | 336 | 4513 | Computer
Z-1979-2014 | 68459 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | CNS-6200 Series Central Nurse Station and accessories. Model Number: CNS-6201 |  |  | June 03, 2014 | 2014 | Nihon Kohden America Inc | Software Version 02-26, when used with the Central Nursing Station (CNS) 6201, (PU-621 RA) may unexpectedly and without warning reboot, resulting in a period of approximately 3 minutes during which time the system does not display patient data and does not emit alarms based on that data should they be appropriate. | Software design | Nihon Kohden sent an Urgent Medical Device Safety Notification and Field Correction letter dated June 3, 2014 to affected customers. Recall notification letter, dated June 3, 2014 was hand-delivered to site by Nihon personnel. The letter provided the following information: product description with codes and software version, reason for recall, field correction information and actions to be taken, MedWatch reporting information, and contact information for recall. |  | 430 | 430 | Worldwide Distribution - US in the state of Oklahoma and in the countries of Bolivia, Chile, Germany, Japan, Korea, Taiwan, Yemen, andSouth Africa. | Terminated | April 28, 2015 | 329 | 430 | Computer
Z-1982-2014 | 68644 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | BrightView XCT: product code 882482, BrightView X- upgraded with the XCT Flat Panel Detector (FPD): product code 882454. Gamma camera for Single Photon Emission Computed Tomography (SPECT) Medical Device for imaging. |  |  | June 23, 2014 | 2014 | Philips Medical Systems (Cleveland) Inc | Flat Panel Display (FPD) failed to remain securely locked in the deployed position. | Device Design | Philips sent an Field Safety Notice dated June 17, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.This Field Safety Notice (FSN) 88200490 and Addendum is intended to inform you about the following:" what the problem is and under what circumstances it can occur" the actions that should be taken by the customer / user in order to prevent risks for patients or users" the actions planned by Philips to correct the problem.If you need any further information or support concerning this issue, please contact your local Philips Healthcare representative:For North America and Canada contact the Customer Care Solutions Center (1-800-722-9377: follow the prompts).thcare is informing the consignee of this issue through a Field Safety Notice 88200490. The US consignees will receive this notice via certified mail and Philips Healthcare Key Market representatives will distribute the FSN internationally. Additionally, a Field Service Engineer will contact each site to schedule a time to conduct the appropriate field safety correction through Field Change Order (FCO) 88200490 that will address the above identified issue. |  | 407 | 407 | Worldwide Distribution - USA including the states of : AZ, CA, CO, DC, FL, GA, IL, MA, MD, MN, MO, MS, MT, NC, NJ, NY,OH, OK, OR, PA, TX, VT, WA,WI. Government Facilities: AZ, CA, CO, FL, HI, IA, KY, MA, MN, MO, NC, ND, NE, NV, OH, OK, SD, TX, and WV. and the countries of : Algeria, Argentina, Australia, Austria, Belgium, Brazil, Canada, Chile, China, Czech Republic, Denmark, Finland, France, Germany, Greece, Indonesia, Ireland, Israel, Italy, Japan, Lebanon, Malaysia, Mexico, Netherlands, New Zealand, Norway, Panama, Poland, Portugal, Russia, Saudi Arabia, Slovakia, South Africa, South Korea, Spain, Sweden, Switzerland, Taiwan, Turkey, UAE and UK. | Terminated | September 24, 2015 | 458 | 407 | Computer
Z-1966-2014 | 68472 | No | NAY | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | Vision Troubleshooting Guide for the da Vinci Standard Surgical System; da Vinci S System and da Vinci Si System. |  |  | May 21, 2014 | 2014 | Intuitive Surgical, Inc. | Multiple updates to User Manuals, Instructions for Use (IFU), Quick Reference Guides (QRG), and Instrument Release Kits (IRK) for use with the da Vinci Surgical System, instruments and accessories. Removal of the Emergency Grip Release Wrench, part number 710142. | Device Design | Intuitive Surgical sent a Field Safety Notice dated June 16, 2014, to all consignees affected by changes in labeling, IFU User manuals and QRG's. Consignees were instructed to take the following actions:Ensure all appropriate personnel are fully informed of this Notification. Forward this letter to your Risk Manager, OR Director, Purchasing Manager, Biomedical Engineering staff and members of your medical staff who perform da Vinci surgery procedures.2. Use Attachment C to identify the labeling that has been updated.3. No manuals or products should be returned. Discard old versions of the User Manuals, IFUs, and QRGs for which you have received updated versions.4. Discard all versions of the Emergency Grip Release Wrench (Part Number 710142). Order additional IRKs from Intuitive Surgical as needed.5. Complete and return the attached Acknowledgement Form to acknowledge that you have received the updated Labeling, IRKs, in-service and have discarded all previous versions of Labeling and the emergency grip release wrench. Please return the completed Acknowledgment Form within two days of the in-service.6. Retain a copy of this Notice and the Acknowledgement Form for your records.If you need further information or support concerning this recall, please contact your Intuitive Surgical Representative or Intuitive Surgical Customer Service at the numbers listed below:North America: 800-876-1310 Option 3 (6 am to 5pm PST)For questions regarding this recall call 408-523-2602. |  | 2014 | All distributed prior to June 2014 | Worldwide Distribution - USA (nationwide) and Internationally to Argentina, Australia, Austria, Belgium, Brazil, Bulgaria, Canada, Chile, China, Colombia, Cyprus, Czech Republic, Denmark, Dominican Republic, Egypt, Finland, France, Germany, Great Britain, Greece, Guadeloupe, India, Indonesia, Ireland, Israel, Italy, Japan, Kuwait, Lebanon, Malaysia, Mexico, Monaco, Netherlands, New Zealand, Norway, Pakistan, Philippines, Indonesia, Portugal, Puerto Rico, Qatar, Romania, Russia, Saudi Arabia, Singapore, Slovak Republic, Slovenia, South Africa, South Korea, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, Uruguay, Venezuela, and Vietnam. | Terminated | May 05, 2015 | 349 | 2014 | Computer
Z-1962-2014 | 68529 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | ACUSON S Family Ultrasound Systems at software versions VC25D, VC30A and VC30B; (S2000, S3000 and S1000 systems) Siemens Medical Solutions USA, Inc. Ultrasound imaging systems. |  |  | June 03, 2014 | 2014 | Siemens Medical Solutions USA, Inc. | There is a potential measurement error with the Auto-Stats (auto statistics) measurement tool on the ACUSON S Family ultrasound system. | Software design | Siemens sent a Customer Safety Advisory Notification letter to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to do the following: To avoid the potential risk of this issue, do not use the Auto-Stats tool; use alternate measurement tools. Only use the Velocity tool on frozen images.Customer Service Engineer will contact the customer to schedule an appointment to replace their existing version of software, free of charge. Customers with questions should contact their local service support person forinformation regarding timelines and status.Customers were instructed to share this information with all personnel withintheir organization who need to be aware of this issue.For questions regarding this recall call 650-694-5135. |  | 790 | 790 units | Worldwide Distribution - USA including CA, MO, VA, PA, WV, and Internationally to Angola, Argentina, Australia, Austria, Belarus, Belgium, Bolivia, Brazil, Canada, Chile, Colombia, Czech Republic, Denmark, Dominican Rep., Ecuador, Egypt, France, Germany, Greece, India, Iraq, Israel, Italy, Japan, Kazakhstan, Malaysia, Mexico, Morocco, Netherlands, Norway, P.R. China, Philippines, Poland, Portugal, Republic Korea, Romania, Russian Fed., Saudi Arabia, Singapore, Slovakia, Slovenia, South Africa, Spain, Sweden, Switzerland, Taiwan, Turkey, U.A.E., Ukraine, United Kingdom. | Open, Classified |  |  | 790 | Computer
Z-1949-2014 | 68393 | No | BZD | Noncontinuous ventilator (IPPB). | Anesthesiology | Anesthesiology | 510(k) | DeVilbiss Model DV5x Series CPAP units. For use in treating Obstructive Sleep Apnea (OSA) by means of application of positive air pressure. Used in home and clinical environments. |  |  | May 13, 2014 | 2014 | DeVilbiss Healthcare LLC | A limited number of DV5x Series CPAP units were manufactured and shipped without the serial number programmed into the unit, causing the unit serial number on LCD display to appear as "RUN-TEST". | Software Manufacturing/Software Deployment | DeVilbiss Healthcare sent an Urgent: Medical Device Correction notice dated May 13, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customers. What to do if you have an affected DV5x Series CPAP unit:1. Verify your inventory of DV5x Series CPAPs and isolate any units that are on the included serial number list, these units will need to have the serial number corrected.2. Please contact DeVilbiss Healthcare Customer Service to either request a Sales Representative to correct the issue on site or return the unit back to DeVilbiss to be corrected.3. If you have already distributed any DV5x Series CPAPs included in this serial number list to end users, a patient communication has been included with this notification - a. Insert your contact information at the bottom of the "User Medical Device Correction Letter" included in this notice.b. Forward a copy of the "User Medical Device Correction Letter'' to each user towhom you have provided an affected DV5x Series CPAP unit.If you have one or more of the affected DV5x Series CPAP unit in your possession (as identified in the serial number list) and have any questions regarding the information provided in this communication, please contact DeVilbiss Healthcare Customer Service at 1-800- 338-1988 from 8:00AM-5:00PM EST. [Note: If you have already spoken with a DeVilbiss Healthcare Representative about this issue, there is no need to call unless you have further questions.] |  | 1809 | 1809 | Worldwide Distribution - USA (nationwide) and the countries of Colombia, Honduras, Iran, Thailand and Uruguay. | Terminated | December 02, 2014 | 203 | 1809 | Computer
Z-1958-2014 | 68347 | No | OCC | Respiratory viral panel multiplex nucleic acid assay. | Microbiology | Microbiology | 510(k) | Simplexa" Flu A/B &amp; RSV Direct, Model Number: MOL2650. |  |  | May 15, 2014 | 2014 | Focus Diagnostics Inc | Focus Diagnostics is providing an urgent safety notice for a correction to the labeling for Simplexa" Flu A/B &amp; RSV Direct (MOL2650). Focus Diagnostics received some customer complaints of Simplexa Flu A/B &amp; RSV Direct assays with sporadic false signals, which may result in a higher false results rate due to a potential software spectral matrix and Direct Amplification Disc storage. | Device Design | Focus Diagnostics sent an Urgent Safety Notice Correction letter dated May 15, 2014, to all affected customers. The notification, informed customers of the product, the problem, and the action to be taken by the customer. To decrease the likelihood of obtaining any false results caused by either manually entering a spectral matrix or storing the DAD improperly after each use, please follow the recommended warnings and precautions provided in the revised and attached Simplexa Flu A/B &amp; RSV Direct (PI.MOL2650 Rev. D) package insert.1. The spectral matrix must be installed in each 3M Integrated Cycler and should not be changed unless an updated QR code for the instrument is provided by Focus Diagnostics. The spectral matrix is unique to each 3M Integrated Cycler. The spectral matrix was provided with the 3M Integrated Cycler instrument on the cover of the 3M Integrated Cycler Hardware Manual. If the matrix label will not scan or cannot be found, contact Focus Diagnostics. The contact information is on the last page of this document. 2. Not installing or changing the spectral matrix can result in false results.3. After each use store DAD discs flat with the numbered foil side up.Customers were instructed to acknowledge that they have received this notification by signing the enclosed acknowledgement form and email the form to DxTS@focusdx.com or fax back to Focus Diagnostics Technical Services at 562-240-6526 within 10 business days. The Notice of Correction and Acknowledgement Form was sent to customers on 05/15/2014 via email and USPS mail. Additional customers were notified on 05/29/2014 via email and USPS mail.For questions regarding this recall call 714-220-1900. |  | 614 | 614 kits | Nationwide Distribution | Terminated | December 31, 2014 | 230 | 614 | Computer
Z-1963-2014 | 68506 | No | LHI | Intravascular administration set. | General Hospital | General Hospital | 510(k) | EXACTAMIX 1200 Compounder, model numbers 1200-DY, 1200-DX, 1200-DYR and 1200-DXR; EM1200 DY Display, Replacement EM 1200 DY Display. EXACTAMIX 2400 Compounder model numbers 2400-DY, 2400-DX, 2400-DYR and 2400-DYX. Baxa Corporation. An automated pumping system that compounds multiple sterile ingredients into a finished solution in a single patient IV administration bag. |  |  | June 10, 2014 | 2014 | Baxter Healthcare Corp. | If the universal ingredient (UI) in an active configuration is changed using the Configuration Editor, a flush of the outlet pump tube will not be initiated by the software. It could result in the original UI remaining in the tube which could be delivered into the next bag. | Software design | An Urgent Device Correction communication via 1st class mail was mailed on July 14, 2014, informing customers of a potential issue if the universal ingredient in an active configuration is changed using the Configuration Editor rather than following the normally prescribed and supported methods as outlined in the Operator Manual. |  | 1758 | 1758 | US Nationwide | Open, Classified |  |  | 1758 | Computer
Z-1060-2014 | 66287 | Yes | BSZ | Gas machine for anesthesia or analgesia. | Anesthesiology | Anesthesiology | 510(k) | McKesson Anesthesia Care |  |  | March 15, 2013 | 2013 | MCKESSON TECHNOLOGIES INC. | There was an occurrence where the patient case data did not match the patient data when the case was recalled in the Anesthesia Care Record (ACR) in that it included data from another case. | Software design | McKesson sent a Clinical Alert dated March 15, 2013 to all affected customers. Additionally, phone calls were placed to each customer, followed up by an email to provide them written copies of the communication and Clinical Alert and obtaining acknowledgement that they have read and understand the issue and preventive actions to take. Customers with questions were instructed to contact McKesson Customer Support at 800-442-6767 (option 3).For questions regarding this recall call 404-338-3556. |  | 10 | 10 | Nationwide Distribution including CA, GA, IN, KY, MD, MI, OH, OK, and WA. | Terminated | May 01, 2014 | 412 | 10 | Computer
Z-1944-2014 | 68394 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | syngo Lab Data Manager, SMN 10800057, 10803189, and 10804573. |  |  | May 20, 2014 | 2014 | Siemens Healthcare Diagnostics, Inc. | Certain configuration parameters may cause the reference ranges to not hold results as expected and release them to the Laboratory Information System (LIS) that should have been held for manual review. | Device Design | An urgent medical device correction notice, dated May 20, 2014, was sent to end users to provide instructions to avoid potential risks associated with the software defect, until changes can be made by a Siemens technical representative. |  | 336 | 336 | Worldwide Distribution. US Nationwide, Canada, Australia, Belgium, Czech Republic, Germany, Denmark, Finland, France, United Kingdom, Italy, Netherlands, Portugal, Slovakia, Singapore, Japan, and Spain. | Terminated | August 10, 2015 | 447 | 336 | Computer
Z-1081-2014 | 67217 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | RT Desktop |  |  | November 25, 2013 | 2013 | Elekta, Inc. | During the treatment table, gantry, and collimator calibration procedures, it is possible to type incorrect values or invert a value. | Software design | Elekta sent an Important Field Safety Notice November 2013 to all affected customers. Corrective Action #1 - Release of Important Field Safety Notice 200 01 502 041 - "Recommended quality assurance after a calibration procedure" Released 25 November, 2013. Important Field Safety Notice is in the process of being distributed to customers advising them of the recommended Quality Assurance checks to avoid/detect the problem. The permanent solution to this problem is currently been developed. Corrective Action #2 - Permanent Solution - All users of Desktop Pro software will be informed of a mandatory upgrade to Integrity 1.2 upon its release which will prevent calibration from being inverted.For questions regarding this recall call 770-300-9725. |  | 489 | 489 | Worldwide Distribution - USA including AL, CA, CT, GA, FL, IL, IN, KS, KY, LA, MD, MA, MI, MN, MO, NJ, NY, OH, PA, TN, TX, VA, and Internationally to Argentina, Austria, Bangladesh, Belgium, Botswana, Brazil, Bulgaria, Canada, Chile, China, Colombia, Cuba, Czech Republic, Egypt, Finland, France, Germany, Greece, Hungary, India, Indonesia, Iran, Israel, Italy, Japan, Lebanon, Mexico, Morocco, Netherland, Nigeria, Pakistan, Panama, Peru, Philippines, Portugal, Russia, Serbia, Solvenia, South Africa, South Korea, Spain, Sweden, Switzerland, Syria, Taiwan, Thailand, Turkey, United Kingdom, Venezuela, Vietnam, and Yemen. | Terminated | August 13, 2015 | 626 | 489 | Computer
Z-1920-2014 | 68503 | No | JXG | Central nervous system fluid shunt and components. | Neurology | Neurology | 510(k) | Medtronic Duet External Drainage and Monitoring System, Interlink Injection Sites, Catalog No. 46913. |  |  | June 09, 2014 | 2014 | Medtronic Neurosurgery | Medtronic Neurosurgery is recalling the Medtronic Duet External Drainage and Monitoring Systems because the patient line tubing in the Duet External Drainage and Monitoring Systems may become disconnected from the patient line stopcock. | Device Design | An Urgent Medical Device Recall letter, dated 6/9/14, was sent to all their customers who purchased the Medtronic Duet External Drainage and Monitoring Systems. The letter informs the customers of the problem identified and the actions to be taken. Customers are instructed to cease use of the affected product, return any used units, account for any product used and verify that connections are in the condition as stated in the Instructions for Use for product currently in use. Additionally, customers are to complete the Customer Product Accountability Form. Customers with questions about the recall letter are instructed to contact Bob Shokoohi at (805) 571-8725. |  | 3026 | 3,026 units | Worldwide Distribution -- USA, Australia, Belgium, Canada, Croatia, Cyprus, Denmark, Egypt, Finland, France, Greece, Iran, Israel, Italy, Kuwait, Netherlands, Norway, Poland, Puerto Rico, Romania, Russian Federation, Saudi Arabia, Singapore, South Africa, Switzerland, Turkey, and United Kingdom. | Terminated | January 14, 2015 | 219 | 10752 | Computer
Z-1934-2014 | 68457 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | NexGen Complete Knee Solution, MIS Total Knee Procedure, Stemmed Tibial Component, Precoat, Fixed Bearing, Size 4. Part of the NexGen systems of semi-constrained, non-linked condylar knee prosthesis. |  |  | May 19, 2014 | 2014 | Zimmer, Inc. | Zimmer is initiating a voluntary recall of specific identified lots due to the potential that the threads may be out of specification. | Equipment maintenance | Zimmer sent an Urgent Medical Device Recall letter June 2014 to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to review the notification and ensure affected personnel are aware of the contents. Locate and quarantine any affected product and notify their Zimmer sales representative. Their Zimmer sales representative will remove the recalled product from their facility. For patients that previously had this product implanted, it is recommended that they continue their normal operative follow up routine. Customers with questions were instructed to call 1-877-946-2761. |  | 41180 | 41,180 total | Worldwide Distribution - USA (nationwide) and Internationally to Australia, China, Hong Kong, India, Japan, Korea, Malaysia, Nicaragua, Singapore, and Taiwan. | Terminated | October 29, 2015 | 528 | 41180 | Computer
Z-2043-2014 | 68632 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | Symbia E camera system. Emission Computed Tomography System used to detect or image the distribution of radionuclides in the body or organ, using the following techniques: planar imaging, whole body imaging, tomographic imaging for isotopes with energies up to 588 keV. Nuclear medicine diagnostic device. |  |  | April 23, 2014 | 2014 | Siemens Medical Solutions USA, Inc. | Siemens Medical Solutions USA, Inc. has received reports indicating the collimator cart rear casters may become loose. | Device Design | A Customer Safety Advisory Notice dated April 23, 2014 was sent to every direct account. The letter included instructions for direct accounts to discontinue use of the collimator cart and notify the local service representative if they notice or suspect tha one or more casters are loose (which would be indicated by shaking while the cart is in motion). Otherwise, a service representative will be contacting the direct account within the following 3 months to schedule a retention bracket installation. Direct accounts with questions can contact Siemens Medical Solutions USA, Inc. at 800-888-7436. |  | 531 | 531 systems | Worldwide Distribution- US: Nationwide (AK ; AL ; AR ; AZ ; CA ; CO ; CT ; DE ; FL ; GA ; HI ; IA ; ID ; IL ; IN ; KS ; KY ; LA ; MA ; MD ; ME ; MI ; MN ; MO ; MS ; MT ; NC ; ND ; NE ; NH ; NJ ; NM ; NV ; NY ; OH ; OK ; OR ; PA ; RI ; SC ; SD ; TN ; TX ; UT ; VA ; VT ; WA ; WI ; WV ; WY) + District of Columbia + Puerto Rico; FOREIGN: Albania ; Algeria ; Angola ; Argentina ; Australia ; Austria ; Azerbaijan ; Bahrain ; Bangladesh ; Belarus ; Belgium ; Botswana ; Brazil ; Bulgaria ; Canada ; Chile ; China ; Colombia ; Costa Rica ; Croatia ; Cuba ; Czech Republic ; Dominican Rep. ; Ecuador ; Egypt ; El Salvador ; Finland ; France ; Georgia ; Germany ; Ghana ; Greece ; Hong Kong ; India ; Indonesia ; Iran ; Iraq ; Ireland ; Israel ; Italy ; Jamaica ; Japan ; Jordan ; Kazakhstan ; Kuwait ; Macedonia ; Malaysia ; Mexico ; Morocco ; Myanmar ; Netherlands ; New Zealand ; Niger ; Nigeria ; Norway ; Oman ; Pakistan ; Palestinian Territory ; Peru ; Philippines ; Poland ; Portugal ; Qatar ; Romania ; Russia ; Saudi Arabia ; Serbia ; Singapore ; Slovakia ; Slovenia ; South Africa ; South Korea ; Spain ; Sri Lanka ; Sweden ; Switzerland ; Syria ; Taiwan ; Thailand ; Trinidad and Tobago ; Tunisia ; Turkey ; Ukraine ; United Arab Emirates ; United Kingdom ; Uruguay ; Vietnam ; Yemen. | Open, Classified |  |  | 2989 | Computer
Z-2044-2014 | 68570 | No | EFB | Dental handpiece and accessories. | Dental | Dental | 510(k) | Lynx TM20 Torquemaster Low Speed Handpiece MTI Dental Products |  |  | May 07, 2012 | 2012 | MTI Precision Products LLC. | Sharp edge on the body/housing of the Lynx TM20 TorqueMaster Low Speed Handpiece. | Software design (manufacturing process) | MTI notified their customers via e-mail on 5/712 and 5/8/12. MTI sent end-users recall letters on or about 5/9/12. |  | 134 | 134 devices | US Nationwide Distribution in the states of CA, FL, IL, KY, MA, MD NJ, NY, and RI. | Terminated | March 12, 2015 | 1039 | 134 | Computer
Z-1108-2014 | 67481 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | M3536A HeartStart MRx with Wireless Link Upgrade (861289) |  |  | February 18, 2014 | 2014 | Philips Medical Systems, Inc. | Wireless Link may be unable to transmit data when configured for a specific data flow. | Software design | Philips Healthcare sent an Urgent Medical Device Correction letter to all consignees on February 18, 2014. The Letter informs customers of the issue, identifies details of the units affected, gives instructions on actions to be taken by the customer and identifies what action Philips plans to take to remedy the issue.Philips is initiating a correction to affected devices. Philips will release a new version of Wireless Link software V2.0.0.3 sub1 that corrects the issue. The correction will be provided free of charge to all units affected by this issue. A Philips Healthcare representative will contact customers with affected devices to arrange for installation of the software upgrade.Customers with questons were instructed to call 1-800-722-9377.For questions regarding this recall call 800-722-9377. |  | 220 | 220 devices | Worldwide Distribution - USA including AZ, CA, FL, IL, IN, KS, KY, MA, MD, MI, MN, MO, NC, NY, OH, OK, OR, PA, SC, TN, TX,and VA and Internationally to Australia, Norway, and Spain. | Terminated | November 12, 2015 | 632 | 220 | Computer
Z-1111-2014 | 67441 | No | OOI | Instrumentation for clinical multiplex test systems. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | cobasÂż PCR Urine Kit 100 PKT IVD |  |  | January 06, 2014 | 2014 | Roche Molecular Systems, Inc. | Numerous complaints have been filed that leaking cobas Âż PCR media 4.3 mL IVD have been received from different lots of thecobas Âż PCR Urine kits, cobas Âż PCR Female Swab kits, and cobas Âż PCR media kits. | Device Design | Roche Molecular Diagnostics Inc. sent an Urgent Medical Device Correction letter/Fax-back form dated January 6, 2014. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to follow the test-skpecific Instruction For Use for handling of cobas PCR media 4.3 ml IVD. The use of protective equipment should be worn as indicated in the labeling. Inspection each tube and/or blister packet for evidence of leakage prior to use. If leakage of the cobas PCR media 4.3 mL IVD is observed, do not use the tubes. Discard the affected tubes per local guidelines and contact Roche Support Network Customer Support Center at 1-800-526-1247. Complete the attached fax form and fax it to 1-855-695-8564. File this Urgent Medical Device Correction (UMDC) for future reference. Customers with questions were instructed to call 1-800-526-1247.For questions regarding this recall call 908-253-7200. |  | 84 | 84 kit batches | Nationwide Distribution | Terminated | August 07, 2014 | 213 | 84 | Computer
Z-1059-2014 | 67231 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Synergy XVI |  |  | November 25, 2013 | 2013 | Elekta, Inc. | XVI can incorrectly calculate the target position of the treatment table. | Software design | Elekta sent an Important Field Safety Notice dated November 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customers. Potential collision risk when using XVI and external beam shaping devices released November 25, 2013 (Only software versions R4.5, R4.5.1, R4.6 and R5.0 are applicable to this recall). (All software versions R4.6 in this notices are applicable to this recall). Recommendations to Users - Monitor the Actual column of the Table Move Assistant dialog box. If the Actual is blank, do not press the ASU button the FKP. To correct the error - (1) Restart XVI. (2) Put the patient in their initial setup position. (3) Scan the patient again, and continue the usual workflow. Always monitor the patient when you do RATM. When you do Volume View acquisitions with the Apex collimator attached, make sure that the stop angle of the Volume View parameters is in the lower half of the gantry rotation. This decreases the risk of a collision with the patient. Read and follow the advice int he Important Field Safety Notices and acknowledge receipt to your Elekta representative within 30 days. Corrective Action #2 - Permanent Solution - XVI Risk control measure - changes to the software to prevent table movement if the start position has not been recorded correctly. Recommendations for changes to product are - (1) Table zero should default to NIHII not zero. (2) Check that the table zero is valid because calculating absolute table positions. (3) Check prescribed table position is not more than 2.5cm from actual table position when permitting ASU from outside room. Users of software version 4.5, 4.5.1 and 4.6 will receive a mandatory upgrade to XVI 5.0. The estimated release date for the fix for the XVI 5.0 is December 2014.For further questions call ( 770 ) 300-9725. |  | 1028 | 1028 | Worldwide Distribution - USA (nationwide) including Puerto Rico and the states of AL, AK, AZ, AR, CA, CT, FL, GA, ID, IL, KY, LA, MD, MA, MI, MN, MO, NH, NJ, NM, NY, NC, ND, OH, OK, OR, PA, RI, SD, TN, TX, VT, VA, WA, WV, WI, WY and DC., and the countries of Algeria, Australia, Austria, Bahrain, Belarus, Belgium, Brazil, Bulgaria, Canada, China, Columbia, Croatia, Cuba, Cyprus, Czech Republic, Denmark, Ecuador, Estonia, Finland, France, Germany, Greece, Hong Kong, Hungary, India, Iraq, Ireland, Israel, Italy, Japan, Jordan, Malaysia, Mexico, Morocco, Netherlands, New Zealand, Norway, Pakistan, Poland, Portugal, Romania, Russia, Saudi Arabia, Singapore, Slovenia, South Africa, South Korea, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, United Kingdom and Venezuela | Open, Classified |  |  | 1028 | Computer
Z-2139-2014 | 68790 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | VITROS 3600 Immunodiagnostic System --- For use in the in vitro quantitative, semi-quantitative and qualitative measurement of a variety of analytes of clinical interest, using VITROS Immunodiagnostic Products Reagents. The VITROS Software is a part of the VITROS 3600, 4600, &amp; 5600 Systems. The VITROS 3600, 4600, &amp; 5600 Systems share the same software base. VITROS Sample Programming consists of Sample Identification (SID), test requests, and patient demographics. Patient demographics include patient name, address, date of birth, gender, physician, and patient identification (PID). |  |  | July 07, 2014 | 2014 | Ortho-Clinical Diagnostics | Ortho Clinical Diagnostics has identified an anomaly with VITROS System Software Version 3.0 and below. Internal testing confirmed that under very specific conditions, incorrect patient demographics were associated with a sample identification (ID) number. | Software in the Use Environment | On 7/07/2014, Customer Letter (Ref. CL2014-191) was sent via FedEx overnight courier or ORTHO PLUS e-Communications and/or US Postal Service Priority Mail (for PO Boxes only) to US consignees (OCD Direct, Drop-Ship and US Federal Government) to inform them of the issue and advise them of the required actions. On 7/07/2014, Distributor Letter (Ref. DL2014-191 was sent via US Postal Service Priority Mail and/or FedEx overnight courier to US Distributor consignees instructing them to notify their consignees of the issue and required actions. Foreign affiliates were informed by e-mail on 7/07/2014 of the issue and instructed them to notify their consignees of the issue and required actions. For questions, please contact Customer Technical Services at 1-800-421-3311. |  | 678 | Domestic: 142, Foreign: 536 | Worldwide Distribution-USA (nationwide) including Puerto Rico and the countries of Australia, Brazil, Canada, Chile, China, Colombia, England, France, Germany, India, Italy, Japan, Mexico, Panama, Singapore, Spain and Venezuela. | Open, Classified |  |  | 5568 | Computer
Z-2131-2014 | 68657 | No | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Zimmer Periarticular Locking Plate System Distal Radial Volar Plate Lateral Column (1st generation stainless steel only). Orthopedic fixation plate. For temporary internal fixation and stabilization of osteotomies and fractures. |  |  | June 10, 2014 | 2014 | Zimmer, Inc. | A complaint was received in which it was reported that an 8-hole plate from lot 62276042 had an enlarged hole that is unable to lock the screw, which resulted in the screw head pulling through plate hole. | Device Design | Zimmer sent an Urgent Medical Device Recall letter dated June 17, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customers.They instructed their distributors to do the following: 1. Please locate, quarantine, and return the affected product (part number 00-2358-014-04; lot 62276042) from all inventories.2. Please note that any hospitals that received direct shipments of this product from Zimmer will be sent a copy of the Risk Manager and Surgeon letter directly. Additionally, Risk Managers and Surgeon that were identified via invoice data will also receive a letter directly. 3. Please review the list of hospital and surgeons to be provided to your territory via email.4. Provide the name, address, and phone number of any additional hospitals or surgeons that may have received or used the affected product, or confirm that the list provided is complete.5. Complete and return Attachment 1.6. Ensure that a copy of the Risk Managers and Surgeons letter is left with the facilities from which product is removed by providing them with a copy of the letter, if they are not identified as a facility receiving a letter directly from Zimmer.7.If after reviewing this notification you have questions or concerns, please call 1-877-946-2761.They instructed the Surgeons and Risk Managers to do the following: 1. Review the notification and ensure affected personnel are aware of the contents.2. Inspect inventory for part number 00-2358-014-04; lot number 62276042. Quarantine the device and contact your Zimmer sales representative to remove and return the affected device. 3. If after reviewing this notification you have questions concerning this notification, please call 1-877-946-2761. |  | 28 | 28 units | Worldwide Distribution - US including the states of CA and OH., and the countries of BRAZIL, CHINA, KOREA and TAIWAN. | Terminated | January 09, 2015 | 213 | 28 | Computer
Z-2134-2014 | 68737 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Patient Data Manager 2.0 (Content manager 2.0, Patient Browser 4.0, DICOM Viewer 2.0) Brainlab Digital Lightbox. A system for the display of medical images. Model numbers 25100-05 and 25100-06. The software can transfer images to and from picture archiving and communication systems (PACS), file servers, or removable storage media. |  |  | June 27, 2014 | 2014 | Brainlab AG | The following defects have been identified for data loaded with Patient Data Manager version 2.0 (including subversions 2.0.0, 2.0.1 and 2.0.2): 1) When loading non-square pixel image data with PDM 2.0, the data set is not correctly transferred and displayed in the corresponding Brainlab navigation or planning software. When a data set is loaded, the pixels of the data set are sorted in order to | Software design | FIELD SAFETY NOTICE / PRODUCT NOTIFICATION letters dated June 27, 2014 were sent to all direct accounts. The letters instruct product users to not upload image data sets with non-square pixels and to verify with their radiology department the modalities of the different imaging devices (e.g. CT, MRT, PET, ultrasound) from which data may be loaded onto a Brainlab device. Brain lab will provide a software update with this issue solved to affected customers with a tentative planned time line of availability being July 2014. Customers with questions can contact Brainlab by telephone at 800-597-5911 or via e-mail at us.support@brainlab.com. |  | 154 | 154 systems | Worldwide Distribution-USA (nationwide) including the states of AL, AK, AZ, CA, FL, GA, IA, IL, KS, MD, MO, MS, NC, NY, OH, PA, SD, TX, VA, VT, WA, and WI, and the countries of Australia, Austria, Belgium, Brazil, Chile, China, Ecuador, France, Germany, Japan, Netherlands, New Zealand, Norway, Qatar, Russia, Spain, Switzerland, Taiwan, United Arab Emirates, and United Kingdom. | Open, Classified |  |  | 154 | Computer
Z-1068-2014 | 66893 | No | OMP | Powered suction pump. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Renasys EZ PLUS, a Canister Component of of Renasys EZ Negative Pressure Wound Therapy System |  |  | November 14, 2013 | 2013 | Smith &amp; Nephew Inc. | Modification of the bacterial overflow guard (filter) and related changes to product labeling. | Device Design | The correction strategy is as follows:" An Urgent Medical Device Correction Notice has been prepared to be distributed to all US consignees via Federal Express delivery for confirmation of the delivery." A copy of the Urgent Medical Device Correction Notice is to be posted on Smith &amp; Nephew's websites, URL corporate, http://www.smith-nephew.com/news-and-media/, and myrenasys.com, concurrent with the distribution of the Notice to consignees." US distribution of the original design of Canister Kits codes (66800912, 66800913, 66801066, 66800423 and 66800058) subject to this correction ceased November 1, 2013." The correction has been implemented on a change being affected basis. The modified design is being distributed to US customers effective November 1, 2013. No product will be removed from the market." Periodic progress reports to FDA, as required per 21 CFR 806 will be submitted monthly until completion of the correction actions.. |  | 6871375 | 6,871,375 canisters | Worldwide Distribution: USA (nationwide) including states of: AL, AZ, CA, CO, CT, FL, GA, HI, IA, IL, IN, KY, LA, MA, MI, MN, MO, MS, MT, NC, NH, NV, NY, OH, OR, PA, RI, SC, SD, TN, TX, VA, WA, WV, and Puerto Rico; and the countries of: Austria, Australia, Belgium, Switzerland, China, Cyprus, Germany, Denmark, Spain, Finland, France, UK, Greece, Israel, Italy, Malta, Netherlands, Norway, Paraguay, Russia, Saudi Arabia, Sweden, Thailand, Turkey, South Africa, Mexico, and Canada. | Terminated | August 04, 2014 | 263 | 13670515 | Computer
Z-2130-2014 | 68697 | No | CCW | Rigid laryngoscope. | Anesthesiology | Anesthesiology | 510(K) Exempt | RUSCH, TruLite, One-Piece Disposable Laryngoscope Handle Blade Combination, Non-Sterile, Rx Only, Distributed by: Teleflex Medical. |  |  | June 30, 2014 | 2014 | Teleflex Medical | There is a possibility the handle of the laryngoscope may heat up and melt the end cap. | Device Design | Consignees were notified via letter on/about 06/30/2014. |  | 6246 | 6,246 units | U.S. Nationwide | Terminated | March 21, 2017 | 995 | 6246 | Computer
Z-2121-2014 | 68799 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Horizon Medical Imaging |  |  | May 05, 2014 | 2014 | Mckesson Medical Immaging | The firm has identified a software issue which may cause a discrepancy between the index of images expected by the application, and the actual content of the folder storing the images. | Software design | McKesson sent an Urgent Field Safety Notice dated May 5, 2014, to all affected customers informing them of the issue and advising them of the immediate actions to take to reduce the risk to the patient. These actions include - Users ensure that all expected images are display in the Image Viewer and PACS Administrators should engage McKesson Support for all issues with studies. McKesson is also providing a software solution to HMI systems that may be potentially affected by the problem to prevent future recurrence. Customers were instructed to contact McKesson Support as soon as possible at 1-800-663-2533 to acknowledge that they have read and understood the Field Safety Notice, and to arrange installation of the software update. Customers with questions were instructed to contact Customer Support and ask to speak to their Support Manager.For questions regarding this recall call 404-338-3556. |  | 659 | 659 units | Worldwide Distribution - USA including AL, AK, AZ, AR, CA, CO, CT, FL, GA, HI, ID, IL, IN, IA, KS, KY, LA, ME, MD, MA, MI, MN, MS, MO, MT, NE, NV, NH, NM, NJ, NY, NC, ND, OH, OR, PA, RI, SC, TN, TX, VT, VA, WA, WV, WI, WY,and France, Canada, Puerto Rico, Guam, Great Britain, Ireland, Australia. | Terminated | April 13, 2017 | 1074 | 659 | Computer
Z-2122-2014 | 68470 | No | FNL | AC-powered adjustable hospital bed. | General Hospital | General Hospital | 510(K) Exempt | InTouch Critical Care Bed, Stryker Medical. Model Numbers 2130, 2131, 2140, 2141, 2151, 2152, 2155, and 2156. |  |  | June 04, 2014 | 2014 | Stryker Medical Division of Stryker Corporation | Stryker has received complaints from the field alleging instances where the In Touch beds are hoisted or lifted from the floor and the casters have fallen off. The investigations into these casters have shown that in these cases the caster has broken on the top of the caster stem where the brake rod inserts into the caster. If more than one caster on a unit is damaged, brake functions can be affec | Device Design | On 06/30/1014 the firm sent "Urgent Medical Device Notification" letters to their customers. |  | 11632 | 11,632 units | Worldwide Distribution -USA (nationwide) including DC and the states of AK, AL, AR, AZ, CA, CO, CT, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI &amp; WV, and the countries of Argentina, Brazil, Canada, Chile, China, EMEA, Greece, Italy, Japan, Korea, Latin America, Mexico, Poland, Singapore, South Africa, Taiwan, and UK. | Terminated | March 14, 2016 | 649 | 11632 | Computer
Z-1078-2014 | 67401 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | GE Healthcare Precision 500D and Proteus XR/a equipped with Wireless Digital Radiography Option (WDR1). |  |  | January 20, 2014 | 2014 | GE Healthcare, LLC | GE Healthcare has recently become aware of a software issue associated with patient selection from the worklist on your WDR1(Precision 500D and Proteus XR/a equipped with Wireless Digital Radiography Option) that may impact patient safety.Intermittently when a large number of patient exams are retrieved from the HIS/RIS system, the exams when displayed on the WDR1 are not properly displayed. I | Software design | GE Healthcare sent an "Urgent Medical Device Correction" letter dated January 20, 2014. The letter was addressed to Hospital Administrators / Risk Mangers, Radiology Department Mangers and Radiologists. The letter described the Safety Issue, Affected Product Details, Safety Instructions, Product Correction and Contact Information ( In USA call 800-437-1171, For other countries contact your local GE Healthcare Service Representative). |  | 359 | 359 (333 US, 26 OUS). | Worldwide Distribution - USA (nationwide) including the states of AL, AK, AZ, AR, CA, CO, DC, FL. GA, HI. IA, ID, IL, IN, LA, KY, ME, MD, MA, MI, MN, MS, MO, NE, NV, NH, NJ, NM, NY, NC, OH, OK, OR, PA, TN ,KY, TX, UT, VA, WA, WV, WI and WY., and the countries of ARGENTINA, AUSTRALIA, CANADA, FRANCE, IRELAND, ITALY, JORDAN, ROMANIA, RUSSIA, UNITED ARAB EMIRATES and UNITED KINGDOM. | Terminated | November 03, 2014 | 287 | 359 | Computer
Z-2116-2014 | 68586 | No | FDS | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | EG-3630U Ultrasound Video Gastroscope |  |  | June 12, 2014 | 2014 | Pentax Medical Company | Aspiration needles used in combination with Ultrasound Gastroscopes can derail from the accessory elevator at the distal end of the endoscope. This may result in patient injury if the Instruction for Use is not followed carefully by the user. | Device Design | Pentax Medical sent an Urgent Medical Device Field Safety Corrective Action dated June 12, 2014 to affected customers. The letter identified the affected product, problem and actions to be taken. Pentax also posted their notification on their website. For questions regarding this action please feel free to contact Paul Silva, RecallCoordinator atTel: 800-431-5880 ext 2064 (8:30am  5:00pm Mon-Fri, EST)Fax: 201-799-4063Email : paul.silva@pentaxmedical.com |  | 531 | 531 units | US Nationwide Distribution | Terminated | November 17, 2015 | 523 | 531 | Computer
Z-2114-2014 | 68378 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | GE Healthcare, CARESCAPE Monitor B650. A multi-parameter patient monitor intended for use in multiple areas and intrahospital transport within a professional healthcare facility. |  |  | July 07, 2014 | 2014 | GE Healthcare, LLC | Possible sound loss associated with the CARESCAPE Monitor B650. When an external display is connected to the CARESCAPE Monitor B650 using a USB extender cable with an active USB hub, the system may experience a sound loss including alarm sounds from the patient monitor. This may occur if the external display is powered off and on during normal operation of the monitor. | Software design | GE Healthcare "Urgent Medical Device Correction" letter #36102 dated July 7, 2014 to affected customers. The letter was addressed to Healthcare Administrator / Risk Manger, Chief of Nursing and Director of Biomedical Engineering. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction and Contact Information. For questions they can contact Technical Support at 1-800-558-7044 or your local Service Representative. |  | 21799 | 21,799 (7,635 units US, 14,164 units OUS) | Worldwide distribution. US nationwide including Puerto Rico, ZAMBIA, YEMEN, VIETNAM, VENEZUELA, URUGUAY, UNITED KINGDOM, UNITED ARAB EMIRATES, TURKEY, THAILAND, UNITED REPUBLIC OF TANZANIA, TAJIKISTAN, TAIWAN, SYRIA, SWITZERLAND, SWEDEN, SRILANKA, SPAIN, SOUTH AFRICA, SLOVENIA, SLOVAKIA, SINGAPORE, SAUDI ARABIA, RUSSIAN FÂżDÂżRATION, ROMANIA, QATAR, PORTUGAL, POLAND, PHILIPPINES, PERU, PARAGUAY, PANAMA, PAKISTAN, OMAN, NORWAY, NICARAGUA, NEW ZEALAND, NETHERLAND, NAMIBIA, MYANMAR, MOROCCO, MEXICO, MAURITIUS, MALAYSIA, LITHUANIA, LIBYAN ARAB JAMAHIRIYA, LEBANON, LATVIA, KUWAIT, REPUBLIC OF KOREA, KAZAKHSTAN, JORDAN, Italy, JAPAN, JAMAICA, ISRAEL, IRELAND, IRAQ, INDONESIA, INDIA, ICELAND, HUNGARY, HONG KONG, HONDURAS, GUATEMALA, GREECE, GERMANY, GEORGIA, FRANCE, FINLAND, ESTONIA, EL SALVADOR, Egypt, ECUADOR, DENMARK, CZECH REPUBLIC, CYPRUS, CROATIA, COSTA RICA, COLOMBIA, CHINA, CHILE, CANADA, BULGARIA, BRUNEI, BRAZIL, BOLIVIA, BERMUDA, BELGIUM, BANGLADESH, BAHRAIN, BAHAMAS, AZERBAIJAN, AUSTRIA, AUSTRALIA, ARGENTINA. | Open, Classified |  |  | 21799 | Computer
Z-2110-2014 | 68756 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | QLAB 10.1 Core Module PN 453561728001. QLAB Quantification Software is a software application package designed to view and quantify image data acquired on Philips Medical Systems ultrasound products. |  |  | July 09, 2014 | 2014 | Philips Ultrasound, Inc. | The QLAB a2DQ and aCMQ features have a defect that may use incorrect measurement values under specific conditions. | Software change control | Philips Medical Systems notified customers of the voluntary recall by sending a certified URGENT-Medical Device Correction letter, dated 2014 JUL 09, to all US and Canadian consignees. Philips representatives will distribute the letter internationally. Follow-up letters to wholesale customers will be provided and wholesalers are requested to forward copies of recall letters and response forms to their customers. Philips Medical Systems will also have Customer Reply Cards returned for tracking purposes. Customers are directed to call local Philips representative or Philips Customer Service at 1-866-767-7822. |  | 59 | 59 units total (12 units in the US and 47 units outside the US) | Distributed in the states of CA, IA, KY, MD, NY, PA, UT, VA, WA and WI.Also distributed in Australia, Belgium, Canada, India, Italy, Japan, South Korea, Netherlands, New Zealand, Spain, Sweden, and Switzerland. | Terminated | November 06, 2015 | 485 | 59 | Computer
Z-2111-2014 | 68769 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Siemens syngo Imaging XS, version VA70A or higher |  |  | June 30, 2014 | 2014 | Siemens Medical Solutions USA, Inc | Siemens has become aware of the potential malfunction when using syngo Imaging XS, version VA70A or higher. In some cases a miscalculation may occur for the grey scale values in the functions Region of Interest , Pixel Lens , Edge enhancement and Histograms . This potential malfunction may occur for images with a pixel depth greater than 12 bit. All values calculated for images with a pixel d | Software design | Siemens sent a Customer letter dated June 30th, to all affected customers to inform them of a new software version VA70B_0114 that will be introduced to all syngo Imaging XS customers with version VA70A or higher to solve a potential issue regarding miscalculations done for the grey scale values in the functions "Region of Interest", "Pixel Lens", "Edge Enhancement", and "Histograms".For questions regarding this recall call 610-219-6300. |  | 141 | 141 | Nationwide Distribution | Terminated | February 24, 2015 | 239 | 141 | Computer
Z-1074-2014 | 67307 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | Single and Dual use Emission Computer Tomography gamma cameras (BrightView XCT, BrightView X &amp; BrightView SPECT) Philips Healthcare Systems, Cleveland, OH. |  |  | December 19, 2013 | 2013 | Philips Medical Systems (Cleveland) Inc | Three problems were reported to the firm:When positioning for a relative 180 degree non circular acquisition the second (bottom) detector was not positioned correctly causing an inappropriate rotation orbit calculation. When the acquisition was started and the detectors began to move into position, the incorrect calculation resulted in the detectors coming in contact with the patient.As | Software design | On 12/19/2013 the firm sent Field Safety Notices to their customers. |  | 1143 | 1,143 units | Worldwide Distribution-USA (nationwide) including DC and Puerto Rico and the states of AK, AR, AZ, CA, CO, CT, DE, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MI, MN, MO, MS, MT, NC, ND, NE, NJ, NM, NV, NY, OH, OK, OR, PA, SC, SD, TN, TX, VA, VT, WA, WI, WV &amp; WY, and the countries of Algeria, Argentina, Australia, Austria, Bahrain, Belgium, Brazil, Canada, China, Czech Republic, Denmark, Dominican Republic, Egypt, Finland, France, Gabon, Germany, Greece, India, Indonesia, Iran, Ireland, Israel, Italy, Japan, Korea, Lebanon, Lithuania, Malaysia, Mexico, Netherlands, New Zealand, Norway, Oman, Palestine, Panama, Philippines, Poland, Portugal, Russia, Saudi Arabia, Singapore, Slovakia, Slovenia, South Africa, Spain, Sweden, Switzerland, Syrian Arab Republic, Taiwan, Thailand, Turkey, Ukraine, United Arab Emirates, United Kingdom &amp; Viet Nam. | Terminated | January 12, 2015 | 389 | 1143 | Computer
Z-1079-2014 | 66430 | No | KWT | Shoulder joint metal/polymer non-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Trabecular Metal" Reverse Shoulder System |  |  | October 17, 2013 | 2013 | Zimmer, Inc. | As a result of complaints for disassociation of the TM Reverse Glenosphere from the Baseplate, Zimmer updated the related surgical technique and released a video to provide additional instructions to users on how to properly assemble the Glenosphere to the Baseplate. | Device Design | In November and December 2012 Sales staff were provided updated surgical technique and associated video to share with current customers, surgeons, and risk managers. October 17, 2013 Distributors ( via E-mail),and Risk Managers and Surgeons (via traceable courier) were sent written Urgent Recall notifications of these labeling changes. Current customers are identified through an invoice search and distribution history for the affected devices for the past 12 months, since 9/1/12. |  | 73002 | 73,002 units | Worldwide Distribution-USA (nationwide) and the countries of Germany, India, Malaysia, China, Singapore, Korea, Taiwan, and Brazil. | Terminated | September 12, 2014 | 330 | 73002 | Computer
Z-1080-2014 | 67175 | No | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | Siemens Artis zee Ceiling Systems. |  |  | December 16, 2013 | 2013 | Siemens Medical Solutions USA, Inc | During the lifetime of Artis zee Ceiling systems, there is an increased wear due to potential sub-optimal cable guidance that can lead to cable damage. | Device Design | In December 2013, a recall notification letter was sent to Users. |  | 284 | 284 | Nationwide Distribution. | Terminated | March 09, 2015 | 448 | 284 | Computer
Z-2084-2014 | 68440 | No | KWA | Hip joint metal/metal semi-constrained, with an uncemented acetabular component, prosthesis. | Orthopedic | Orthopedic | PMA | RECLAIM DISTAL TAPERED Distal Stem, various sizes. Taper Protector Sleeve. RECLAIM Revision Hip System Orthopedic Manual Surgical Instrument, used to ensure the locking taper surface of the Distal Stem is protected during proximal preparation. |  |  | June 03, 2014 | 2014 | DePuy Orthopaedics, Inc. | The product can be difficult to remove from the Distal Stem both out of the package andafter proximal reaming. | Device Design | DePuy Synthes sent an Urgent Information - Device Correction Notice letter dated June 24, 2014, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to:1. Immediate Remedy: Discard the RECLAIMÂż Taper Sleeve Protector instrument before stem implantation.2. Changes to the literature and IFU: Documents are revised as follows:" RECLAIMÂż Revision Hip Systems Surgical Technique # DSUS/JRC/0614/0200  pages 11, 12, 13, 16,30, 34, 35 (figure numbering from pages 11 to 27 were also revised): Removed information regardingtaper sleeve protector." RECLAIMÂż Revision Hip Systems Surgical Technique FAQ # DSUS/JRC/0614/0202  pages 1 and 2:Removed information regarding taper sleeve protector." RECLAIMÂż Revision Hip Systems IFU #090200799 REV E  Page 5: Verbiage will be revised andimplemented as soon as possible, depending on regulatory approvals.3. Change: The company consulted with design surgeons and developed a strategy to eliminate the taper sleeve protector from the system. As a result of removing the taper sleeve protector, the company isevaluating whether there is a need for a protective component to be used during the distributionCustomers with questions were instructed to call 574-371-4917.For questions regarding this recall call 574-371-4577. |  | 3623 | 3,623 | Worldwide Distribution - USA including AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, SC, SD, TN, TX, UT, VA, VT, WA, WI &amp; WV and Internationally to AUSTRALIA, AUSTRIA, CANADA, CHILE, CHINA, CZECH REPUBLIC, DENMARK, FRANCE, GERMANY, IBERIA, IRELAND, ISRAEL, ITALY, NEW ZEALAND, NORWAY, POLAND, SINGAPORE, SLOVENIA, SOUTH AFRICA, SPAIN, SWITZERLAND &amp; UK. | Terminated | May 17, 2017 | 1079 | 14904 | Computer
Z-1090-2014 | 67429 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | HeartStart MRx Monitor/Defibrillator; Model(s), Catalogue, or Code Number: M3535A, M3536A, M3536J, M3536MC |  |  | February 18, 2014 | 2014 | Philips Medical Systems, Inc. | A component of the MRx Processor Board may be susceptible to damage from electrostatic discharge (ESD), which can disrupt ECG and SpO2 functionality leading to an inability to perform. | Device Design | The firm, Philips, sent an "URGENT - Medical Device Correction" letter dated February, 2014 with an attached Field Safety Notice to each known affected customer on 2/18/14. The letter describes the issue, identifies details of the units affected, gives instructions on actions to be taken by the customer and identifies what action Philips plans to take to remedy the issue.The customers were instructed to continue to use your MRx prior to receiving the hardware upgrade; if a failure has been detected that may disable some ECG and SpO2 functionality during patient care, treat your patient per existing protocols and if the condition persists, take the device out of use and call for service from Philips authorized service personnel.Philips is initiating a correction to affected devices. A hardware upgrade will be provided free of charge for all units affected by this issue. A Philips Healthcare representative will contact customers with affected devices to arrange for installation of the hardware upgrade.If you have questions regarding this notification or need any further information or support, please contact your local Philips representative or call us at 1-800-722-9377. |  | 39128 | 39,128 | Worldwide Distribution: US (Nationwide) including states of: AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, GU, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, and WY; and Internationally to: Canada, AFGHANISTAN, ALBANIA, ARGENTINA, ARUBA, AUSTRALIA, AUSTRIA, AZERBAIJAN, BAHRAIN, BELGIUM, BOSNIA AND HERZEGOWINA, BRAZIL, BRUNEI DARUSSALAM, BULGARIA, CHILE, CHINA, COLOMBIA, CROATIA (local name: Hrvatska), CYPRUS, CZECH REPUBLIC, DENMARK, ECUADOR, EGYPT, ESTONIA, FINLAND, FRANCE, GERMANY, GREECE, GUADELOUPE, HONG KONG, ICELAND, INDIA, INDONESIA, IRELAND, ISRAEL, ITALY, JAPAN, JORDAN, KENYA, KOREA, REPUBLIC OF, KUWAIT, LATVIA, LEBANON, LIBYAN ARAB JAMAHIRIYA, LITHUANIA, LUXEMBOURG, MACEDONIA, MALAYSIA, MEXICO, MOROCCO, NAMIBIA, NETHERLANDS, NETHERLANDS ANTILLES, NEW CALEDONIA, NEW ZEALAND, NICARAGUA, NORWAY, OMAN, PAKISTAN, PALESTINE, PANAMA, PAPUA NEW GUINEA, PERU, PHILIPPINES, POLAND, PORTUGAL, PUERTO RICO, QATAR, REUNION, ROMANIA, RUSSIAN FEDERATION, SAUDI ARABIA, SERBIA, SINGAPORE, SLOVENIA, SOUTH AFRICA, SPAIN, SRI LANKA, SWEDEN, SWITZERLAND, TAIWAN, PROVINCE OF CHINA, THAILAND, TRINIDAD AND TOBAGO, TUNISIA, TURKEY, UGANDA, UKRAINE, UNITED ARAB EMIRATES, UNITED KINGDOM, VENEZUELA, VIET NAM, and YEMEN. | Open, Classified |  |  | 39128 | Computer
Z-2085-2014 | 68569 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | GE Healthcare, CARESCAPE Monitor B850, B650 or B450. Patient Monitor. |  |  | June 13, 2014 | 2014 | GE Healthcare, LLC | Partial loss of monitoring with the CARESCAPE Monitor B850, B650 or B450. When connecting the PDM (Patient Data Module) to the CARESCAPE B850, B650 or B450 monitor with software version 2 the PDM may not start communicating to the monitor and loss of the PDM parameters could occur. | Software design | Consignees were sent on 6/13/2014 a GE Healthcare "Urgent Medical Device Correction" letter #36103, 36104 and 36105 dated June 6, 2014. The letter was addressed to Healthcare Administrator / Risk Manger, Chief of Nursing and Director of Biomedical Engineering. The letter described the Safety Issue, Safety Instructions, Affected Products Details, Product Correction and Contact Information. Consignees were advised to contact Technical Support at 800-558-7044 in regard to questions on this recall. |  | 41992 | 41,992 (20,509 US; 21,483 OUS) | Worldwide Distribution: US nationwide, UNITED KINGDOM, UNITED ARAB EMIRATES, TURKEY, TAIWAN, SWITZERLAND, SWEDEN, SPAIN, SOUTH AFRICA, SLOVENIA, SLOVAKIA, SINGAPORE, SERBIA, SAUDI ARABIA, RUSSIAN FEDERATION, ROMANIA, REUNION, QATAR, PORTUGAL, POLAND, PHILIPPINEs, PARAGUAY, PANAMA, NORWAY, NEW ZEALAND, NETHERLANDS, MYANMAR, MOROCCO, MEXICO, MALAYSIA, LITHUANIA, LEBANON, LATVIA, JORDAN, JAPAN, JAMAICA, ITALY, ISRAEL, IRELAND, IRAQ, INDONESIA, INDIA, ICELAND, HUNGARY, GREECE, GIBRALTAR, GERMANY, FRANCE, FINLAND, ESTONIA, EGYPT, ECUADOR, DENMARK, CZECH REPUBLIC, CYPRUS, CROATIA, CHILE, BRUNEI, BERMUDA, BELGIUM, BANGLADESH, AUSTRIA, AUSTRALIA. | Open, Classified |  |  | 41992 | Computer
Z-2088-2014 | 68659 | No | JPA | Multipurpose system for in vitro coagulation studies. | Hematology | Hematology | 510(k) | Hemochron Jr. Whole Blood Microcoagulation System APTT Cuvette. For in vitro diagnostic coagulation test. Catalog number J103. |  |  | June 20, 2014 | 2014 | ITC-Nexus Dx | Test results may be erroneously elevated or report an out-of-range high (OOR-H) error code in non heparinized blood samples. | Device Design | ITC issued an Urgent Medical Device Recall Letters/Response Forms dated June 20, 2014, to End-users and US Distributors were notified. ITC asks that this communication be forwarded to all those within your organization who need to be aware of this matter. If any affected products have been forwarded to another facility, a copy of this communication should be forwarded to them immediately.Customers should check their inventory for the affected lot. The affected lot should not be used and quarantined until return to ITC. The attached form should be completed and returned to ITC within 10 days. ITC will review the form and contact the customer to coordinate any returns. For questions, please contact Technical Support: Telephone 732-548-5700, Extension 4011 E-mail: Techsupport@itcmed.comNormal business hours are Monday through Friday, 8:00 AM to 6:30 PM, Eastern Time.For urgent matters, Technical Support is available 24 hours per day, 7 days a week. |  | 3150 | US- 315 cuvettes; OUS - 2835 cuvettes | Worldwide Distribution - US including the states of Georgia, Virginia and New York., and the countries of Canada, France, Italy, Liechtenstein and Romania. | Terminated | December 11, 2014 | 174 | 3150 | Computer
Z-1091-2014 | 67342 | No | EHD | Extraoral source x-ray system. | Dental | Radiology | 510(k) | The FOCUS is a microprocessor controlled intra-oral x-ray unit with a Load Matching Frequency (LMF) DC generator, which produces dental images on digital and film media. For the digital images the FOCUS is connected to the Sigma USB terminal which is connected to a PC, which reads the image data from a sensor and displays the image on the PC monitor. In order to make installation and service faster and easier the electronics is integrated into the horizontal arm. |  |  | January 14, 2014 | 2014 | Instrumentarium Dental, PaloDEx Group Oy | Instrumentarium Dental has delivered a number of FOCUS 3.0 Intraoral X-ray units in November 2013 with a potential flaw in the mounting hardware. A number of wall mount units were delivered with FOCUS X-ray machines that had not been properly welded at our supplier. As a result of the incomplete welding operation, the wall mount unit on the FOCUS X-ray unit may not properly support the unit as exp | Device Design | The firm, Instrumentarium, sent an "URGENT MEDICAL DEVICE RECALL" letter dated January 14, 2014 was sent to all PaloDEx direct accounts (including subsidiaries). The letter inested recipients to perform a field inspection on potentially defective FOCUS 3.0 X-ray units (units with serial numbers F21221 - F21641 and F21643). If a defective wall mount is discovered during the inspection, the FOCUS X-ray unit must be taken out of service and the wall mount replaced before continuing use of the device. Replacement wall mounts will be made available free of charge for the field repair. A Technical Bulletin including instructions on how to perform the inspection was included with the letter. Direct accounts were further instructed to return the enclosed Acknowledgement Letter confirming receipt of the URGENT MEDICAL DEVICE RECALL letter by mail to Instrumentarium Dental, Jussi Tanskanen, Nahkelantie 160, Fl-04300 Finland; or, via e-mail to Jussi.tanskanen@instrudental.com. The direct accounts are to complete and return the enclosed Confirmation Letter to document that all of the potentially affected units have been inspected and repaired if necessary to Instrumentarium Dental, Jussi Tanskanen, Nahkelantie 160, Fl-04300 Finland; or, via e-mail to Jussi.tanskanen@instrudental.com. A nearly identical URGENT MEDICAL DEVICE RECALL letter dated January 22, 2014 was sent to all US customers. The instructions were the same as noted above in the January 14, 2014 letter to PaloDEx direct accounts.If there are any questions, do not hesitate to contact us at the following: Marketing Director, General dentistryMika.anttila@instrudental.comTel. +358 40 768 3340For more information you can contact also:Director, Quality and Regulatory AffairsTarja.paivinen@instrudental.comTel. +358 40 719 0513After sales ManagerMika.Kyllonen@instrudental.comTel. + 358 40 080 5305 |  | 101 | 101 units (US Distribution) | Worldwide Distribution: US (Nationwide) including states of: AK, AZ, CA, CO, GA, IL, KY, MD, MT, NC, NV, OR, PA, TN, TX, UT, VA, WA,and WI; and Internationally to: Belgium Brazil, Chile, China, Czech Republic, France, Germany, Italy, Spain, Sweden, Switzerland, and United Kingdom. | Terminated | May 22, 2015 | 493 | 101 | Computer
Z-1086-2014 | 67388 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Specialty Straight Static Inserter Assembly |  |  | December 18, 2013 | 2013 | Stryker Spine | Stryker has received complaints from customers relating to fracture along the weld that binds the handle to the shaft of the static inserter. | Device Design | Stryker sent Urgent Medical Device Removal Notification letter and to their customers via email on 12/18/2013. The letter identified the affected product, problem and actions to be taken. Customers were instructed to examine their inventory, return affected product using enclosed pre-paid mailing label, complete Customer Response Form and fax to Regulatory Compliance at 855-632-9049 or mail it to Spine-RegulatoryActions@Stryker.com. For questions call 201-760-8287, 201-760-8298 or 201-760-8041. |  | 21 | 21 units | US Nationwide Distribution | Terminated | August 25, 2014 | 250 | 21 | Computer
Z-2068-2014 | 68344 | No | CCK | Carbon dioxide gas analyzer. | Anesthesiology | Anesthesiology | 510(k) | DATEX-OHMEDA S/5 SINGLE-WIDTH AIRWAY MODULE, E-MINIC AND ACCESSORIES. |  |  | June 11, 2014 | 2014 | GE Healthcare, LLC | Failure of the CO2 detector in Single-width Airway and Extension modules, which may cause a slow continuous decrease of measured Et-/FiCO2 values. Incorrect Et CO2/Fi CO2 value may impair clinical decision making for both mechanically and spontaneously ventilated patients because of incorrect low CO2 values. | Device Design | GE Healthcare sent an Urgent Medical Device Correction letter dated June 11, 2014, to all affected consignees. The letter described the Safety Issue, Safety Instructions, Affected Product Details, and Product Correction and Contact Information. For questions customers were instructed to contact Technical Support at 1-800-558-7044 or their local Service Representative. |  | 2948 | 2,948 (761 units US, 2,187 units OUS) | Worldwide Distribution - USA nationwide, ARGENTINA, AUSTRALIA, AUSTRIA, BAHRAIN, BRAZIL, CANADA, CHILE, CHINA, Croatia, CZECH REPUBLIC, DENMARK, EGYPT, EL SALVADOR, FINLAND, FRANCE, GERMANY, GREECE, GUATEMALA, HONG KONG, INDIA, IRAQ, ISRAEL, ITALY, JAPAN, KAZAKHSTAN, REPUBLIC OF KOREA, MALAYSIA, NETHERLANDS, NEW ZEALAND, OMAN, PERU, POLAND, ROMANIA, RUSSIAN FEDERATION, SAUDI ARABIA, SINGAPORE, SPAIN, SWEDEN, SWITZERLAND, SYRIAN ARAB REPUBLIC, TAIWAN, TURKEY, UNITED ARAB EMIRATES, UNITED KINGDOM, VENEZUELA, LEBANON, REPUBLIC OF MALDOVIA, COSTA RICA, GABON, INDONESIA, JAMAICA, NIGERIA, NORWAY, NEPAL, PAKISTAN, PANAMA, PHILIPPINES, SENEGAL, SOUTH AFRICA, THAILAND, TUNISIA, and CAMBODIA. | Terminated | March 29, 2016 | 657 | 2948 | Computer
Z-2075-2014 | 68395 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Siemens syngo RT Dosimetrist 2.7 system |  |  | May 15, 2014 | 2014 | Siemens Medical Solutions USA, Inc | An update to the Virtual Simulation software of the syngo RT Dosimetrist is necessary to address a safety issue. In special scenarios within the Structure Operations function the system did not store the selected name from the list of predefined names. After reloading of the data, this structure name had been changed by the system. Additionally, a performance issue was addressed where import of DI | Software design | Siemens sent a Medical Device Correction Field Safety Notice of Improvement letter to affected customers via USPS Certified Mail. The letter described the affected product, problem, and actions to be taken. |  | 7 | 7 | US Nationwide Distribution in the states of NH, TN, OH, VA, and TX. | Terminated | December 15, 2014 | 214 | 7 | Computer
Z-2071-2014 | 68353 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Siemens RAPIDCommÂż Data Management System. |  |  | May 22, 2014 | 2014 | Siemens Healthcare Diagnostics Inc | RAPIDCommÂż 5.0 Screen Layout will display incorrect test names | Software design | Siemens Healthcare issued an Urgent Field Safety Notice by e-mail to all affected Siemens Healthcare Diagnostics Regional offices on May 22, 2014, both in the United States and elsewhere for communication with affected customers. This notice informs customers of the issue with the RAPIDCommÂż Data Management System (Version 5.0.A software correction will be issued to address the problem.Please retain this letter with your laboratory records. Complete and return the Field Correction Effectiveness Check attached to this letter within 30 days.We apologize for the inconvenience this situation has caused. If you have any questions, please contact your Siemens Customer Care Center or your local Siemens technical support representative.For further questions please call (781) 269-3000. |  | 6 | 6 systems | Worldwide Distribution - US including the states of OR and PA., and the countries of Denmark and UK. | Terminated | May 26, 2015 | 369 | 6 | Computer
Z-2059-2014 | 68582 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | V-Series Patient Monitoring System, Manufactured for Mindray DS USA, Inc. Used for monitoring of human physiological parameters. The V-Series Monitor also has the capability of performing IV Drug and Hemodynamic Calculations and interfacing with network devices. |  |  | May 19, 2014 | 2014 | Mindray DS USA, Inc. dba Mindray North America | ECG interpretation feature issues. False measurements reported when the algorithm cannot determine P wave or QRS axis. The Median Complex Report displays the representative beat 25% larger than actual signal input. When a patient date of birth is imported from an external source, the ECG interpretation algorithm uses the default age of 50, and interpretive statements may be incorrect. | Software design | Mindray issued "Urgent: V-Series Monitoring System Corrective Action" notices dated May 19, 2014, to customers that received the affected product via certified mail with return receipt verification. The notification informs the customer of the issues with the affected product and how to identify the affected product. The notification recommends that customers discontinue use of the 12-lead ECG interpretation feature until a software correction is available (anticipate August 2014). A point of contact is provided to customers in case they have questions.For further questions please call 1-800-288-2121. |  | 113 | 113 units (57 United States; 56 International) | Worldwide Distribution - US including the states of Louisiana, New Jersey, Montana, South Carolina and Wyoming. | Open, Classified |  |  | 113 | Computer
Z-2060-2014 | 68578 | No | GKP | Coagulation instrument. | Hematology | Hematology | 510(k) | ACL TOP (Base), automated coagulation laboratory instrument. |  |  | June 03, 2014 | 2014 | Instrumentation Laboratory Co. | Potential for sample misidentification. | Device Design | Consignees were notified via phone starting on June 3, 2014. On June 4, a USPS mailing was sent to all US/Canadian customers. On June 4, 2014, email notifications were sent to the IL Affiliates and International Distributors/Dealers (through the International Group) requesting them to translate the Field Safety Notice and provide notification to their customer base and secondary distributors.The Urgent Product Correction letter informs consignees of the issue and provides workaround instructions. Consignees are asked to complete a Mandatory Response Tracking Record and return it via Fax (781)861-4207 or email: ra-usa@ilww.com. For technical questions, consignees should contact the IL Technical Support Center at 1-800-678-0710. |  | 1074 | 1074 | Worldwide Distribution-USA (nationwide) including the states of AL, AR, AZ, CA, CO, CT, DE, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, and WV, and the countries of Canada, Mexico, Argentina, Australia, Austria, Belarus, Belgium, Bolivia, Brazil, Brunei, Chile, China, Costa Rica, Croatia, Czech Republic, Denmark, Dominican Republic, El Salvador, Finland, France, Germany, Gibraltar, Great Britain, Greece, Guatemala, Hong Kong, Hungary, India, Indonesia, Ireland, Israel, Italy, Japan, Korea, Kuwait, Libya, Lithuania, Luxembourg, Malaysia, Mexico, Netherlands, New Zealand, Norway, Panama, Peru, Philippines, Poland, Portugal, Quatar, Romania, Russia, Saudi Arabia, Serbia, Slovakia, Slovenia, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, Tunisia, Turkey, UAE, Uruguay, Venezuela, Vietnam. | Terminated | June 11, 2015 | 373 | 6647 | Computer
Z-0620-2014 | 64807 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | The Centricity Laboratory System is intended to be an information system designed to support the clinical and administrative activities associated with the provision and utilization of clinical laboratory services and facilities, e.g., the storing and delivering of analytical results. It is a specially designed data program application (software), which is supplied for installation in existing mainframe or decentralized computers or a computer network. This software system will not be the software directly used to run, steer, or control any specific laboratory analyzer or equipment. It is intended for use in laboratory facilities, including a central laboratory or in a multiple laboratory environment servicing satellite laboratories, reference laboratories, clinics, etc. Centricity Laboratory System is not intended for direct patient contact. |  |  | March 12, 2013 | 2013 | GE Healthcare It | GE Healthcare is aware of a potential safety issue with the use of GE Centricity Laboratory Instrument Interface (IF) where free-text sent from IM is not being transcribed into free-text in iNET. The instrument may upload results to IF/iNET when order/request is not yet in IF/iNET. In such a situation, test level supporting free-text sent from the instrument is discarded (not captured). | Device Design | GE Healthcare sent an URGENT MEDICAL DEVICE CORRECTION letter dated March 12, 2013 to all affected customers. The letter notified customers of the problem, suggested that customers experiencing this problem should retransmit the result from the instrument once the order exists in iNET and verify that the item level free-text is present in iNET, and to contact their GE service representative to request or install the updated software. Customers with additional questions can contact the GE Healthcare IT Help Desk at 1-847-589-8496 or tgg.nahelpdesk@med.ge.com.For questions regarding this recall call 847-589-8496. |  | 5 | 5 units | Worldwide Distribution - US including MN and Internationally to India, Australia, Canada, and England. | Open, Classified |  |  | 5 | Computer